## YSI MPS
YSI 556 MPS Audit Log
Date | 2005-03-18 | Time
Name | Drexell Barnes - City of Bend
Meter #
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
temperature | COB SN 1622 | 26.3 | 26.3 | 0 | 2 | Pass
conductivity | COB YSI 3200 | 152.4 | 146 | 6.4 | 22.86 | Pass
pH | COB Orion SN 27985 | 7.01 | 6.92 | 0.09 | 0.5 | Pass
dissolved oxygen | Winkler | 7.02 | 7.07 | -0.05 | 1.0 mg/L | Pass
dissolved oxygen |  |  | 78.4% Sat
Comments: Audits performed at the City of Bend Water Quality Laboratory

## Hach Turbidity
Hach 2100P Audit Log
Date | 2005-03-18 | Time
Name | Drexell Barnes - City of Bend
Meter #
COB Hach 2100P SN 040300035008
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
Turbidity 0-1000 | Gelex secondary | 520 | 521 | -1 | 156 | Pass
Turbidity 0-100 | Gelex secondary | 52.7 | 52.7 | 0 | 15.81 | Pass
Turbidity 0-10 | Gelex secondary | 5.83 | 5.8 | 0.03 | 1.749 | Pass
Comments: Audits performed at the City of Bend Water Quality Laboratory