## YSI MPS
YSI 556 MPS Audit Log
Date | 2005-06-14 | Time | 8:26
Name | Drexell Barnes - City of Bend
Meter #
SN 04C1624 AX
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
temperature | COB SN 1622 | 18.7 | 18.68 | 0.02 | 2 | PASS
conductivity | COB YSI 3200 | 146 | 148 | -2 | 21.9 | PASS
pH | COB Orion SN 27985 | 7.02 | 7.06 | -0.04 | 0.5 | PASS
dissolved oxygen | Winkler | 7.4 | 7.42 | -0.02 | 1.0 mg/L | PASS
dissolved oxygen
Comments: Analysis performed in City of Bend Water Quality Laboratory

## Hach Turbidity
Hach 2100P Audit Log
Date | 2005-06-14 | Time | 8:30
Name | Drxell Barnes - City of Bend
Meter #
COB Hach 2100P SN 040300035008
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
Turbidity 0-1000 | Gelex secondary | 520 | 523 | -3 | 156 | PASS
Turbidity 0-500 | Gelex secondary | 52.7 | 52.8 | -0.1 | 15.81 | PASS
Turbidity 0-5 | Gelex secondary | 5.83 | 5.89 | -0.06 | 1.749 | PASS
Comments: Analysis performed in City of Bend Water Quality Laboratory