## YSI MPS
YSI 556 MPS Audit Log
Date | 2005-09-12 | Time | 8:35
Name | Drexell Barnes - COB Water Quality Lab
Meter #
SN 04C1624 AX
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
temperature | COB SN 1622 | 18 | 17.96 | 0.04 | 2 | PASS
conductivity | COB YSI 3200 | 145.7 | 146 | -0.3 | 21.855 | PASS
pH | COB Orion SN 27985 | 7.01 | 7.01 | 0 | 0.5 | PASS
dissolved oxygen | Winkler | 7.9 | 8.1 | -0.2 | 1.0 mg/L | PASS
dissolved oxygen |  |  | 87.1%
Comments: Analysis performed in City of Bend Water Quality Laboratory

## Hach Turbidity
Hach 2100P Audit Log
Date | 2005-09-13 | Time | 7:35
Name | Drexell Barnes - COB Water Quality Lab
Meter #
COB Hach 2100P SN 040300035008
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
Turbidity 0-1000 | Gelex secondary | 520 | 524 | -4 | 156 | PASS
Turbidity 0-500 | Gelex secondary | 52.7 | 53.4 | -0.7 | 15.81 | PASS
Turbidity 0-5 | Gelex secondary | 5.83 | 5.95 | -0.12 | 1.749 | PASS
Comments: Analysis performed in City of Bend Water Quality Laboratory