## YSI MPS
YSI 556 MPS Audit Log
Date | 2005-09-15 | Time | 7:15
Name | Drexell Barnes - COB Water Quality Lab
Meter #
SN 04C1624 AX
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
temperature | COB SN 1622 | 18.6 | 18.55 | 0.05 | 2 | PASS
conductivity | COB YSI 3200 | 144.1 | 146 | -1.9 | 21.615 | PASS
pH | COB Orion SN 27985 | 7.05 | 7.07 | -0.02 | 0.5 | PASS
dissolved oxygen | Winkler | 8.2 | 8.43 | -0.23 | 1.0 mg/L | PASS
dissolved oxygen |  |  | 87.9%
Comments: Analysis performed in City of Bend Water Quality Laboratory

## Hach Turbidity
Hach 2100P Audit Log
Date | 2005-09-15 | Time | 7:35
Name | Drexell Barnes - COB Water Quality Lab
Meter #
COB Hach 2100P SN 040300035008
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
Turbidity 0-1000 | Gelex secondary | 520 | 528 | -8 | 156 | PASS
Turbidity 0-500 | Gelex secondary | 52.7 | 53.4 | -0.7 | 15.81 | PASS
Turbidity 0-5 | Gelex secondary | 5.83 | 5.93 | -0.1 | 1.749 | PASS
Comments: Analysis performed in City of Bend Water Quality Laboratory