## YSI MPS
YSI 556 MPS Audit Log
Date | 2005-07-15 | Time | 16:35
Name | Drexell Barnes - COB Water Quality Lab
Meter #
SN 04C1624 AX
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
temperature | COB SN 1622 | 25.1 | 25.1 | 0 | 2 | PASS
conductivity | COB YSI 3200 | 141 | 147 | -6 | 21.15 | PASS
pH | COB Orion SN 27985 | 6.98 | 7.1 | -0.12 | 0.5 | PASS
dissolved oxygen | Winkler | 6.7 | 6.31 | 0.39 | 1.0 mg/L | PASS
dissolved oxygen |  |  | 75.1%
Comments: Analysis performed in City of Bend Water Quality Laboratory

## Hach Turbidity
Hach 2100P Audit Log
Date | 2005-07-15 | Time | 16:44
Name | Drexell Barnes - COB Water Quality Lab
Meter #
COB Hach 2100P SN 040300035008
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
Turbidity 0-1000 | Gelex secondary | 520 | 523 | -3 | 156 | PASS
Turbidity 0-500 | Gelex secondary | 52.7 | 53.1 | -0.4 | 15.81 | PASS
Turbidity 0-5 | Gelex secondary | 5.83 | 5.89 | -0.06 | 1.749 | PASS
Comments: Analysis performed in City of Bend Water Quality Laboratory