## YSI MPS
YSI 556 MPS Audit Log
Date | 2005-07-22 | Time | 13:13
Name | Drexell Barnes - COB Water Quality Lab
Meter #
SN 04C1624 AX
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
temperature | COB SN 1622 | 22.6 | 22.67 | -0.07 | 2 | PASS
conductivity | COB YSI 3200 | 146.4 | 147 | -0.6 | 21.96 | PASS
pH | COB Orion SN 27985 | 7.01 | 7.02 | -0.01 | 0.5 | PASS
dissolved oxygen | Winkler | 7.2 | 7.4 | -0.2 | 1.0 mg/L | PASS
dissolved oxygen |  |  | 83.9%
Comments: Analysis performed in City of Bend Water Quality Laboratory

## Hach Turbidity
Hach 2100P Audit Log
Date | 2005-07-22 | Time | 13:13
Name | Drexell Barnes - COB Water Quality Lab
Meter #
COB Hach 2100P SN 040300035008
Parameter | Reference | Reference Reading | Reading | Difference | Criteria +/- | Precision Pass / Fail
Turbidity 0-1000 | Gelex secondary | 520 | 524 | -4 | 156 | PASS
Turbidity 0-500 | Gelex secondary | 52.7 | 53.1 | -0.4 | 15.81 | PASS
Turbidity 0-5 | Gelex secondary | 5.83 | 5.91 | -0.08 | 1.749 | PASS
Comments: Analysis performed in City of Bend Water Quality Laboratory