## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River at Lower Bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 #AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-03 |  | 2007-04-03 |  | A
Time | Computer set to atomic time |  | 15:13 |  | 15:13 |  | A
Temp | VWR NIST traceable surface water |  | 13.1 |  | 13.3 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 11.5 |  | A
SpC | 148 µS/mL |  | 149 |  | 149 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2007-04-04 12:31 | 7.5 |  | 12.71 | 106 | 63 | 40.9
| 2007-04-04 12:31 | 7.46 | 7.73 |  |  | 63 | 41
Grade |  | A | OR |  |  | A
Audit 2 | 2007-04-09 16:22 | 10.89 |  | 11.42 | 103.3 | 61.5 | 40
| 2007-04-09 16:31 | 10.84 | 8.03 |  |  | 60.1 | 39.1
Grade |  | A | OR |  |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
OR = over ride |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River at Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 #05G1083 AM
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-03 |  | 2007-04-03 |  | A
Time | Computer set to atomic time |  | 17:15 |  | 17:15 |  | A
Temp | VWR NIST traceable surface water |  | 9.7 |  | 9.7 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 12.1 |  | B
SpC | 148 µS/mL |  | 149 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2007-04-04 12:31 | 7.5 |  | 12.71 | 106 | 63 | 40.9
2
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-04 |  | 2007-04-04 |  | A
Time | Computer set to atomic time |  | 16:50 |  | 16:51 |  | A
Temp | VWR NIST traceable surface water |  | 16.2 |  | 15.7 |  | A
DO | Winkler titration of surface water |  | 11.1 |  | 11.3 |  | A
SpC | 148 µS/mL |  | 149 |  | 138 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River at Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 #05G1083 AM
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-09 |  | 2007-04-09 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 15 |  | 15.1 |  | A
DO | Winkler titration of surface water |  | 11.4 |  | 11.3 |  | A
SpC | 148 µS/mL |  | 149 |  | 142 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2007-04-09 16:22 | 10.89 | 3.63 | 11.42 | 103.3 | 61.5 | 40
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-16 |  | 2007-04-16 |  | A
Time | Computer set to atomic time |  | 13:10 |  | 13:11 |  | A
Temp | VWR NIST traceable surface water |  | 15.3 |  | 15.4 |  | A
DO | Winkler titration of surface water |  | 10.6 |  | 10.6 |  | A
SpC | 148 µS/mL |  | 149 |  | 139 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO % | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2007-04-04 | 12:31 | 7.46 | 7.73 |  |  | 63 | 41
2007-04-04 | 12:46 | 7.62 | 7.75 |  |  | 62.9 | 40.9
2007-04-04 | 13:00 | 7.77 | 7.78 |  |  | 62.9 | 40.9
2007-04-04 | 13:15 | 7.94 | 7.8 |  |  | 62.9 | 40.9
2007-04-04 | 13:30 | 8.08 | 7.83 |  |  | 62.9 | 40.9
2007-04-04 | 13:45 | 8.22 | 7.85 |  |  | 62.9 | 40.9
2007-04-04 | 14:00 | 8.31 | 7.87 |  |  | 62.9 | 40.9
2007-04-04 | 14:15 | 8.37 | 7.89 |  |  | 62.9 | 40.9
2007-04-04 | 14:30 | 8.48 | 7.91 |  |  | 62.9 | 40.9
2007-04-04 | 14:45 | 8.56 | 7.92 |  |  | 62.9 | 40.9
2007-04-04 | 15:00 | 8.69 | 7.93 |  |  | 62.9 | 40.9
2007-04-04 | 15:15 | 8.73 | 7.93 |  |  | 62.9 | 40.9
2007-04-04 | 15:30 | 8.77 | 7.93 |  |  | 62.9 | 40.9
2007-04-04 | 15:45 | 8.85 | 7.94 |  |  | 62.9 | 40.9
2007-04-04 | 16:00 | 8.99 | 7.95 |  |  | 62.9 | 40.9
2007-04-04 | 16:15 | 9.11 | 7.96 |  |  | 62.9 | 40.9
2007-04-04 | 16:30 | 9.2 | 7.96 |  |  | 62.9 | 40.9
2007-04-04 | 16:45 | 9.21 | 7.95 |  |  | 62.9 | 40.9
2007-04-04 | 17:00 | 9.21 | 7.94 |  |  | 63 | 40.9
2007-04-04 | 17:15 | 9.21 | 7.93 |  |  | 63 | 40.9
2007-04-04 | 17:30 | 9.26 | 7.93 |  |  | 63 | 40.9
2007-04-04 | 17:45 | 9.29 | 7.93 |  |  | 63 | 41
2007-04-04 | 18:00 | 9.35 | 7.93 |  |  | 63 | 41
2007-04-04 | 18:15 | 9.38 | 7.92 |  |  | 63.1 | 41
2007-04-04 | 18:30 | 9.39 | 7.91 |  |  | 63.1 | 41
2007-04-04 | 18:45 | 9.37 | 7.89 |  |  | 63.2 | 41.1
2007-04-04 | 19:00 | 9.34 | 7.87 |  |  | 63.2 | 41.1
2007-04-04 | 19:15 | 9.3 | 7.84 |  |  | 63.3 | 41.2
2007-04-04 | 19:30 | 9.25 | 7.82 |  |  | 63.5 | 41.3
2007-04-04 | 19:45 | 9.21 | 7.79 |  |  | 63.5 | 41.3
2007-04-04 | 20:00 | 9.16 | 7.77 |  |  | 63.5 | 41.3
2007-04-04 | 20:15 | 9.12 | 7.75 |  |  | 63.6 | 41.3
2007-04-04 | 20:30 | 9.09 | 7.73 |  |  | 63.6 | 41.4
2007-04-04 | 20:45 | 9.07 | 7.72 |  |  | 63.7 | 41.4
2007-04-04 | 21:00 | 9.06 | 7.71 |  |  | 63.7 | 41.4
2007-04-04 | 21:15 | 9.05 | 7.7 |  |  | 63.7 | 41.4
2007-04-04 | 21:30 | 9.05 | 7.69 |  |  | 63.7 | 41.4
2007-04-04 | 21:45 | 9.05 | 7.68 |  |  | 63.8 | 41.5
2007-04-04 | 22:00 | 9.04 | 7.67 |  |  | 63.8 | 41.5
2007-04-04 | 22:15 | 9.03 | 7.67 |  |  | 63.8 | 41.5
2007-04-04 | 22:30 | 9.02 | 7.66 |  |  | 63.8 | 41.5
2007-04-04 | 22:45 | 9 | 7.66 |  |  | 63.9 | 41.5
2007-04-04 | 23:00 | 8.98 | 7.65 |  |  | 63.9 | 41.5
2007-04-04 | 23:15 | 8.95 | 7.64 |  |  | 63.9 | 41.5
2007-04-04 | 23:30 | 8.92 | 7.64 |  |  | 63.9 | 41.5
2007-04-04 | 23:45 | 8.88 | 7.64 |  |  | 63.9 | 41.5
2007-04-05 | 0:00 | 8.84 | 7.63 |  |  | 63.9 | 41.5
2007-04-05 | 0:15 | 8.8 | 7.63 |  |  | 63.9 | 41.5
2007-04-05 | 0:30 | 8.76 | 7.63 |  |  | 63.9 | 41.5
2007-04-05 | 0:45 | 8.72 | 7.62 |  |  | 64.3 | 41.8
2007-04-05 | 1:01 | 8.68 | 7.62 |  |  | 64.3 | 41.8
2007-04-05 | 1:16 | 8.65 | 7.62 |  |  | 63.9 | 41.6
2007-04-05 | 1:31 | 8.62 | 7.61 |  |  | 63.9 | 41.6
2007-04-05 | 1:46 | 8.58 | 7.61 |  |  | 63.9 | 41.6
2007-04-05 | 2:01 | 8.55 | 7.61 |  |  | 63.9 | 41.5
2007-04-05 | 2:16 | 8.52 | 7.6 |  |  | 63.9 | 41.5
2007-04-05 | 2:31 | 8.48 | 7.6 |  |  | 63.8 | 41.5
2007-04-05 | 2:46 | 8.44 | 7.6 |  |  | 63.8 | 41.5
2007-04-05 | 3:01 | 8.39 | 7.59 |  |  | 63.8 | 41.4
2007-04-05 | 3:16 | 8.33 | 7.59 |  |  | 63.7 | 41.4
2007-04-05 | 3:31 | 8.27 | 7.59 |  |  | 63.7 | 41.4
2007-04-05 | 3:46 | 8.21 | 7.59 |  |  | 63.6 | 41.4
2007-04-05 | 4:01 | 8.15 | 7.59 |  |  | 63.6 | 41.3
2007-04-05 | 4:16 | 8.1 | 7.59 |  |  | 63.6 | 41.3
2007-04-05 | 4:31 | 8.04 | 7.59 |  |  | 63.5 | 41.3
2007-04-05 | 4:46 | 8 | 7.59 |  |  | 63.5 | 41.3
2007-04-05 | 5:01 | 7.96 | 7.58 |  |  | 63.5 | 41.3
2007-04-05 | 5:16 | 7.93 | 7.58 |  |  | 63.4 | 41.2
2007-04-05 | 5:31 | 7.9 | 7.58 |  |  | 63.4 | 41.2
2007-04-05 | 5:46 | 7.86 | 7.58 |  |  | 63.4 | 41.2
2007-04-05 | 6:01 | 7.83 | 7.58 |  |  | 63.3 | 41.2
2007-04-05 | 6:16 | 7.8 | 7.58 |  |  | 63.3 | 41.1
2007-04-05 | 6:31 | 7.77 | 7.58 |  |  | 63.3 | 41.1
2007-04-05 | 6:46 | 7.74 | 7.58 |  |  | 63.2 | 41.1
2007-04-05 | 7:01 | 7.7 | 7.58 |  |  | 63.2 | 41.1
2007-04-05 | 7:16 | 7.66 | 7.59 |  |  | 63.2 | 41.1
2007-04-05 | 7:31 | 7.63 | 7.59 |  |  | 63.1 | 41
2007-04-05 | 7:46 | 7.6 | 7.6 |  |  | 63.1 | 41
2007-04-05 | 8:01 | 7.58 | 7.61 |  |  | 63.1 | 41
2007-04-05 | 8:16 | 7.57 | 7.62 |  |  | 63.1 | 41
2007-04-05 | 8:31 | 7.57 | 7.63 |  |  | 63.1 | 41
2007-04-05 | 8:46 | 7.59 | 7.65 |  |  | 63 | 41
2007-04-05 | 9:01 | 7.62 | 7.67 |  |  | 63 | 41
2007-04-05 | 9:16 | 7.66 | 7.68 |  |  | 63 | 40.9
2007-04-05 | 9:31 | 7.72 | 7.7 |  |  | 63 | 40.9
2007-04-05 | 9:46 | 7.78 | 7.72 |  |  | 63 | 40.9
2007-04-05 | 10:01 | 7.84 | 7.73 |  |  | 63 | 40.9
2007-04-05 | 10:16 | 7.91 | 7.75 |  |  | 62.9 | 40.9
2007-04-05 | 10:31 | 7.98 | 7.77 |  |  | 62.9 | 40.9
2007-04-05 | 10:46 | 8.05 | 7.78 |  |  | 62.9 | 40.9
2007-04-05 | 11:01 | 8.14 | 7.79 |  |  | 62.9 | 40.9
2007-04-05 | 11:16 | 8.23 | 7.81 |  |  | 62.9 | 40.9
2007-04-05 | 11:31 | 8.34 | 7.82 |  |  | 62.9 | 40.9
2007-04-05 | 11:46 | 8.46 | 7.84 |  |  | 62.9 | 40.9
2007-04-05 | 12:01 | 8.58 | 7.85 |  |  | 62.8 | 40.8
2007-04-05 | 12:16 | 8.7 | 7.87 |  |  | 62.8 | 40.8
2007-04-05 | 12:31 | 8.83 | 7.89 |  |  | 62.8 | 40.8
2007-04-05 | 12:46 | 8.97 | 7.91 |  |  | 62.8 | 40.8
2007-04-05 | 13:00 | 9.11 | 7.91 |  |  | 62.8 | 40.8
2007-04-05 | 13:15 | 9.24 | 7.93 |  |  | 62.8 | 40.8
2007-04-05 | 13:30 | 9.37 | 7.94 |  |  | 62.7 | 40.8
2007-04-05 | 13:45 | 9.51 | 7.96 |  |  | 62.7 | 40.8
2007-04-05 | 14:00 | 9.65 | 7.97 |  |  | 62.7 | 40.8
2007-04-05 | 14:15 | 9.77 | 7.99 |  |  | 62.7 | 40.8
2007-04-05 | 14:30 | 9.9 | 7.99 |  |  | 62.7 | 40.8
2007-04-05 | 14:45 | 10.02 | 8 |  |  | 62.7 | 40.8
2007-04-05 | 15:00 | 10.12 | 8.01 |  |  | 62.7 | 40.8
2007-04-05 | 15:15 | 10.22 | 8.02 |  |  | 62.7 | 40.8
2007-04-05 | 15:30 | 10.29 | 8.03 |  |  | 62.7 | 40.8
2007-04-05 | 15:45 | 10.37 | 8.03 |  |  | 62.8 | 40.8
2007-04-05 | 16:00 | 10.43 | 8.04 |  |  | 62.8 | 40.8
2007-04-05 | 16:15 | 10.5 | 8.04 |  |  | 62.8 | 40.8
2007-04-05 | 16:30 | 10.55 | 8.04 |  |  | 62.8 | 40.8
2007-04-05 | 16:45 | 10.6 | 8.04 |  |  | 62.8 | 40.8
2007-04-05 | 17:00 | 10.63 | 8.03 |  |  | 62.8 | 40.8
2007-04-05 | 17:15 | 10.66 | 8.03 |  |  | 62.9 | 40.9
2007-04-05 | 17:30 | 10.67 | 8.02 |  |  | 62.9 | 40.9
2007-04-05 | 17:45 | 10.68 | 8.02 |  |  | 62.9 | 40.9
2007-04-05 | 18:00 | 10.68 | 8 |  |  | 62.9 | 40.9
2007-04-05 | 18:15 | 10.66 | 7.99 |  |  | 63 | 40.9
2007-04-05 | 18:30 | 10.63 | 7.96 |  |  | 63 | 41
2007-04-05 | 18:45 | 10.58 | 7.94 |  |  | 63.1 | 41
2007-04-05 | 19:00 | 10.54 | 7.91 |  |  | 63.1 | 41
2007-04-05 | 19:15 | 10.49 | 7.88 |  |  | 63.2 | 41.1
2007-04-05 | 19:30 | 10.45 | 7.85 |  |  | 63.2 | 41.1
2007-04-05 | 19:45 | 10.4 | 7.82 |  |  | 63.3 | 41.1
2007-04-05 | 20:00 | 10.37 | 7.8 |  |  | 63.3 | 41.1
2007-04-05 | 20:15 | 10.33 | 7.77 |  |  | 63.3 | 41.2
2007-04-05 | 20:30 | 10.3 | 7.76 |  |  | 63.4 | 41.2
2007-04-05 | 20:45 | 10.27 | 7.74 |  |  | 63.4 | 41.2
2007-04-05 | 21:00 | 10.24 | 7.72 |  |  | 63.4 | 41.2
2007-04-05 | 21:15 | 10.21 | 7.7 |  |  | 63.4 | 41.2
2007-04-05 | 21:30 | 10.19 | 7.69 |  |  | 63.4 | 41.2
2007-04-05 | 21:45 | 10.17 | 7.69 |  |  | 63.4 | 41.2
2007-04-05 | 22:00 | 10.16 | 7.68 |  |  | 63.4 | 41.2
2007-04-05 | 22:15 | 10.14 | 7.67 |  |  | 63.4 | 41.2
2007-04-05 | 22:30 | 10.13 | 7.66 |  |  | 63.3 | 41.2
2007-04-05 | 22:45 | 10.12 | 7.66 |  |  | 63.3 | 41.1
2007-04-05 | 23:00 | 10.1 | 7.65 |  |  | 63.3 | 41.1
2007-04-05 | 23:15 | 10.08 | 7.64 |  |  | 63.3 | 41.1
2007-04-05 | 23:30 | 10.05 | 7.63 |  |  | 63.3 | 41.1
2007-04-05 | 23:45 | 10.03 | 7.63 |  |  | 63.3 | 41.1
2007-04-06 | 0:00 | 10 | 7.62 |  |  | 63.4 | 41.2
2007-04-06 | 0:15 | 9.96 | 7.62 |  |  | 63.3 | 41.1
2007-04-06 | 0:30 | 9.92 | 7.62 |  |  | 63.3 | 41.1
2007-04-06 | 0:45 | 9.89 | 7.61 |  |  | 63.3 | 41.1
2007-04-06 | 1:01 | 9.85 | 7.61 |  |  | 63.3 | 41.1
2007-04-06 | 1:16 | 9.82 | 7.6 |  |  | 63.3 | 41.1
2007-04-06 | 1:31 | 9.78 | 7.6 |  |  | 63.3 | 41.1
2007-04-06 | 1:46 | 9.75 | 7.6 |  |  | 63.3 | 41.1
2007-04-06 | 2:01 | 9.71 | 7.59 |  |  | 63.2 | 41.1
2007-04-06 | 2:16 | 9.68 | 7.59 |  |  | 63.2 | 41.1
2007-04-06 | 2:31 | 9.64 | 7.59 |  |  | 63.2 | 41.1
2007-04-06 | 2:46 | 9.61 | 7.58 |  |  | 63.2 | 41.1
2007-04-06 | 3:01 | 9.57 | 7.59 |  |  | 63.2 | 41.1
2007-04-06 | 3:16 | 9.53 | 7.58 |  |  | 63.2 | 41.1
2007-04-06 | 3:31 | 9.5 | 7.58 |  |  | 63.2 | 41.1
2007-04-06 | 3:46 | 9.46 | 7.58 |  |  | 63.2 | 41.1
2007-04-06 | 4:01 | 9.42 | 7.58 |  |  | 63.2 | 41.1
2007-04-06 | 4:16 | 9.38 | 7.58 |  |  | 63.1 | 41
2007-04-06 | 4:31 | 9.35 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 4:46 | 9.31 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 5:01 | 9.28 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 5:16 | 9.25 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 5:31 | 9.22 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 5:46 | 9.2 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 6:01 | 9.17 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 6:16 | 9.15 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 6:31 | 9.12 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 6:46 | 9.1 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 7:01 | 9.08 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 7:16 | 9.06 | 7.57 |  |  | 63.1 | 41
2007-04-06 | 7:31 | 9.04 | 7.58 |  |  | 63 | 41
2007-04-06 | 7:46 | 9.03 | 7.58 |  |  | 63 | 41
2007-04-06 | 8:01 | 9.01 | 7.59 |  |  | 63 | 41
2007-04-06 | 8:16 | 9 | 7.61 |  |  | 63 | 41
2007-04-06 | 8:31 | 9.01 | 7.63 |  |  | 63 | 40.9
2007-04-06 | 8:46 | 9.03 | 7.64 |  |  | 63 | 40.9
2007-04-06 | 9:01 | 9.06 | 7.66 |  |  | 63 | 40.9
2007-04-06 | 9:16 | 9.1 | 7.68 |  |  | 63 | 40.9
2007-04-06 | 9:31 | 9.15 | 7.7 |  |  | 62.9 | 40.9
2007-04-06 | 9:46 | 9.21 | 7.72 |  |  | 62.9 | 40.9
2007-04-06 | 10:01 | 9.27 | 7.74 |  |  | 63 | 40.9
2007-04-06 | 10:16 | 9.34 | 7.75 |  |  | 63 | 40.9
2007-04-06 | 10:31 | 9.41 | 7.78 |  |  | 63 | 40.9
2007-04-06 | 10:46 | 9.49 | 7.79 |  |  | 63 | 40.9
2007-04-06 | 11:01 | 9.58 | 7.81 |  |  | 63 | 40.9
2007-04-06 | 11:16 | 9.68 | 7.82 |  |  | 63 | 40.9
2007-04-06 | 11:31 | 9.79 | 7.84 |  |  | 63 | 40.9
2007-04-06 | 11:46 | 9.9 | 7.87 |  |  | 63 | 40.9
2007-04-06 | 12:01 | 10.02 | 7.88 |  |  | 63 | 40.9
2007-04-06 | 12:16 | 10.14 | 7.9 |  |  | 63 | 40.9
2007-04-06 | 12:31 | 10.28 | 7.91 |  |  | 63 | 40.9
2007-04-06 | 12:46 | 10.41 | 7.93 |  |  | 62.9 | 40.9
2007-04-06 | 13:00 | 10.54 | 7.95 |  |  | 62.9 | 40.9
2007-04-06 | 13:15 | 10.69 | 7.97 |  |  | 62.9 | 40.9
2007-04-06 | 13:30 | 10.83 | 7.98 |  |  | 62.9 | 40.9
2007-04-06 | 13:45 | 10.98 | 7.99 |  |  | 62.9 | 40.9
2007-04-06 | 14:00 | 11.12 | 8.01 |  |  | 62.9 | 40.9
2007-04-06 | 14:15 | 11.25 | 8.03 |  |  | 62.9 | 40.9
2007-04-06 | 14:30 | 11.39 | 8.03 |  |  | 62.9 | 40.9
2007-04-06 | 14:45 | 11.5 | 8.05 |  |  | 62.9 | 40.9
2007-04-06 | 15:00 | 11.62 | 8.06 |  |  | 62.9 | 40.9
2007-04-06 | 15:15 | 11.73 | 8.07 |  |  | 62.9 | 40.9
2007-04-06 | 15:30 | 11.83 | 8.08 |  |  | 63 | 40.9
2007-04-06 | 15:45 | 11.88 | 8.07 |  |  | 63 | 40.9
2007-04-06 | 16:00 | 12 | 8.08 |  |  | 63 | 41
2007-04-06 | 16:15 | 12.07 | 8.08 |  |  | 63 | 41
2007-04-06 | 16:30 | 12.13 | 8.08 |  |  | 63.1 | 41
2007-04-06 | 16:45 | 12.17 | 8.07 |  |  | 63.1 | 41
2007-04-06 | 17:00 | 12.19 | 8.06 |  |  | 63.1 | 41
2007-04-06 | 17:15 | 12.23 | 8.05 |  |  | 63.2 | 41.1
2007-04-06 | 17:30 | 12.26 | 8.05 |  |  | 63.2 | 41.1
2007-04-06 | 17:45 | 12.26 | 8.04 |  |  | 63.2 | 41.1
2007-04-06 | 18:00 | 12.25 | 8.02 |  |  | 63.3 | 41.1
2007-04-06 | 18:15 | 12.22 | 7.99 |  |  | 63.3 | 41.2
2007-04-06 | 18:30 | 12.19 | 7.96 |  |  | 63.4 | 41.2
2007-04-06 | 18:45 | 12.17 | 7.93 |  |  | 63.4 | 41.2
2007-04-06 | 19:00 | 12.15 | 7.9 |  |  | 63.5 | 41.3
2007-04-06 | 19:15 | 12.12 | 7.87 |  |  | 63.6 | 41.3
2007-04-06 | 19:30 | 12.09 | 7.84 |  |  | 63.6 | 41.3
2007-04-06 | 19:45 | 12.06 | 7.81 |  |  | 63.7 | 41.4
2007-04-06 | 20:00 | 12.03 | 7.78 |  |  | 63.7 | 41.4
2007-04-06 | 20:15 | 12 | 7.76 |  |  | 63.7 | 41.4
2007-04-06 | 20:30 | 11.98 | 7.74 |  |  | 63.8 | 41.5
2007-04-06 | 20:45 | 11.95 | 7.73 |  |  | 63.8 | 41.5
2007-04-06 | 21:00 | 11.93 | 7.71 |  |  | 63.8 | 41.5
2007-04-06 | 21:15 | 11.91 | 7.69 |  |  | 63.9 | 41.5
2007-04-06 | 21:30 | 11.88 | 7.68 |  |  | 63.9 | 41.5
2007-04-06 | 21:45 | 11.87 | 7.67 |  |  | 64 | 41.6
2007-04-06 | 22:00 | 11.84 | 7.66 |  |  | 64 | 41.6
2007-04-06 | 22:15 | 11.82 | 7.65 |  |  | 64 | 41.6
2007-04-06 | 22:30 | 11.78 | 7.65 |  |  | 64.8 | 42.1
2007-04-06 | 22:45 | 11.76 | 7.64 |  |  | 64 | 41.6
2007-04-06 | 23:00 | 11.73 | 7.64 |  |  | 64 | 41.6
2007-04-06 | 23:15 | 11.7 | 7.63 |  |  | 63.9 | 41.6
2007-04-06 | 23:30 | 11.66 | 7.63 |  |  | 63.9 | 41.5
2007-04-06 | 23:45 | 11.63 | 7.62 |  |  | 63.9 | 41.5
2007-04-07 | 0:00 | 11.6 | 7.61 |  |  | 63.9 | 41.5
2007-04-07 | 0:15 | 11.57 | 7.6 |  |  | 63.9 | 41.5
2007-04-07 | 0:30 | 11.54 | 7.6 |  |  | 63.9 | 41.5
2007-04-07 | 0:45 | 11.51 | 7.6 |  |  | 63.9 | 41.5
2007-04-07 | 1:01 | 11.48 | 7.59 |  |  | 63.9 | 41.5
2007-04-07 | 1:16 | 11.45 | 7.59 |  |  | 63.9 | 41.5
2007-04-07 | 1:31 | 11.43 | 7.59 |  |  | 63.9 | 41.5
2007-04-07 | 1:46 | 11.4 | 7.58 |  |  | 63.8 | 41.5
2007-04-07 | 2:01 | 11.36 | 7.58 |  |  | 63.9 | 41.5
2007-04-07 | 2:16 | 11.33 | 7.58 |  |  | 63.9 | 41.5
2007-04-07 | 2:31 | 11.29 | 7.57 |  |  | 63.8 | 41.5
2007-04-07 | 2:46 | 11.26 | 7.57 |  |  | 63.8 | 41.5
2007-04-07 | 3:01 | 11.23 | 7.57 |  |  | 63.8 | 41.5
2007-04-07 | 3:16 | 11.2 | 7.57 |  |  | 63.8 | 41.5
2007-04-07 | 3:31 | 11.18 | 7.57 |  |  | 63.8 | 41.4
2007-04-07 | 3:46 | 11.16 | 7.56 |  |  | 63.7 | 41.4
2007-04-07 | 4:01 | 11.15 | 7.56 |  |  | 63.7 | 41.4
2007-04-07 | 4:16 | 11.13 | 7.56 |  |  | 63.7 | 41.4
2007-04-07 | 4:31 | 11.11 | 7.56 |  |  | 63.7 | 41.4
2007-04-07 | 4:46 | 11.1 | 7.55 |  |  | 63.7 | 41.4
2007-04-07 | 5:01 | 11.09 | 7.56 |  |  | 63.6 | 41.3
2007-04-07 | 5:16 | 11.07 | 7.56 |  |  | 63.6 | 41.3
2007-04-07 | 5:31 | 11.05 | 7.55 |  |  | 63.5 | 41.3
2007-04-07 | 5:46 | 11.04 | 7.55 |  |  | 63.5 | 41.3
2007-04-07 | 6:01 | 11.03 | 7.55 |  |  | 63.4 | 41.2
2007-04-07 | 6:16 | 11.01 | 7.55 |  |  | 63.4 | 41.2
2007-04-07 | 6:31 | 10.99 | 7.55 |  |  | 63.3 | 41.2
2007-04-07 | 6:46 | 10.98 | 7.55 |  |  | 63.3 | 41.2
2007-04-07 | 7:01 | 10.97 | 7.55 |  |  | 63.3 | 41.1
2007-04-07 | 7:16 | 10.96 | 7.55 |  |  | 63.3 | 41.1
2007-04-07 | 7:31 | 10.95 | 7.55 |  |  | 63.2 | 41.1
2007-04-07 | 7:46 | 10.94 | 7.56 |  |  | 63.2 | 41.1
2007-04-07 | 8:01 | 10.92 | 7.56 |  |  | 63.3 | 41.1
2007-04-07 | 8:16 | 10.9 | 7.56 |  |  | 63.2 | 41.1
2007-04-07 | 8:31 | 10.89 | 7.57 |  |  | 63.2 | 41.1
2007-04-07 | 8:46 | 10.9 | 7.59 |  |  | 63.2 | 41.1
2007-04-07 | 9:01 | 10.9 | 7.61 |  |  | 63.2 | 41.1
2007-04-07 | 9:16 | 10.91 | 7.63 |  |  | 63.2 | 41
2007-04-07 | 9:31 | 10.92 | 7.65 |  |  | 63.2 | 41
2007-04-07 | 9:46 | 10.94 | 7.66 |  |  | 63.1 | 41
2007-04-07 | 10:01 | 10.96 | 7.69 |  |  | 63.1 | 41
2007-04-07 | 10:16 | 11 | 7.71 |  |  | 63.1 | 41
2007-04-07 | 10:31 | 11.04 | 7.73 |  |  | 63.1 | 41
2007-04-07 | 10:46 | 11.07 | 7.75 |  |  | 63.1 | 41
2007-04-07 | 11:01 | 11.1 | 7.77 |  |  | 63.1 | 41
2007-04-07 | 11:16 | 11.11 | 7.77 |  |  | 63.1 | 41
2007-04-07 | 11:31 | 11.11 | 7.78 |  |  | 63.1 | 41
2007-04-07 | 11:46 | 11.08 | 7.77 |  |  | 63 | 41
2007-04-07 | 12:01 | 11.06 | 7.76 |  |  | 63 | 41
2007-04-07 | 12:16 | 11.04 | 7.76 |  |  | 63 | 40.9
2007-04-07 | 12:31 | 11.02 | 7.75 |  |  | 63 | 40.9
2007-04-07 | 12:46 | 11 | 7.75 |  |  | 62.9 | 40.9
2007-04-07 | 13:00 | 11 | 7.75 |  |  | 62.9 | 40.9
2007-04-07 | 13:15 | 11.02 | 7.77 |  |  | 62.9 | 40.9
2007-04-07 | 13:30 | 11.05 | 7.78 |  |  | 62.9 | 40.9
2007-04-07 | 13:45 | 11.08 | 7.79 |  |  | 62.9 | 40.9
2007-04-07 | 14:00 | 11.11 | 7.8 |  |  | 62.9 | 40.9
2007-04-07 | 14:15 | 11.12 | 7.8 |  |  | 62.9 | 40.9
2007-04-07 | 14:30 | 11.13 | 7.8 |  |  | 62.9 | 40.9
2007-04-07 | 14:45 | 11.14 | 7.81 |  |  | 62.8 | 40.8
2007-04-07 | 15:00 | 11.16 | 7.82 |  |  | 62.8 | 40.8
2007-04-07 | 15:15 | 11.17 | 7.82 |  |  | 62.8 | 40.8
2007-04-07 | 15:30 | 11.2 | 7.84 |  |  | 62.7 | 40.8
2007-04-07 | 15:45 | 11.24 | 7.85 |  |  | 62.7 | 40.8
2007-04-07 | 16:00 | 11.33 | 7.87 |  |  | 62.7 | 40.7
2007-04-07 | 16:15 | 11.42 | 7.9 |  |  | 62.6 | 40.7
2007-04-07 | 16:30 | 11.46 | 7.91 |  |  | 62.6 | 40.7
2007-04-07 | 16:45 | 11.51 | 7.92 |  |  | 62.7 | 40.7
2007-04-07 | 17:00 | 11.54 | 7.94 |  |  | 62.7 | 40.7
2007-04-07 | 17:15 | 11.55 | 7.93 |  |  | 62.6 | 40.7
2007-04-07 | 17:30 | 11.59 | 7.94 |  |  | 62.7 | 40.7
2007-04-07 | 17:45 | 11.55 | 7.93 |  |  | 62.7 | 40.8
2007-04-07 | 18:00 | 11.56 | 7.94 |  |  | 62.7 | 40.8
2007-04-07 | 18:15 | 11.53 | 7.93 |  |  | 62.7 | 40.8
2007-04-07 | 18:30 | 11.51 | 7.92 |  |  | 62.7 | 40.8
2007-04-07 | 18:45 | 11.47 | 7.9 |  |  | 62.7 | 40.8
2007-04-07 | 19:00 | 11.43 | 7.88 |  |  | 62.8 | 40.8
2007-04-07 | 19:15 | 11.39 | 7.86 |  |  | 62.8 | 40.8
2007-04-07 | 19:30 | 11.34 | 7.84 |  |  | 62.8 | 40.8
2007-04-07 | 19:45 | 11.28 | 7.81 |  |  | 62.8 | 40.8
2007-04-07 | 20:00 | 11.22 | 7.78 |  |  | 62.8 | 40.8
2007-04-07 | 20:15 | 11.16 | 7.76 |  |  | 62.9 | 40.9
2007-04-07 | 20:30 | 11.11 | 7.74 |  |  | 62.9 | 40.9
2007-04-07 | 20:45 | 11.06 | 7.72 |  |  | 62.8 | 40.8
2007-04-07 | 21:00 | 11.01 | 7.7 |  |  | 62.9 | 40.9
2007-04-07 | 21:15 | 10.96 | 7.69 |  |  | 62.8 | 40.8
2007-04-07 | 21:30 | 10.9 | 7.68 |  |  | 62.8 | 40.8
2007-04-07 | 21:45 | 10.85 | 7.67 |  |  | 62.8 | 40.8
2007-04-07 | 22:00 | 10.8 | 7.66 |  |  | 62.8 | 40.8
2007-04-07 | 22:15 | 10.75 | 7.65 |  |  | 62.8 | 40.8
2007-04-07 | 22:30 | 10.7 | 7.65 |  |  | 62.8 | 40.8
2007-04-07 | 22:45 | 10.65 | 7.64 |  |  | 62.7 | 40.8
2007-04-07 | 23:00 | 10.6 | 7.64 |  |  | 62.7 | 40.8
2007-04-07 | 23:15 | 10.55 | 7.63 |  |  | 63.2 | 41.1
2007-04-07 | 23:30 | 10.51 | 7.63 |  |  | 62.6 | 40.7
2007-04-07 | 23:45 | 10.46 | 7.62 |  |  | 62.6 | 40.7
2007-04-08 | 0:00 | 10.42 | 7.62 |  |  | 62.6 | 40.7
2007-04-08 | 0:15 | 10.38 | 7.61 |  |  | 62.5 | 40.6
2007-04-08 | 0:30 | 10.33 | 7.61 |  |  | 62.5 | 40.6
2007-04-08 | 0:45 | 10.29 | 7.61 |  |  | 62.5 | 40.6
2007-04-08 | 1:01 | 10.24 | 7.61 |  |  | 62.4 | 40.6
2007-04-08 | 1:16 | 10.19 | 7.6 |  |  | 62.4 | 40.6
2007-04-08 | 1:31 | 10.13 | 7.6 |  |  | 62.4 | 40.6
2007-04-08 | 1:46 | 10.07 | 7.59 |  |  | 62.4 | 40.5
2007-04-08 | 2:01 | 10.01 | 7.59 |  |  | 62.4 | 40.5
2007-04-08 | 2:16 | 9.94 | 7.59 |  |  | 62.4 | 40.5
2007-04-08 | 2:31 | 9.88 | 7.58 |  |  | 62.3 | 40.5
2007-04-08 | 2:46 | 9.81 | 7.58 |  |  | 62.3 | 40.5
2007-04-08 | 3:01 | 9.75 | 7.58 |  |  | 62.3 | 40.5
2007-04-08 | 3:16 | 9.69 | 7.58 |  |  | 62.2 | 40.5
2007-04-08 | 3:31 | 9.63 | 7.58 |  |  | 62.2 | 40.4
2007-04-08 | 3:46 | 9.56 | 7.58 |  |  | 62.2 | 40.4
2007-04-08 | 4:01 | 9.5 | 7.58 |  |  | 62.2 | 40.4
2007-04-08 | 4:16 | 9.44 | 7.57 |  |  | 62.2 | 40.4
2007-04-08 | 4:31 | 9.38 | 7.58 |  |  | 62.2 | 40.4
2007-04-08 | 4:46 | 9.32 | 7.57 |  |  | 62.2 | 40.4
2007-04-08 | 5:01 | 9.26 | 7.57 |  |  | 62.1 | 40.4
2007-04-08 | 5:16 | 9.2 | 7.57 |  |  | 62.1 | 40.4
2007-04-08 | 5:31 | 9.14 | 7.57 |  |  | 62.1 | 40.3
2007-04-08 | 5:46 | 9.09 | 7.57 |  |  | 62 | 40.3
2007-04-08 | 6:01 | 9.03 | 7.57 |  |  | 62 | 40.3
2007-04-08 | 6:16 | 8.98 | 7.57 |  |  | 61.9 | 40.3
2007-04-08 | 6:31 | 8.93 | 7.57 |  |  | 61.9 | 40.2
2007-04-08 | 6:46 | 8.88 | 7.57 |  |  | 61.8 | 40.2
2007-04-08 | 7:01 | 8.83 | 7.57 |  |  | 61.8 | 40.2
2007-04-08 | 7:16 | 8.78 | 7.57 |  |  | 61.8 | 40.2
2007-04-08 | 7:31 | 8.76 | 7.58 |  |  | 61.7 | 40.1
2007-04-08 | 7:46 | 8.74 | 7.59 |  |  | 61.7 | 40.1
2007-04-08 | 8:01 | 8.73 | 7.61 |  |  | 61.7 | 40.1
2007-04-08 | 8:16 | 8.73 | 7.62 |  |  | 61.6 | 40.1
2007-04-08 | 8:31 | 8.75 | 7.64 |  |  | 61.6 | 40
2007-04-08 | 8:46 | 8.78 | 7.66 |  |  | 61.6 | 40
2007-04-08 | 9:01 | 8.83 | 7.68 |  |  | 61.6 | 40
2007-04-08 | 9:16 | 8.9 | 7.7 |  |  | 61.6 | 40
2007-04-08 | 9:31 | 8.97 | 7.71 |  |  | 61.6 | 40
2007-04-08 | 9:46 | 9.05 | 7.74 |  |  | 61.6 | 40
2007-04-08 | 10:01 | 9.14 | 7.75 |  |  | 61.6 | 40
2007-04-08 | 10:16 | 9.22 | 7.77 |  |  | 61.6 | 40
2007-04-08 | 10:31 | 9.32 | 7.79 |  |  | 61.6 | 40
2007-04-08 | 10:46 | 9.43 | 7.8 |  |  | 61.6 | 40
2007-04-08 | 11:01 | 9.53 | 7.83 |  |  | 61.6 | 40
2007-04-08 | 11:16 | 9.64 | 7.84 |  |  | 61.6 | 40
2007-04-08 | 11:31 | 9.76 | 7.86 |  |  | 61.6 | 40
2007-04-08 | 11:46 | 9.87 | 7.88 |  |  | 61.6 | 40
2007-04-08 | 12:01 | 10 | 7.89 |  |  | 61.6 | 40
2007-04-08 | 12:16 | 10.12 | 7.91 |  |  | 61.6 | 40
2007-04-08 | 12:31 | 10.26 | 7.92 |  |  | 61.6 | 40
2007-04-08 | 12:46 | 10.38 | 7.94 |  |  | 61.5 | 40
2007-04-08 | 13:00 | 10.49 | 7.97 |  |  | 61.6 | 40
2007-04-08 | 13:15 | 10.62 | 7.98 |  |  | 61.5 | 40
2007-04-08 | 13:30 | 10.75 | 7.99 |  |  | 61.5 | 40
2007-04-08 | 13:45 | 10.86 | 8.01 |  |  | 61.5 | 40
2007-04-08 | 14:00 | 10.96 | 8.03 |  |  | 61.4 | 39.9
2007-04-08 | 14:15 | 11.04 | 8.03 |  |  | 61.4 | 39.9
2007-04-08 | 14:30 | 11.13 | 8.05 |  |  | 61.4 | 39.9
2007-04-08 | 14:45 | 11.19 | 8.05 |  |  | 61.4 | 39.9
2007-04-08 | 15:00 | 11.23 | 8.06 |  |  | 61.4 | 39.9
2007-04-08 | 15:15 | 11.28 | 8.07 |  |  | 61.3 | 39.9
2007-04-08 | 15:30 | 11.37 | 8.09 |  |  | 61.3 | 39.9
2007-04-08 | 15:45 | 11.47 | 8.1 |  |  | 61.3 | 39.8
2007-04-08 | 16:00 | 11.56 | 8.11 |  |  | 61.3 | 39.8
2007-04-08 | 16:15 | 11.62 | 8.11 |  |  | 61.3 | 39.8
2007-04-08 | 16:30 | 11.65 | 8.11 |  |  | 61.3 | 39.8
2007-04-08 | 16:45 | 11.63 | 8.08 |  |  | 62.1 | 40.4
2007-04-08 | 17:00 | 11.62 | 8.06 |  |  | 62 | 40.3
2007-04-08 | 17:15 | 11.58 | 8.03 |  |  | 62.1 | 40.4
2007-04-08 | 17:30 | 11.54 | 8 |  |  | 61.5 | 40
2007-04-08 | 17:45 | 11.51 | 7.96 |  |  | 61.5 | 40
2007-04-08 | 18:00 | 11.48 | 7.93 |  |  | 61.6 | 40
2007-04-08 | 18:15 | 11.45 | 7.91 |  |  | 61.7 | 40.1
2007-04-08 | 18:30 | 11.41 | 7.88 |  |  | 61.7 | 40.1
2007-04-08 | 18:45 | 11.38 | 7.87 |  |  | 61.8 | 40.2
2007-04-08 | 19:00 | 11.34 | 7.84 |  |  | 61.9 | 40.2
2007-04-08 | 19:15 | 11.3 | 7.82 |  |  | 61.9 | 40.2
2007-04-08 | 19:30 | 11.27 | 7.79 |  |  | 62 | 40.3
2007-04-08 | 19:45 | 11.23 | 7.77 |  |  | 62 | 40.3
2007-04-08 | 20:00 | 11.2 | 7.74 |  |  | 62.1 | 40.3
2007-04-08 | 20:15 | 11.16 | 7.73 |  |  | 62.1 | 40.4
2007-04-08 | 20:30 | 11.12 | 7.71 |  |  | 62.1 | 40.4
2007-04-08 | 20:45 | 11.09 | 7.7 |  |  | 62.1 | 40.4
2007-04-08 | 21:00 | 11.06 | 7.69 |  |  | 62.2 | 40.4
2007-04-08 | 21:15 | 11.03 | 7.68 |  |  | 62.2 | 40.4
2007-04-08 | 21:30 | 11.01 | 7.67 |  |  | 62.2 | 40.4
2007-04-08 | 21:45 | 10.99 | 7.66 |  |  | 62.2 | 40.4
2007-04-08 | 22:00 | 10.97 | 7.65 |  |  | 62.2 | 40.4
2007-04-08 | 22:15 | 10.94 | 7.65 |  |  | 62.2 | 40.4
2007-04-08 | 22:30 | 10.91 | 7.64 |  |  | 62.2 | 40.4
2007-04-08 | 22:45 | 10.88 | 7.63 |  |  | 62.2 | 40.4
2007-04-08 | 23:00 | 10.84 | 7.63 |  |  | 62.2 | 40.4
2007-04-08 | 23:15 | 10.81 | 7.62 |  |  | 62.2 | 40.4
2007-04-08 | 23:30 | 10.77 | 7.62 |  |  | 62.2 | 40.4
2007-04-08 | 23:45 | 10.72 | 7.61 |  |  | 62.2 | 40.4
2007-04-09 | 0:00 | 10.66 | 7.6 |  |  | 62.2 | 40.4
2007-04-09 | 0:15 | 10.6 | 7.6 |  |  | 62.2 | 40.4
2007-04-09 | 0:30 | 10.55 | 7.6 |  |  | 62.1 | 40.4
2007-04-09 | 0:45 | 10.51 | 7.6 |  |  | 62.1 | 40.4
2007-04-09 | 1:01 | 10.47 | 7.59 |  |  | 62.1 | 40.4
2007-04-09 | 1:16 | 10.44 | 7.59 |  |  | 62.1 | 40.4
2007-04-09 | 1:31 | 10.4 | 7.58 |  |  | 62.1 | 40.4
2007-04-09 | 1:46 | 10.37 | 7.58 |  |  | 62.1 | 40.3
2007-04-09 | 2:01 | 10.35 | 7.58 |  |  | 62.1 | 40.3
2007-04-09 | 2:16 | 10.33 | 7.58 |  |  | 62 | 40.3
2007-04-09 | 2:31 | 10.31 | 7.58 |  |  | 62 | 40.3
2007-04-09 | 2:46 | 10.28 | 7.58 |  |  | 62 | 40.3
2007-04-09 | 3:01 | 10.24 | 7.58 |  |  | 62 | 40.3
2007-04-09 | 3:16 | 10.19 | 7.57 |  |  | 62 | 40.3
2007-04-09 | 3:31 | 10.13 | 7.57 |  |  | 62 | 40.3
2007-04-09 | 3:46 | 10.08 | 7.57 |  |  | 62 | 40.3
2007-04-09 | 4:01 | 10.03 | 7.57 |  |  | 61.9 | 40.3
2007-04-09 | 4:16 | 9.98 | 7.57 |  |  | 61.9 | 40.2
2007-04-09 | 4:31 | 9.93 | 7.57 |  |  | 61.9 | 40.2
2007-04-09 | 4:46 | 9.9 | 7.57 |  |  | 61.9 | 40.2
2007-04-09 | 5:01 | 9.88 | 7.57 |  |  | 61.9 | 40.2
2007-04-09 | 5:16 | 9.86 | 7.57 |  |  | 61.8 | 40.2
2007-04-09 | 5:31 | 9.85 | 7.57 |  |  | 61.8 | 40.2
2007-04-09 | 5:46 | 9.82 | 7.57 |  |  | 61.8 | 40.2
2007-04-09 | 6:01 | 9.8 | 7.56 |  |  | 61.8 | 40.2
2007-04-09 | 6:16 | 9.76 | 7.56 |  |  | 61.8 | 40.1
2007-04-09 | 6:31 | 9.72 | 7.57 |  |  | 61.8 | 40.1
2007-04-09 | 6:46 | 9.65 | 7.57 |  |  | 61.8 | 40.1
2007-04-09 | 7:01 | 9.58 | 7.57 |  |  | 61.8 | 40.1
2007-04-09 | 7:16 | 9.55 | 7.58 |  |  | 61.8 | 40.1
2007-04-09 | 7:31 | 9.49 | 7.58 |  |  | 61.8 | 40.1
2007-04-09 | 7:46 | 9.42 | 7.58 |  |  | 61.7 | 40.1
2007-04-09 | 8:01 | 9.36 | 7.59 |  |  | 61.8 | 40.1
2007-04-09 | 8:16 | 9.33 | 7.61 |  |  | 61.7 | 40.1
2007-04-09 | 8:31 | 9.32 | 7.63 |  |  | 61.7 | 40.1
2007-04-09 | 8:46 | 9.29 | 7.64 |  |  | 61.7 | 40.1
2007-04-09 | 9:01 | 9.29 | 7.66 |  |  | 61.7 | 40.1
2007-04-09 | 9:16 | 9.26 | 7.67 |  |  | 61.7 | 40.1
2007-04-09 | 9:31 | 9.27 | 7.69 |  |  | 61.6 | 40.1
2007-04-09 | 9:46 | 9.33 | 7.71 |  |  | 61.6 | 40.1
2007-04-09 | 10:01 | 9.43 | 7.74 |  |  | 61.6 | 40
2007-04-09 | 10:16 | 9.48 | 7.76 |  |  | 61.6 | 40
2007-04-09 | 10:31 | 9.52 | 7.78 |  |  | 61.6 | 40
2007-04-09 | 10:46 | 9.6 | 7.81 |  |  | 61.5 | 40
2007-04-09 | 11:01 | 9.57 | 7.81 |  |  | 61.5 | 40
2007-04-09 | 11:16 | 9.59 | 7.83 |  |  | 61.5 | 40
2007-04-09 | 11:31 | 9.63 | 7.84 |  |  | 61.5 | 40
2007-04-09 | 11:46 | 9.74 | 7.87 |  |  | 61.5 | 40
2007-04-09 | 12:01 | 9.86 | 7.89 |  |  | 61.5 | 39.9
2007-04-09 | 12:16 | 9.97 | 7.92 |  |  | 61.5 | 40
2007-04-09 | 12:31 | 10.08 | 7.93 |  |  | 61.5 | 40
2007-04-09 | 12:46 | 10.18 | 7.95 |  |  | 61.5 | 39.9
2007-04-09 | 13:00 | 10.29 | 7.96 |  |  | 61.5 | 39.9
2007-04-09 | 13:15 | 10.4 | 7.98 |  |  | 61.5 | 40
2007-04-09 | 13:30 | 10.49 | 8 |  |  | 61.4 | 39.9
2007-04-09 | 13:45 | 10.51 | 8 |  |  | 61.5 | 40
2007-04-09 | 14:00 | 10.54 | 8 |  |  | 61.4 | 39.9
2007-04-09 | 14:15 | 10.48 | 7.99 |  |  | 61.4 | 39.9
2007-04-09 | 14:30 | 10.52 | 8 |  |  | 61.5 | 39.9
2007-04-09 | 14:45 | 10.63 | 8.02 |  |  | 61.4 | 39.9
2007-04-09 | 15:00 | 10.75 | 8.03 |  |  | 61.4 | 39.9
2007-04-09 | 15:15 | 10.82 | 8.05 |  |  | 61.4 | 39.9
2007-04-09 | 15:30 | 10.87 | 8.05 |  |  | 61.4 | 39.9
2007-04-09 | 15:45 | 10.87 | 8.04 |  |  | 61.4 | 39.9
2007-04-09 | 16:00 | 10.85 | 8.03 |  |  | 61.5 | 40
2007-04-09 | 16:15 | 10.85 | 8.02 |  |  | 61.5 | 40
2007-04-09 | 16:30 | 10.84 | 8.03 |  |  | 60.1 | 39.1