## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek OWRD gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 #AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-03 |  | 2007-04-03 |  | A
Time | Computer set to atomic time |  | 15:13 |  | 15:13 |  | A
Temp | VWR NIST traceable surface water |  | 13.1 |  | 13.3 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 11.5 |  | A
SpC | 148 µS/mL |  | 149 |  | 149 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2007-04-26 16:42 | 8.26 | - | - | - | 29.5 | 19.2
| 2007-04-26 16:31 | 8.11 | 7.56 | 11.16 | 94.5 | 30 | 19.5
Grade |  | A | OR | OR |  | A
Audit 2 | 2007-05-13 13:44 | 6.42 | - | - | - | 24.6 | 16
| 2007-05-13 13:46 | 6.35 | 7.49 | 12.27 | 99.4 | 25.1 | 16.3
Grade |  | A | OR | OR |  | A
Audit 3 | 2007-06-07 14:42 | 5.93 | 0 | - | - | 23.8 | 15.4
| 2007-06-07 14:46 | 5.9 | 7.5 | 12.48 | 100 | 18.1 | 11.8
Grade |  | A | OR | OR |  | Fail
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
OR = over ride |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek OWRD gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-25 |  | 2007-04-25 |  | A
Time | Computer set to atomic time |  | 16:29 |  | 16:29 |  | A
Temp | VWR NIST traceable surface water |  | 15 |  | 15.1 |  | A
DO | Winkler titration of surface water |  | 10.2 |  | 10.9 |  | B
SpC | 148 µS/mL |  | 147 |  | 142 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2007-04-26 16:42 | 8.26 | - | - | - | 29.5 | 19.2
-
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-04-29 |  | 2007-04-29 |  | A
Time | Computer set to atomic time |  | 10:57 |  | 10:57 |  | A
Temp | VWR NIST traceable surface water |  | 13.8 |  | 13.9 |  | A
DO | Winkler titration of surface water |  | 10 |  | 10.8 |  | B
SpC | 148 µS/mL |  | 147 |  | 142 |  | A
Ph | pH 7 standard solution |  | 7.03 |  | 7.1 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek OWRD gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-05-13 |  | 2007-05-13 |  | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water |  | 15.3 |  | 15.1 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.3 |  | A
SpC | 148 µS/mL |  | 147 |  | 134 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2007-05-13 13:44 | 6.42 | - | - | - | 24.6 | 16
-
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-05-16 |  | 2007-05-16 |  | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water |  | 10 |  | 10.6 |  | B
SpC | 148 µS/mL |  | 147 |  | 140 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek OWRD gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-05-16 |  | 2007-05-16 |  | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water |  | 10 |  | 10.6 |  | B
SpC | 148 µS/mL |  | 147 |  | 140 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2007-06-07 14:42 | 5.93 |  | - | - | 23.8 | 15.4
-
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2007-06-08 |  | 2007-06-08 |  | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water |  | 15 |  | 14.8 |  | A
DO | Winkler titration of surface water |  | 11.2 |  | 11.6 |  | B
SpC | 148 µS/mL |  | 147 |  | 135 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.3 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO % | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2007-04-26 | 16:30 | 8.11 | 7.56 | 11.16 | 94.5 | 30 | 19.5
2007-04-26 | 16:45 | 8.23 | 7.57 | 11.18 | 94.9 | 30 | 19.5
2007-04-26 | 17:00 | 8.35 | 7.57 | 11.16 | 95 | 30 | 19.5
2007-04-26 | 17:15 | 8.44 | 7.58 | 11.14 | 95.1 | 30 | 19.5
2007-04-26 | 17:30 | 8.5 | 7.58 | 11.12 | 95 | 30 | 19.5
2007-04-26 | 17:45 | 8.54 | 7.58 | 11.12 | 95.1 | 30 | 19.5
2007-04-26 | 18:00 | 8.59 | 7.59 | 11.06 | 94.7 | 30 | 19.5
2007-04-26 | 18:15 | 8.65 | 7.59 | 11.07 | 94.9 | 30 | 19.5
2007-04-26 | 18:30 | 8.69 | 7.59 | 11.05 | 94.9 | 30 | 19.5
2007-04-26 | 18:45 | 8.69 | 7.59 | 11.05 | 94.9 | 30 | 19.5
2007-04-26 | 19:00 | 8.66 | 7.59 | 11.06 | 94.9 | 30 | 19.5
2007-04-26 | 19:15 | 8.61 | 7.58 | 11.07 | 94.9 | 30 | 19.5
2007-04-26 | 19:30 | 8.52 | 7.58 | 11.09 | 94.9 | 30 | 19.5
2007-04-26 | 19:45 | 8.4 | 7.58 | 11.12 | 94.8 | 30 | 19.5
2007-04-26 | 20:00 | 8.26 | 7.57 | 11.16 | 94.8 | 30 | 19.5
2007-04-26 | 20:15 | 8.1 | 7.56 | 11.19 | 94.7 | 29.9 | 19.5
2007-04-26 | 20:30 | 7.93 | 7.56 | 11.23 | 94.7 | 29.9 | 19.4
2007-04-26 | 20:45 | 7.77 | 7.55 | 11.28 | 94.7 | 29.9 | 19.5
2007-04-26 | 21:00 | 7.61 | 7.54 | 11.31 | 94.6 | 29.9 | 19.4
2007-04-26 | 21:15 | 7.45 | 7.54 | 11.35 | 94.5 | 29.8 | 19.4
2007-04-26 | 21:30 | 7.3 | 7.53 | 11.41 | 94.7 | 29.8 | 19.4
2007-04-26 | 21:45 | 7.14 | 7.53 | 11.45 | 94.7 | 29.8 | 19.4
2007-04-26 | 22:00 | 6.99 | 7.52 | 11.47 | 94.5 | 29.8 | 19.3
2007-04-26 | 22:15 | 6.84 | 7.52 | 11.54 | 94.7 | 29.7 | 19.3
2007-04-26 | 22:30 | 6.7 | 7.51 | 11.57 | 94.6 | 29.7 | 19.3
2007-04-26 | 22:45 | 6.57 | 7.51 | 11.6 | 94.5 | 29.7 | 19.3
2007-04-26 | 23:00 | 6.43 | 7.51 | 11.65 | 94.6 | 29.6 | 19.2
2007-04-26 | 23:15 | 6.29 | 7.51 | 11.69 | 94.6 | 29.6 | 19.2
2007-04-26 | 23:30 | 6.16 | 7.5 | 11.73 | 94.6 | 29.6 | 19.2
2007-04-26 | 23:45 | 6.03 | 7.5 | 11.77 | 94.6 | 29.5 | 19.2
2007-04-27 | 0:00 | 5.9 | 7.5 | 11.82 | 94.7 | 29.5 | 19.2
2007-04-27 | 0:15 | 5.76 | 7.49 | 11.84 | 94.6 | 29.4 | 19.1
2007-04-27 | 0:30 | 5.62 | 7.49 | 11.9 | 94.6 | 29.4 | 19.1
2007-04-27 | 0:45 | 5.48 | 7.48 | 11.94 | 94.7 | 29.4 | 19.1
2007-04-27 | 1:01 | 5.35 | 7.49 | 11.98 | 94.6 | 29.4 | 19.1
2007-04-27 | 1:16 | 5.22 | 7.48 | 12.02 | 94.7 | 29.3 | 19.1
2007-04-27 | 1:31 | 5.09 | 7.48 | 12.06 | 94.7 | 29.4 | 19.1
2007-04-27 | 1:46 | 4.96 | 7.47 | 12.1 | 94.7 | 29.3 | 19.1
2007-04-27 | 2:01 | 4.84 | 7.48 | 12.14 | 94.6 | 29.4 | 19.1
2007-04-27 | 2:16 | 4.71 | 7.47 | 12.17 | 94.6 | 29.3 | 19.1
2007-04-27 | 2:31 | 4.6 | 7.47 | 12.21 | 94.6 | 29.3 | 19.1
2007-04-27 | 2:46 | 4.48 | 7.48 | 12.24 | 94.6 | 29.3 | 19.1
2007-04-27 | 3:01 | 4.37 | 7.47 | 12.29 | 94.7 | 29.3 | 19.1
2007-04-27 | 3:16 | 4.25 | 7.47 | 12.31 | 94.6 | 29.4 | 19.1
2007-04-27 | 3:31 | 4.15 | 7.47 | 12.34 | 94.5 | 29.4 | 19.1
2007-04-27 | 3:46 | 4.05 | 7.47 | 12.38 | 94.6 | 29.4 | 19.1
2007-04-27 | 4:01 | 3.95 | 7.47 | 12.41 | 94.5 | 29.4 | 19.1
2007-04-27 | 4:16 | 3.86 | 7.47 | 12.44 | 94.6 | 29.4 | 19.1
2007-04-27 | 4:31 | 3.78 | 7.46 | 12.49 | 94.7 | 29.4 | 19.1
2007-04-27 | 4:46 | 3.69 | 7.47 | 12.5 | 94.6 | 29.4 | 19.1
2007-04-27 | 5:01 | 3.61 | 7.46 | 12.53 | 94.6 | 29.4 | 19.1
2007-04-27 | 5:16 | 3.53 | 7.46 | 12.55 | 94.6 | 29.4 | 19.1
2007-04-27 | 5:31 | 3.46 | 7.46 | 12.61 | 94.8 | 29.5 | 19.2
2007-04-27 | 5:46 | 3.39 | 7.47 | 12.6 | 94.6 | 29.5 | 19.2
2007-04-27 | 6:01 | 3.33 | 7.46 | 12.63 | 94.7 | 29.5 | 19.2
2007-04-27 | 6:16 | 3.28 | 7.46 | 12.65 | 94.6 | 29.5 | 19.2
2007-04-27 | 6:31 | 3.23 | 7.46 | 12.68 | 94.8 | 29.5 | 19.2
2007-04-27 | 6:46 | 3.18 | 7.46 | 12.68 | 94.7 | 29.5 | 19.2
2007-04-27 | 7:01 | 3.14 | 7.46 | 12.7 | 94.7 | 29.5 | 19.2
2007-04-27 | 7:16 | 3.11 | 7.46 | 12.72 | 94.8 | 29.5 | 19.2
2007-04-27 | 7:31 | 3.09 | 7.47 | 12.74 | 94.9 | 29.6 | 19.2
2007-04-27 | 7:46 | 3.1 | 7.48 | 12.71 | 94.6 | 29.6 | 19.2
2007-04-27 | 8:01 | 3.11 | 7.48 | 12.73 | 94.8 | 29.6 | 19.2
2007-04-27 | 8:16 | 3.13 | 7.48 | 12.71 | 94.7 | 29.6 | 19.2
2007-04-27 | 8:31 | 3.16 | 7.49 | 12.73 | 94.9 | 29.6 | 19.2
2007-04-27 | 8:46 | 3.19 | 7.5 | 12.7 | 94.8 | 29.6 | 19.2
2007-04-27 | 9:01 | 3.24 | 7.51 | 12.7 | 94.9 | 29.6 | 19.3
2007-04-27 | 9:16 | 3.31 | 7.51 | 12.67 | 94.9 | 29.6 | 19.2
2007-04-27 | 9:31 | 3.39 | 7.51 | 12.65 | 95 | 29.6 | 19.3
2007-04-27 | 9:46 | 3.49 | 7.52 | 12.63 | 95 | 29.6 | 19.3
2007-04-27 | 10:01 | 3.6 | 7.52 | 12.6 | 95.1 | 29.6 | 19.3
2007-04-27 | 10:16 | 3.73 | 7.52 | 12.57 | 95.2 | 29.6 | 19.3
2007-04-27 | 10:31 | 3.9 | 7.53 | 12.52 | 95.3 | 29.7 | 19.3
2007-04-27 | 10:46 | 4.07 | 7.53 | 12.46 | 95.2 | 29.6 | 19.3
2007-04-27 | 11:01 | 4.26 | 7.53 | 12.4 | 95.2 | 29.6 | 19.3
2007-04-27 | 11:16 | 4.45 | 7.53 | 12.34 | 95.2 | 29.6 | 19.3
2007-04-27 | 11:31 | 4.66 | 7.53 | 12.28 | 95.3 | 29.6 | 19.3
2007-04-27 | 11:46 | 4.89 | 7.54 | 12.2 | 95.2 | 29.7 | 19.3
2007-04-27 | 12:01 | 5.13 | 7.55 | 12.15 | 95.4 | 29.6 | 19.3
2007-04-27 | 12:16 | 5.37 | 7.55 | 12.07 | 95.4 | 29.6 | 19.3
2007-04-27 | 12:31 | 5.61 | 7.55 | 12.01 | 95.5 | 29.7 | 19.3
2007-04-27 | 12:46 | 5.87 | 7.55 | 11.94 | 95.6 | 29.7 | 19.3
2007-04-27 | 13:00 | 6.12 | 7.56 | 11.8 | 95.1 | 29.7 | 19.3
2007-04-27 | 13:15 | 6.42 | 7.56 | 11.74 | 95.3 | 29.7 | 19.3
2007-04-27 | 13:30 | 6.72 | 7.56 | 11.69 | 95.6 | 29.7 | 19.3
2007-04-27 | 13:45 | 7.01 | 7.57 | 11.61 | 95.6 | 29.7 | 19.3
2007-04-27 | 14:00 | 7.32 | 7.57 | 11.52 | 95.7 | 29.7 | 19.3
2007-04-27 | 14:15 | 7.63 | 7.58 | 11.44 | 95.7 | 29.7 | 19.3
2007-04-27 | 14:30 | 7.82 | 7.58 | 11.37 | 95.6 | 29.7 | 19.3
2007-04-27 | 14:45 | 7.94 | 7.58 | 11.33 | 95.5 | 29.7 | 19.3
2007-04-27 | 15:00 | 8.1 | 7.59 | 11.27 | 95.4 | 29.7 | 19.3
2007-04-27 | 15:15 | 8.3 | 7.59 | 11.23 | 95.6 | 29.7 | 19.3
2007-04-27 | 15:30 | 8.57 | 7.59 | 11.16 | 95.6 | 29.6 | 19.3
2007-04-27 | 15:45 | 8.85 | 7.59 | 11.1 | 95.6 | 29.7 | 19.3
2007-04-27 | 16:00 | 9.04 | 7.6 | 11.04 | 95.6 | 29.7 | 19.3
2007-04-27 | 16:15 | 9.18 | 7.6 | 10.99 | 95.5 | 29.7 | 19.3
2007-04-27 | 16:30 | 9.32 | 7.6 | 10.95 | 95.5 | 29.7 | 19.3
2007-04-27 | 16:45 | 9.47 | 7.61 | 10.93 | 95.6 | 29.7 | 19.3
2007-04-27 | 17:00 | 9.49 | 7.6 | 10.9 | 95.4 | 29.7 | 19.3
2007-04-27 | 17:15 | 9.39 | 7.6 | 10.92 | 95.3 | 29.7 | 19.3
2007-04-27 | 17:30 | 9.28 | 7.59 | 10.95 | 95.3 | 29.6 | 19.3
2007-04-27 | 17:45 | 9.14 | 7.59 | 10.98 | 95.3 | 29.6 | 19.3
2007-04-27 | 18:00 | 9.06 | 7.59 | 11.01 | 95.4 | 29.6 | 19.3
2007-04-27 | 18:15 | 9.03 | 7.59 | 10.99 | 95.1 | 29.6 | 19.2
2007-04-27 | 18:30 | 9.06 | 7.59 | 10.98 | 95.1 | 29.6 | 19.2
2007-04-27 | 18:45 | 9.1 | 7.59 | 10.97 | 95.1 | 29.5 | 19.2
2007-04-27 | 19:00 | 9.17 | 7.59 | 10.93 | 95 | 29.5 | 19.2
2007-04-27 | 19:15 | 9.35 | 7.59 | 10.9 | 95 | 29.5 | 19.2
2007-04-27 | 19:30 | 9.58 | 7.59 | 10.83 | 95 | 29.4 | 19.1
2007-04-27 | 19:45 | 9.78 | 7.59 | 10.77 | 94.9 | 29.4 | 19.1
2007-04-27 | 20:00 | 9.86 | 7.58 | 10.76 | 95 | 29.3 | 19.1
2007-04-27 | 20:15 | 9.83 | 7.57 | 10.76 | 95 | 29.2 | 19
2007-04-27 | 20:30 | 9.72 | 7.57 | 10.78 | 94.9 | 29.2 | 19
2007-04-27 | 20:45 | 9.57 | 7.56 | 10.85 | 95.1 | 29.1 | 18.9
2007-04-27 | 21:00 | 9.41 | 7.57 | 10.88 | 95 | 29 | 18.9
2007-04-27 | 21:15 | 9.24 | 7.55 | 10.93 | 95.1 | 28.9 | 18.8
2007-04-27 | 21:30 | 9.07 | 7.54 | 10.96 | 95 | 28.8 | 18.7
2007-04-27 | 21:45 | 8.9 | 7.53 | 11 | 94.9 | 28.7 | 18.7
2007-04-27 | 22:00 | 8.72 | 7.52 | 11.06 | 95 | 28.7 | 18.6
2007-04-27 | 22:15 | 8.54 | 7.51 | 11.1 | 95 | 28.5 | 18.5
2007-04-27 | 22:30 | 8.36 | 7.5 | 11.15 | 94.9 | 28.4 | 18.5
2007-04-27 | 22:45 | 8.18 | 7.5 | 11.2 | 94.9 | 28.3 | 18.4
2007-04-27 | 23:00 | 8 | 7.5 | 11.24 | 94.9 | 28.2 | 18.3
2007-04-27 | 23:15 | 7.84 | 7.49 | 11.28 | 94.9 | 28.1 | 18.3
2007-04-27 | 23:30 | 7.68 | 7.49 | 11.33 | 94.9 | 28.1 | 18.3
2007-04-27 | 23:45 | 7.52 | 7.48 | 11.31 | 94.4 | 28 | 18.2
2007-04-28 | 0:00 | 7.37 | 7.48 | 11.39 | 94.7 | 27.9 | 18.2
2007-04-28 | 0:15 | 7.22 | 7.48 | 11.45 | 94.8 | 27.9 | 18.1
2007-04-28 | 0:30 | 7.05 | 7.48 | 11.5 | 94.8 | 27.8 | 18.1
2007-04-28 | 0:45 | 6.88 | 7.47 | 11.55 | 94.9 | 27.8 | 18.1
2007-04-28 | 1:01 | 6.71 | 7.47 | 11.58 | 94.8 | 27.8 | 18.1
2007-04-28 | 1:16 | 6.56 | 7.47 | 11.63 | 94.8 | 27.8 | 18.1
2007-04-28 | 1:31 | 6.4 | 7.47 | 11.67 | 94.7 | 27.7 | 18
2007-04-28 | 1:46 | 6.25 | 7.46 | 11.75 | 95 | 27.7 | 18
2007-04-28 | 2:01 | 6.1 | 7.46 | 11.78 | 94.9 | 27.7 | 18
2007-04-28 | 2:16 | 5.97 | 7.45 | 11.81 | 94.8 | 27.7 | 18
2007-04-28 | 2:31 | 5.84 | 7.46 | 11.85 | 94.8 | 27.7 | 18
2007-04-28 | 2:46 | 5.72 | 7.45 | 11.89 | 94.9 | 27.7 | 18
2007-04-28 | 3:01 | 5.6 | 7.45 | 11.91 | 94.7 | 27.8 | 18
2007-04-28 | 3:16 | 5.5 | 7.45 | 11.95 | 94.8 | 27.8 | 18.1
2007-04-28 | 3:31 | 5.4 | 7.45 | 11.98 | 94.8 | 27.8 | 18.1
2007-04-28 | 3:46 | 5.31 | 7.45 | 11.99 | 94.7 | 27.8 | 18.1
2007-04-28 | 4:01 | 5.21 | 7.45 | 12.04 | 94.8 | 27.8 | 18.1
2007-04-28 | 4:16 | 5.12 | 7.45 | 12.07 | 94.8 | 27.8 | 18.1
2007-04-28 | 4:31 | 5.04 | 7.45 | 12.02 | 94.2 | 27.8 | 18.1
2007-04-28 | 4:46 | 4.95 | 7.45 | 12.11 | 94.7 | 27.9 | 18.1
2007-04-28 | 5:01 | 4.86 | 7.45 | 12.14 | 94.7 | 27.9 | 18.1
2007-04-28 | 5:16 | 4.78 | 7.45 | 12.17 | 94.8 | 27.9 | 18.1
2007-04-28 | 5:31 | 4.7 | 7.45 | 12.21 | 94.9 | 27.9 | 18.1
2007-04-28 | 5:46 | 4.62 | 7.45 | 12.24 | 94.9 | 27.9 | 18.1
2007-04-28 | 6:01 | 4.54 | 7.45 | 12.25 | 94.8 | 27.9 | 18.2
2007-04-28 | 6:16 | 4.46 | 7.45 | 12.28 | 94.9 | 27.9 | 18.2
2007-04-28 | 6:31 | 4.38 | 7.45 | 12.31 | 94.9 | 28 | 18.2
2007-04-28 | 6:46 | 4.32 | 7.45 | 12.33 | 94.9 | 28 | 18.2
2007-04-28 | 7:01 | 4.26 | 7.45 | 12.36 | 94.9 | 28 | 18.2
2007-04-28 | 7:16 | 4.22 | 7.45 | 12.38 | 95 | 28 | 18.2
2007-04-28 | 7:31 | 4.18 | 7.46 | 12.39 | 95 | 28 | 18.2
2007-04-28 | 7:46 | 4.15 | 7.46 | 12.41 | 95 | 28 | 18.2
2007-04-28 | 8:01 | 4.13 | 7.46 | 12.41 | 95 | 28 | 18.2
2007-04-28 | 8:16 | 4.11 | 7.47 | 12.43 | 95.1 | 28.1 | 18.2
2007-04-28 | 8:31 | 4.1 | 7.47 | 12.42 | 95 | 28.1 | 18.3
2007-04-28 | 8:46 | 4.1 | 7.48 | 12.42 | 95 | 28.1 | 18.3
2007-04-28 | 9:01 | 4.12 | 7.49 | 12.44 | 95.2 | 28.1 | 18.3
2007-04-28 | 9:16 | 4.15 | 7.5 | 12.43 | 95.2 | 28.1 | 18.3
2007-04-28 | 9:31 | 4.17 | 7.49 | 12.42 | 95.1 | 28.2 | 18.3
2007-04-28 | 9:46 | 4.25 | 7.5 | 12.41 | 95.3 | 28.1 | 18.3
2007-04-28 | 10:01 | 4.35 | 7.5 | 12.4 | 95.4 | 28.2 | 18.3
2007-04-28 | 10:16 | 4.43 | 7.51 | 12.36 | 95.4 | 28.2 | 18.3
2007-04-28 | 10:31 | 4.54 | 7.51 | 12.33 | 95.4 | 28.2 | 18.3
2007-04-28 | 10:46 | 4.65 | 7.51 | 12.32 | 95.6 | 28.2 | 18.3
2007-04-28 | 11:01 | 4.79 | 7.51 | 12.25 | 95.4 | 28.2 | 18.3
2007-04-28 | 11:16 | 4.94 | 7.51 | 12.21 | 95.4 | 28.2 | 18.3
2007-04-28 | 11:31 | 5.1 | 7.52 | 12.18 | 95.6 | 28.2 | 18.3
2007-04-28 | 11:46 | 5.27 | 7.53 | 12.13 | 95.6 | 28.2 | 18.4
2007-04-28 | 12:01 | 5.48 | 7.52 | 12.07 | 95.7 | 28.3 | 18.4
2007-04-28 | 12:16 | 5.7 | 7.53 | 12 | 95.6 | 28.2 | 18.4
2007-04-28 | 12:31 | 5.93 | 7.53 | 11.96 | 95.9 | 28.3 | 18.4
2007-04-28 | 12:46 | 6.15 | 7.54 | 11.88 | 95.8 | 28.3 | 18.4
2007-04-28 | 13:00 | 6.39 | 7.54 | 11.81 | 95.8 | 28.3 | 18.4
2007-04-28 | 13:15 | 6.6 | 7.54 | 11.75 | 95.8 | 28.3 | 18.4
2007-04-28 | 13:30 | 6.88 | 7.54 | 11.67 | 95.8 | 28.3 | 18.4
2007-04-28 | 13:45 | 7.13 | 7.55 | 11.6 | 95.9 | 28.3 | 18.4
2007-04-28 | 14:00 | 7.42 | 7.55 | 11.54 | 96.1 | 28.3 | 18.4
2007-04-28 | 14:15 | 7.71 | 7.55 | 11.44 | 95.9 | 28.3 | 18.4
2007-04-28 | 14:30 | 7.98 | 7.56 | 11.39 | 96.1 | 28.4 | 18.4
2007-04-28 | 14:45 | 8.27 | 7.56 | 11.28 | 95.9 | 28.4 | 18.5
2007-04-28 | 15:00 | 8.57 | 7.56 | 11.19 | 95.8 | 28.4 | 18.5
2007-04-28 | 15:15 | 8.83 | 7.57 | 11.13 | 95.9 | 28.4 | 18.4
2007-04-28 | 15:30 | 9.06 | 7.58 | 11.06 | 95.9 | 28.4 | 18.4
2007-04-28 | 15:45 | 9.28 | 7.58 | 11 | 95.8 | 28.4 | 18.5
2007-04-28 | 16:00 | 9.47 | 7.58 | 10.95 | 95.8 | 28.4 | 18.5
2007-04-28 | 16:15 | 9.63 | 7.58 | 10.89 | 95.6 | 28.4 | 18.5
2007-04-28 | 16:30 | 9.75 | 7.59 | 10.86 | 95.7 | 28.4 | 18.5
2007-04-28 | 16:45 | 9.87 | 7.58 | 10.83 | 95.7 | 28.4 | 18.5
2007-04-28 | 17:00 | 9.97 | 7.58 | 10.81 | 95.7 | 28.5 | 18.5
2007-04-28 | 17:15 | 10.06 | 7.58 | 10.78 | 95.6 | 28.4 | 18.5
2007-04-28 | 17:30 | 10.13 | 7.58 | 10.77 | 95.7 | 28.4 | 18.5
2007-04-28 | 17:45 | 10.17 | 7.58 | 10.75 | 95.6 | 28.4 | 18.5
2007-04-28 | 18:00 | 10.17 | 7.58 | 10.73 | 95.5 | 28.4 | 18.5
2007-04-28 | 18:15 | 10.15 | 7.58 | 10.72 | 95.3 | 28.4 | 18.5
2007-04-28 | 18:30 | 10.1 | 7.58 | 10.74 | 95.4 | 28.4 | 18.5
2007-04-28 | 18:45 | 10.04 | 7.57 | 10.75 | 95.3 | 28.4 | 18.5
2007-04-28 | 19:00 | 9.95 | 7.56 | 10.77 | 95.3 | 28.4 | 18.4
2007-04-28 | 19:15 | 9.83 | 7.56 | 10.79 | 95.2 | 28.3 | 18.4
2007-04-28 | 19:30 | 9.67 | 7.56 | 10.83 | 95.2 | 28.3 | 18.4
2007-04-28 | 19:45 | 9.47 | 7.56 | 10.88 | 95.2 | 28.3 | 18.4
2007-04-28 | 20:00 | 9.25 | 7.55 | 10.94 | 95.2 | 28.2 | 18.4
2007-04-28 | 20:15 | 9.03 | 7.52 | 11 | 95.2 | 28.2 | 18.3
2007-04-28 | 20:30 | 8.8 | 7.52 | 11.05 | 95.1 | 28.1 | 18.3
2007-04-28 | 20:45 | 8.58 | 7.51 | 11.09 | 95 | 28.1 | 18.2
2007-04-28 | 21:00 | 8.39 | 7.5 | 11.16 | 95.1 | 28 | 18.2
2007-04-28 | 21:15 | 8.2 | 7.5 | 11.21 | 95.1 | 27.9 | 18.1
2007-04-28 | 21:30 | 8.03 | 7.49 | 11.25 | 95.1 | 27.8 | 18.1
2007-04-28 | 21:45 | 7.86 | 7.49 | 11.31 | 95.1 | 27.8 | 18
2007-04-28 | 22:00 | 7.69 | 7.48 | 11.35 | 95.1 | 27.7 | 18
2007-04-28 | 22:15 | 7.52 | 7.48 | 11.4 | 95.1 | 27.6 | 17.9
2007-04-28 | 22:30 | 7.36 | 7.47 | 11.44 | 95.1 | 27.5 | 17.9
2007-04-28 | 22:45 | 7.2 | 7.47 | 11.48 | 95.1 | 27.4 | 17.8
2007-04-28 | 23:00 | 7.05 | 7.46 | 11.53 | 95.1 | 27.3 | 17.8
2007-04-28 | 23:15 | 6.89 | 7.46 | 11.58 | 95.1 | 27.3 | 17.7
2007-04-28 | 23:30 | 6.73 | 7.45 | 11.62 | 95 | 27.2 | 17.7
2007-04-28 | 23:45 | 6.57 | 7.45 | 11.67 | 95.1 | 27.2 | 17.7
2007-04-29 | 0:00 | 6.41 | 7.45 | 11.71 | 95 | 27.1 | 17.6
2007-04-29 | 0:15 | 6.26 | 7.45 | 11.75 | 95 | 27.1 | 17.6
2007-04-29 | 0:30 | 6.11 | 7.44 | 11.8 | 95 | 27.1 | 17.6
2007-04-29 | 0:45 | 5.96 | 7.45 | 11.85 | 95.1 | 27 | 17.6
2007-04-29 | 1:01 | 5.82 | 7.44 | 11.88 | 95 | 27 | 17.6
2007-04-29 | 1:16 | 5.69 | 7.44 | 11.92 | 95 | 27 | 17.6
2007-04-29 | 1:31 | 5.57 | 7.43 | 11.96 | 95 | 27 | 17.5
2007-04-29 | 1:46 | 5.45 | 7.43 | 11.99 | 95 | 27 | 17.5
2007-04-29 | 2:01 | 5.33 | 7.43 | 12.02 | 95 | 27 | 17.5
2007-04-29 | 2:16 | 5.23 | 7.43 | 12.06 | 95 | 27 | 17.6
2007-04-29 | 2:31 | 5.12 | 7.43 | 12.1 | 95.1 | 27 | 17.6
2007-04-29 | 2:46 | 5.02 | 7.43 | 12.12 | 94.9 | 27 | 17.6
2007-04-29 | 3:01 | 4.93 | 7.43 | 12.15 | 94.9 | 27 | 17.6
2007-04-29 | 3:16 | 4.84 | 7.43 | 12.17 | 94.9 | 27 | 17.6
2007-04-29 | 3:31 | 4.75 | 7.43 | 12.2 | 94.9 | 27 | 17.6
2007-04-29 | 3:46 | 4.66 | 7.43 | 12.23 | 94.9 | 27.1 | 17.6
2007-04-29 | 4:01 | 4.58 | 7.43 | 12.26 | 94.9 | 27.1 | 17.6
2007-04-29 | 4:16 | 4.51 | 7.43 | 12.28 | 95 | 27.1 | 17.6
2007-04-29 | 4:31 | 4.43 | 7.43 | 12.31 | 95 | 27.1 | 17.6
2007-04-29 | 4:46 | 4.36 | 7.42 | 12.32 | 94.9 | 27.2 | 17.7
2007-04-29 | 5:01 | 4.3 | 7.42 | 12.35 | 94.9 | 27.2 | 17.7
2007-04-29 | 5:16 | 4.24 | 7.42 | 12.37 | 95 | 27.2 | 17.7
2007-04-29 | 5:31 | 4.2 | 7.43 | 12.4 | 95 | 27.2 | 17.7
2007-04-29 | 5:46 | 4.15 | 7.42 | 12.4 | 95 | 27.2 | 17.7
2007-04-29 | 6:01 | 4.11 | 7.42 | 12.42 | 95 | 27.3 | 17.7
2007-04-29 | 6:16 | 4.07 | 7.42 | 12.37 | 94.5 | 27.3 | 17.7
2007-04-29 | 6:31 | 4.04 | 7.42 | 12.42 | 94.8 | 27.3 | 17.8
2007-04-29 | 6:46 | 4.02 | 7.42 | 12.44 | 95 | 27.3 | 17.8
2007-04-29 | 7:01 | 4 | 7.42 | 12.46 | 95 | 27.4 | 17.8
2007-04-29 | 7:16 | 3.99 | 7.43 | 12.47 | 95.1 | 27.4 | 17.8
2007-04-29 | 7:31 | 3.99 | 7.43 | 12.48 | 95.2 | 27.4 | 17.8
2007-04-29 | 7:46 | 4 | 7.44 | 12.49 | 95.3 | 27.4 | 17.8
2007-04-29 | 8:01 | 4.01 | 7.45 | 12.48 | 95.2 | 27.5 | 17.9
2007-04-29 | 8:16 | 4.02 | 7.45 | 12.48 | 95.3 | 27.5 | 17.8
2007-04-29 | 8:31 | 4.05 | 7.46 | 12.48 | 95.3 | 27.5 | 17.9
2007-04-29 | 8:46 | 4.09 | 7.47 | 12.48 | 95.4 | 27.5 | 17.9
2007-04-29 | 9:01 | 4.14 | 7.48 | 12.46 | 95.4 | 27.5 | 17.9
2007-04-29 | 9:16 | 4.19 | 7.48 | 12.46 | 95.5 | 27.5 | 17.9
2007-04-29 | 9:31 | 4.24 | 7.48 | 12.43 | 95.4 | 27.5 | 17.9
2007-04-29 | 9:46 | 4.31 | 7.49 | 12.4 | 95.4 | 27.6 | 17.9
2007-04-29 | 10:01 | 4.41 | 7.49 | 12.39 | 95.6 | 27.6 | 17.9
2007-04-29 | 10:16 | 4.55 | 7.5 | 12.34 | 95.5 | 27.6 | 17.9
2007-04-29 | 10:31 | 4.67 | 7.5 | 12.31 | 95.5 | 27.6 | 17.9
2007-04-29 | 10:46 | 4.83 | 7.5 | 12.26 | 95.6 | 27.6 | 17.9
2007-04-29 | 11:01 | 4.99 | 7.5 | 12.21 | 95.5 | 27.6 | 17.9
2007-04-29 | 11:16 | 5.16 | 7.51 | 12.17 | 95.7 | 27.6 | 18
2007-04-29 | 11:31 | 5.36 | 7.51 | 12.1 | 95.6 | 27.7 | 18
2007-04-29 | 11:46 | 5.58 | 7.52 | 12.04 | 95.7 | 27.7 | 18
2007-04-29 | 12:01 | 5.78 | 7.52 | 11.97 | 95.7 | 27.7 | 18
2007-04-29 | 12:16 | 6.01 | 7.52 | 11.93 | 95.9 | 27.7 | 18
2007-04-29 | 12:31 | 6.24 | 7.53 | 11.87 | 96 | 27.7 | 18
2007-04-29 | 12:46 | 6.49 | 7.53 | 11.78 | 95.9 | 27.8 | 18
2007-04-29 | 13:00 | 6.6 | 7.53 | 11.76 | 95.9 | 27.7 | 18
2007-04-29 | 13:15 | 6.77 | 7.53 | 11.7 | 95.9 | 27.8 | 18.1
2007-04-29 | 13:30 | 6.88 | 7.53 | 11.67 | 95.8 | 27.8 | 18
2007-04-29 | 13:45 | 6.94 | 7.53 | 11.64 | 95.7 | 27.8 | 18
2007-04-29 | 14:00 | 7.12 | 7.54 | 11.6 | 95.9 | 27.8 | 18.1
2007-04-29 | 14:15 | 7.45 | 7.54 | 11.52 | 96 | 27.8 | 18.1
2007-04-29 | 14:30 | 7.72 | 7.55 | 11.46 | 96.1 | 27.8 | 18.1
2007-04-29 | 14:45 | 7.78 | 7.55 | 11.41 | 95.8 | 27.8 | 18.1
2007-04-29 | 15:00 | 7.92 | 7.55 | 11.37 | 95.8 | 27.8 | 18.1
2007-04-29 | 15:15 | 8.37 | 7.56 | 11.27 | 96 | 27.8 | 18.1
2007-04-29 | 15:30 | 8.86 | 7.56 | 11.13 | 95.9 | 27.8 | 18.1
2007-04-29 | 15:45 | 9.14 | 7.57 | 11.06 | 96 | 27.8 | 18.1
2007-04-29 | 16:00 | 9.25 | 7.57 | 11.02 | 95.9 | 27.8 | 18.1
2007-04-29 | 16:15 | 9.22 | 7.57 | 11.02 | 95.8 | 27.9 | 18.1
2007-04-29 | 16:30 | 9.18 | 7.56 | 11.02 | 95.7 | 27.9 | 18.1
2007-04-29 | 16:45 | 9.11 | 7.57 | 10.97 | 95.1 | 27.9 | 18.1
2007-04-29 | 17:00 | 9.08 | 7.56 | 11.01 | 95.5 | 27.9 | 18.1
2007-04-29 | 17:15 | 9.07 | 7.56 | 11.01 | 95.4 | 27.9 | 18.1
2007-04-29 | 17:30 | 9.06 | 7.56 | 11.04 | 95.6 | 27.9 | 18.1
2007-04-29 | 17:45 | 9.03 | 7.56 | 11.05 | 95.6 | 27.9 | 18.1
2007-04-29 | 18:00 | 8.99 | 7.56 | 11.05 | 95.5 | 27.8 | 18.1
2007-04-29 | 18:15 | 9.08 | 7.56 | 11.01 | 95.4 | 27.8 | 18.1
2007-04-29 | 18:30 | 9.3 | 7.57 | 10.97 | 95.5 | 27.8 | 18.1
2007-04-29 | 18:45 | 9.44 | 7.56 | 10.92 | 95.5 | 27.8 | 18.1
2007-04-29 | 19:00 | 9.52 | 7.56 | 10.91 | 95.5 | 27.8 | 18
2007-04-29 | 19:15 | 9.52 | 7.56 | 10.88 | 95.3 | 27.8 | 18.1
2007-04-29 | 19:30 | 9.46 | 7.56 | 10.92 | 95.5 | 27.7 | 18
2007-04-29 | 19:45 | 9.37 | 7.55 | 10.94 | 95.5 | 27.6 | 18
2007-04-29 | 20:00 | 9.23 | 7.54 | 10.98 | 95.5 | 27.6 | 17.9
2007-04-29 | 20:15 | 9.06 | 7.54 | 11.02 | 95.5 | 27.5 | 17.9
2007-04-29 | 20:30 | 8.86 | 7.53 | 11.07 | 95.5 | 27.4 | 17.8
2007-04-29 | 20:45 | 8.65 | 7.52 | 11.13 | 95.4 | 27.4 | 17.8
2007-04-29 | 21:00 | 8.44 | 7.51 | 11.2 | 95.6 | 27.3 | 17.7
2007-04-29 | 21:15 | 8.23 | 7.51 | 11.25 | 95.5 | 27.2 | 17.7
2007-04-29 | 21:30 | 8.02 | 7.5 | 11.3 | 95.4 | 27.1 | 17.6
2007-04-29 | 21:45 | 7.81 | 7.49 | 11.38 | 95.6 | 27.1 | 17.6
2007-04-29 | 22:00 | 7.59 | 7.48 | 11.43 | 95.5 | 26.9 | 17.5
2007-04-29 | 22:15 | 7.36 | 7.48 | 11.49 | 95.5 | 26.9 | 17.5
2007-04-29 | 22:30 | 7.14 | 7.47 | 11.56 | 95.5 | 26.8 | 17.4
2007-04-29 | 22:45 | 6.94 | 7.47 | 11.63 | 95.6 | 26.7 | 17.4
2007-04-29 | 23:00 | 6.75 | 7.45 | 11.67 | 95.5 | 26.6 | 17.3
2007-04-29 | 23:15 | 6.58 | 7.45 | 11.72 | 95.6 | 26.6 | 17.3
2007-04-29 | 23:30 | 6.41 | 7.46 | 11.78 | 95.6 | 26.5 | 17.2
2007-04-29 | 23:45 | 6.25 | 7.45 | 11.82 | 95.6 | 26.5 | 17.2
2007-04-30 | 0:00 | 6.1 | 7.44 | 11.87 | 95.5 | 26.4 | 17.2
2007-04-30 | 0:15 | 5.95 | 7.44 | 11.9 | 95.5 | 26.4 | 17.2
2007-04-30 | 0:30 | 5.81 | 7.44 | 11.95 | 95.5 | 26.4 | 17.1
2007-04-30 | 0:45 | 5.66 | 7.44 | 11.99 | 95.4 | 26.4 | 17.1
2007-04-30 | 1:01 | 5.52 | 7.43 | 12.03 | 95.5 | 26.4 | 17.1
2007-04-30 | 1:16 | 5.38 | 7.43 | 12.07 | 95.4 | 26.3 | 17.1
2007-04-30 | 1:31 | 5.25 | 7.43 | 12.12 | 95.5 | 26.4 | 17.1
2007-04-30 | 1:46 | 5.12 | 7.43 | 12.16 | 95.5 | 26.3 | 17.1
2007-04-30 | 2:01 | 5 | 7.43 | 12.2 | 95.5 | 26.4 | 17.1
2007-04-30 | 2:16 | 4.87 | 7.42 | 12.24 | 95.5 | 26.4 | 17.1
2007-04-30 | 2:31 | 4.76 | 7.42 | 12.27 | 95.5 | 26.4 | 17.2
2007-04-30 | 2:46 | 4.65 | 7.42 | 12.31 | 95.5 | 26.4 | 17.2
2007-04-30 | 3:01 | 4.54 | 7.42 | 12.36 | 95.6 | 26.4 | 17.2
2007-04-30 | 3:16 | 4.45 | 7.42 | 12.37 | 95.5 | 26.4 | 17.2
2007-04-30 | 3:31 | 4.36 | 7.41 | 12.4 | 95.5 | 26.5 | 17.2
2007-04-30 | 3:46 | 4.28 | 7.41 | 12.43 | 95.5 | 26.5 | 17.2
2007-04-30 | 4:01 | 4.21 | 7.42 | 12.45 | 95.5 | 26.5 | 17.2
2007-04-30 | 4:16 | 4.14 | 7.41 | 12.48 | 95.5 | 26.6 | 17.3
2007-04-30 | 4:31 | 4.05 | 7.41 | 12.49 | 95.4 | 26.6 | 17.3
2007-04-30 | 4:46 | 3.97 | 7.42 | 12.54 | 95.6 | 26.6 | 17.3
2007-04-30 | 5:01 | 3.89 | 7.42 | 12.55 | 95.5 | 26.6 | 17.3
2007-04-30 | 5:16 | 3.82 | 7.41 | 12.58 | 95.5 | 26.6 | 17.3
2007-04-30 | 5:31 | 3.74 | 7.42 | 12.59 | 95.4 | 26.7 | 17.3
2007-04-30 | 5:46 | 3.66 | 7.42 | 12.63 | 95.5 | 26.7 | 17.3
2007-04-30 | 6:01 | 3.59 | 7.42 | 12.65 | 95.5 | 26.7 | 17.4
2007-04-30 | 6:16 | 3.52 | 7.41 | 12.67 | 95.5 | 26.7 | 17.4
2007-04-30 | 6:31 | 3.45 | 7.42 | 12.7 | 95.5 | 26.8 | 17.4
2007-04-30 | 6:46 | 3.4 | 7.42 | 12.71 | 95.4 | 26.8 | 17.4
2007-04-30 | 7:01 | 3.36 | 7.42 | 12.74 | 95.5 | 26.8 | 17.4
2007-04-30 | 7:16 | 3.32 | 7.43 | 12.75 | 95.5 | 26.8 | 17.4
2007-04-30 | 7:31 | 3.29 | 7.43 | 12.76 | 95.5 | 26.9 | 17.5
2007-04-30 | 7:46 | 3.27 | 7.43 | 12.76 | 95.5 | 26.9 | 17.5
2007-04-30 | 8:01 | 3.26 | 7.43 | 12.78 | 95.6 | 26.9 | 17.5
2007-04-30 | 8:16 | 3.26 | 7.44 | 12.79 | 95.7 | 26.9 | 17.5
2007-04-30 | 8:31 | 3.28 | 7.44 | 12.78 | 95.7 | 26.9 | 17.5
2007-04-30 | 8:46 | 3.31 | 7.46 | 12.78 | 95.7 | 27 | 17.5
2007-04-30 | 9:01 | 3.36 | 7.46 | 12.77 | 95.8 | 27 | 17.5
2007-04-30 | 9:16 | 3.43 | 7.46 | 12.75 | 95.8 | 27 | 17.5
2007-04-30 | 9:31 | 3.5 | 7.47 | 12.74 | 95.9 | 27 | 17.6
2007-04-30 | 9:46 | 3.6 | 7.47 | 12.7 | 95.8 | 27 | 17.6
2007-04-30 | 10:01 | 3.72 | 7.48 | 12.68 | 96 | 27 | 17.6
2007-04-30 | 10:16 | 3.84 | 7.47 | 12.63 | 95.9 | 27 | 17.6
2007-04-30 | 10:31 | 3.97 | 7.48 | 12.57 | 95.8 | 27.1 | 17.6
2007-04-30 | 10:46 | 4.13 | 7.48 | 12.55 | 96.1 | 27.1 | 17.6
2007-04-30 | 11:01 | 4.29 | 7.48 | 12.49 | 96 | 27.1 | 17.6
2007-04-30 | 11:16 | 4.46 | 7.49 | 12.45 | 96.1 | 27.1 | 17.6
2007-04-30 | 11:31 | 4.64 | 7.49 | 12.39 | 96.1 | 27.1 | 17.6
2007-04-30 | 11:46 | 4.83 | 7.5 | 12.32 | 96.1 | 27.1 | 17.6
2007-04-30 | 12:01 | 5.01 | 7.5 | 12.27 | 96.1 | 27.1 | 17.6
2007-04-30 | 12:16 | 5.22 | 7.5 | 12.21 | 96.2 | 27.2 | 17.7
2007-04-30 | 12:31 | 5.44 | 7.5 | 12.16 | 96.3 | 27.2 | 17.7
2007-04-30 | 12:46 | 5.66 | 7.51 | 12.07 | 96.1 | 27.2 | 17.7
2007-04-30 | 13:00 | 5.91 | 7.51 | 12.03 | 96.4 | 27.2 | 17.7
2007-04-30 | 13:15 | 6.14 | 7.52 | 11.94 | 96.2 | 27.2 | 17.7
2007-04-30 | 13:30 | 6.33 | 7.52 | 11.87 | 96.2 | 27.3 | 17.7
2007-04-30 | 13:45 | 6.61 | 7.52 | 11.81 | 96.3 | 27.2 | 17.7
2007-04-30 | 14:00 | 6.75 | 7.53 | 11.74 | 96.1 | 27.3 | 17.7
2007-04-30 | 14:15 | 7.05 | 7.53 | 11.66 | 96.2 | 27.3 | 17.7
2007-04-30 | 14:30 | 7.3 | 7.53 | 11.6 | 96.2 | 27.3 | 17.8
2007-04-30 | 14:45 | 7.58 | 7.54 | 11.51 | 96.2 | 27.3 | 17.8
2007-04-30 | 15:00 | 7.86 | 7.54 | 11.45 | 96.3 | 27.3 | 17.8
2007-04-30 | 15:15 | 8.03 | 7.54 | 11.37 | 96.1 | 27.3 | 17.8
2007-04-30 | 15:30 | 8.27 | 7.55 | 11.31 | 96.1 | 27.3 | 17.8
2007-04-30 | 15:45 | 8.52 | 7.55 | 11.25 | 96.2 | 27.4 | 17.8
2007-04-30 | 16:00 | 8.74 | 7.55 | 11.19 | 96.2 | 27.4 | 17.8
2007-04-30 | 16:15 | 8.91 | 7.56 | 11.12 | 96 | 27.4 | 17.8
2007-04-30 | 16:30 | 8.95 | 7.56 | 11.08 | 95.7 | 27.4 | 17.8
2007-04-30 | 16:45 | 8.88 | 7.56 | 11.1 | 95.7 | 27.4 | 17.8
2007-04-30 | 17:00 | 8.85 | 7.55 | 11.1 | 95.6 | 27.4 | 17.8
2007-04-30 | 17:15 | 8.89 | 7.55 | 11.08 | 95.6 | 27.4 | 17.8
2007-04-30 | 17:30 | 9 | 7.55 | 11.04 | 95.5 | 27.4 | 17.8
2007-04-30 | 17:45 | 9.07 | 7.55 | 11.03 | 95.6 | 27.4 | 17.8
2007-04-30 | 18:00 | 9.11 | 7.55 | 11.02 | 95.6 | 27.4 | 17.8
2007-04-30 | 18:15 | 9.15 | 7.55 | 11.01 | 95.6 | 27.4 | 17.8
2007-04-30 | 18:30 | 9.2 | 7.55 | 10.99 | 95.6 | 27.4 | 17.8
2007-04-30 | 18:45 | 9.25 | 7.55 | 10.98 | 95.6 | 27.4 | 17.8
2007-04-30 | 19:00 | 9.24 | 7.55 | 10.97 | 95.4 | 27.4 | 17.8
2007-04-30 | 19:15 | 9.2 | 7.54 | 11 | 95.6 | 27.4 | 17.8
2007-04-30 | 19:30 | 9.13 | 7.54 | 11.01 | 95.5 | 27.3 | 17.8
2007-04-30 | 19:45 | 8.99 | 7.54 | 11.03 | 95.4 | 27.3 | 17.8
2007-04-30 | 20:00 | 8.83 | 7.54 | 11.09 | 95.6 | 27.3 | 17.7
2007-04-30 | 20:15 | 8.66 | 7.53 | 11.14 | 95.6 | 27.2 | 17.7
2007-04-30 | 20:30 | 8.48 | 7.52 | 11.19 | 95.6 | 27.2 | 17.7
2007-04-30 | 20:45 | 8.31 | 7.51 | 11.23 | 95.5 | 27.1 | 17.6
2007-04-30 | 21:00 | 8.14 | 7.5 | 11.28 | 95.6 | 27.1 | 17.6
2007-04-30 | 21:15 | 7.98 | 7.5 | 11.32 | 95.6 | 27 | 17.6
2007-04-30 | 21:30 | 7.82 | 7.49 | 11.38 | 95.7 | 27 | 17.5
2007-04-30 | 21:45 | 7.67 | 7.48 | 11.42 | 95.7 | 26.9 | 17.5
2007-04-30 | 22:00 | 7.53 | 7.48 | 11.45 | 95.6 | 26.9 | 17.5
2007-04-30 | 22:15 | 7.41 | 7.48 | 11.48 | 95.5 | 26.8 | 17.4
2007-04-30 | 22:30 | 7.3 | 7.47 | 11.51 | 95.5 | 26.7 | 17.4
2007-04-30 | 22:45 | 7.18 | 7.47 | 11.55 | 95.6 | 26.7 | 17.4
2007-04-30 | 23:00 | 7.06 | 7.46 | 11.58 | 95.5 | 26.6 | 17.3
2007-04-30 | 23:15 | 6.93 | 7.46 | 11.62 | 95.5 | 26.6 | 17.3
2007-04-30 | 23:30 | 6.8 | 7.46 | 11.65 | 95.5 | 26.6 | 17.3
2007-04-30 | 23:45 | 6.67 | 7.46 | 11.69 | 95.5 | 26.5 | 17.2
2007-05-01 | 0:00 | 6.54 | 7.45 | 11.72 | 95.4 | 26.5 | 17.2
2007-05-01 | 0:15 | 6.4 | 7.45 | 11.77 | 95.5 | 26.5 | 17.2
2007-05-01 | 0:30 | 6.26 | 7.45 | 11.81 | 95.5 | 26.5 | 17.2
2007-05-01 | 0:45 | 6.11 | 7.44 | 11.85 | 95.5 | 26.4 | 17.2
2007-05-01 | 1:01 | 5.97 | 7.44 | 11.9 | 95.5 | 26.5 | 17.2
2007-05-01 | 1:16 | 5.83 | 7.44 | 11.93 | 95.5 | 26.4 | 17.2
2007-05-01 | 1:31 | 5.7 | 7.44 | 11.98 | 95.5 | 26.4 | 17.2
2007-05-01 | 1:46 | 5.58 | 7.44 | 12.01 | 95.4 | 26.4 | 17.2
2007-05-01 | 2:01 | 5.46 | 7.44 | 12.05 | 95.4 | 26.4 | 17.2
2007-05-01 | 2:16 | 5.34 | 7.43 | 12.07 | 95.3 | 26.5 | 17.2
2007-05-01 | 2:31 | 5.23 | 7.43 | 12.1 | 95.3 | 26.5 | 17.2
2007-05-01 | 2:46 | 5.13 | 7.43 | 12.15 | 95.4 | 26.5 | 17.2
2007-05-01 | 3:01 | 5.03 | 7.43 | 12.14 | 95.2 | 26.5 | 17.2
2007-05-01 | 3:16 | 4.94 | 7.43 | 12.21 | 95.4 | 26.5 | 17.2
2007-05-01 | 3:31 | 4.86 | 7.43 | 12.25 | 95.5 | 26.5 | 17.2
2007-05-01 | 3:46 | 4.79 | 7.43 | 12.25 | 95.4 | 26.5 | 17.3
2007-05-01 | 4:01 | 4.73 | 7.43 | 12.28 | 95.5 | 26.6 | 17.3
2007-05-01 | 4:16 | 4.69 | 7.43 | 12.28 | 95.4 | 26.6 | 17.3
2007-05-01 | 4:31 | 4.65 | 7.43 | 12.29 | 95.4 | 26.6 | 17.3
2007-05-01 | 4:46 | 4.61 | 7.43 | 12.3 | 95.4 | 26.6 | 17.3
2007-05-01 | 5:01 | 4.57 | 7.43 | 12.33 | 95.4 | 26.7 | 17.3
2007-05-01 | 5:16 | 4.53 | 7.43 | 12.32 | 95.3 | 26.6 | 17.3
2007-05-01 | 5:31 | 4.5 | 7.43 | 12.34 | 95.4 | 26.7 | 17.3
2007-05-01 | 5:46 | 4.47 | 7.43 | 12.35 | 95.4 | 26.7 | 17.3
2007-05-01 | 6:01 | 4.43 | 7.43 | 12.36 | 95.4 | 26.7 | 17.4
2007-05-01 | 6:16 | 4.39 | 7.43 | 12.37 | 95.4 | 26.7 | 17.4
2007-05-01 | 6:31 | 4.36 | 7.43 | 12.39 | 95.4 | 26.8 | 17.4
2007-05-01 | 6:46 | 4.34 | 7.43 | 12.42 | 95.6 | 26.8 | 17.4
2007-05-01 | 7:01 | 4.33 | 7.43 | 12.39 | 95.3 | 26.8 | 17.4
2007-05-01 | 7:16 | 4.32 | 7.43 | 12.41 | 95.5 | 26.8 | 17.4
2007-05-01 | 7:31 | 4.33 | 7.44 | 12.41 | 95.5 | 26.8 | 17.4
2007-05-01 | 7:46 | 4.33 | 7.44 | 12.41 | 95.5 | 26.8 | 17.4
2007-05-01 | 8:01 | 4.34 | 7.45 | 12.42 | 95.5 | 26.9 | 17.5
2007-05-01 | 8:16 | 4.34 | 7.46 | 12.42 | 95.6 | 26.8 | 17.4
2007-05-01 | 8:31 | 4.34 | 7.47 | 12.43 | 95.7 | 26.9 | 17.5
2007-05-01 | 8:46 | 4.34 | 7.47 | 12.43 | 95.6 | 26.9 | 17.5
2007-05-01 | 9:01 | 4.33 | 7.47 | 12.43 | 95.6 | 26.9 | 17.5
2007-05-01 | 9:16 | 4.34 | 7.48 | 12.43 | 95.7 | 26.9 | 17.5
2007-05-01 | 9:31 | 4.34 | 7.48 | 12.43 | 95.6 | 26.9 | 17.5
2007-05-01 | 9:46 | 4.35 | 7.48 | 12.42 | 95.7 | 26.9 | 17.5
2007-05-01 | 10:01 | 4.37 | 7.48 | 12.42 | 95.6 | 26.9 | 17.5
2007-05-01 | 10:16 | 4.4 | 7.48 | 12.35 | 95.2 | 27 | 17.5
2007-05-01 | 10:31 | 4.46 | 7.49 | 12.32 | 95.1 | 27 | 17.5
2007-05-01 | 10:46 | 4.53 | 7.49 | 12.34 | 95.5 | 27 | 17.5
2007-05-01 | 11:01 | 4.6 | 7.49 | 12.34 | 95.6 | 27 | 17.5
2007-05-01 | 11:16 | 4.64 | 7.49 | 12.33 | 95.6 | 27 | 17.6
2007-05-01 | 11:31 | 4.67 | 7.49 | 12.32 | 95.7 | 27 | 17.6
2007-05-01 | 11:46 | 4.71 | 7.5 | 12.31 | 95.7 | 27 | 17.6
2007-05-01 | 12:01 | 4.78 | 7.49 | 12.3 | 95.8 | 27 | 17.6
2007-05-01 | 12:16 | 4.84 | 7.5 | 12.27 | 95.7 | 27 | 17.6
2007-05-01 | 12:31 | 4.92 | 7.5 | 12.25 | 95.7 | 27 | 17.6
2007-05-01 | 12:46 | 5.01 | 7.5 | 12.22 | 95.7 | 27 | 17.6
2007-05-01 | 13:00 | 5.1 | 7.51 | 12.2 | 95.8 | 27 | 17.6
2007-05-01 | 13:15 | 5.2 | 7.51 | 12.16 | 95.7 | 27 | 17.6
2007-05-01 | 13:30 | 5.3 | 7.5 | 12.15 | 95.8 | 27 | 17.6
2007-05-01 | 13:45 | 5.37 | 7.51 | 12.11 | 95.7 | 27 | 17.6
2007-05-01 | 14:00 | 5.45 | 7.51 | 12.08 | 95.7 | 27 | 17.6
2007-05-01 | 14:15 | 5.53 | 7.51 | 12.05 | 95.7 | 27 | 17.6
2007-05-01 | 14:30 | 5.64 | 7.51 | 12.03 | 95.7 | 27 | 17.5
2007-05-01 | 14:45 | 5.75 | 7.51 | 11.99 | 95.7 | 27 | 17.5
2007-05-01 | 15:00 | 5.84 | 7.51 | 11.95 | 95.6 | 27 | 17.5
2007-05-01 | 15:15 | 5.92 | 7.51 | 11.94 | 95.7 | 27 | 17.5
2007-05-01 | 15:30 | 5.95 | 7.51 | 11.91 | 95.6 | 27 | 17.5
2007-05-01 | 15:45 | 5.98 | 7.51 | 11.91 | 95.7 | 27 | 17.5
2007-05-01 | 16:00 | 6.03 | 7.52 | 11.9 | 95.7 | 27 | 17.5
2007-05-01 | 16:15 | 6.03 | 7.51 | 11.88 | 95.5 | 27 | 17.5
2007-05-01 | 16:30 | 6.04 | 7.51 | 11.87 | 95.4 | 27 | 17.6
2007-05-01 | 16:45 | 6.03 | 7.51 | 11.88 | 95.5 | 27 | 17.6
2007-05-01 | 17:00 | 6.02 | 7.51 | 11.88 | 95.5 | 27 | 17.6
2007-05-01 | 17:15 | 6.02 | 7.51 | 11.88 | 95.5 | 27 | 17.5
2007-05-01 | 17:30 | 6.01 | 7.51 | 11.88 | 95.4 | 27 | 17.5
2007-05-01 | 17:45 | 6.01 | 7.51 | 11.88 | 95.4 | 27 | 17.5
2007-05-01 | 18:00 | 6.01 | 7.51 | 11.88 | 95.5 | 27 | 17.5
2007-05-01 | 18:15 | 6.01 | 7.51 | 11.87 | 95.4 | 27 | 17.5
2007-05-01 | 18:30 | 6.02 | 7.51 | 11.86 | 95.3 | 26.9 | 17.5
2007-05-01 | 18:45 | 6.03 | 7.51 | 11.85 | 95.2 | 26.9 | 17.5
2007-05-01 | 19:00 | 6.04 | 7.5 | 11.84 | 95.2 | 26.9 | 17.5
2007-05-01 | 19:15 | 6.05 | 7.5 | 11.84 | 95.2 | 26.9 | 17.5
2007-05-01 | 19:30 | 6.05 | 7.5 | 11.85 | 95.3 | 26.8 | 17.5
2007-05-01 | 19:45 | 6.04 | 7.49 | 11.82 | 95.1 | 26.8 | 17.4
2007-05-01 | 20:00 | 6.03 | 7.49 | 11.85 | 95.3 | 26.8 | 17.4
2007-05-01 | 20:15 | 6.02 | 7.49 | 11.86 | 95.3 | 26.8 | 17.4
2007-05-01 | 20:30 | 6.01 | 7.48 | 11.85 | 95.2 | 26.7 | 17.4
2007-05-01 | 20:45 | 6 | 7.48 | 11.85 | 95.2 | 26.7 | 17.4
2007-05-01 | 21:00 | 5.98 | 7.47 | 11.86 | 95.3 | 26.7 | 17.4
2007-05-01 | 21:15 | 5.96 | 7.47 | 11.85 | 95.1 | 26.7 | 17.3
2007-05-01 | 21:30 | 5.94 | 7.46 | 11.87 | 95.2 | 26.6 | 17.3
2007-05-01 | 21:45 | 5.91 | 7.45 | 11.87 | 95.2 | 26.6 | 17.3
2007-05-01 | 22:00 | 5.88 | 7.45 | 11.89 | 95.2 | 26.6 | 17.3
2007-05-01 | 22:15 | 5.84 | 7.45 | 11.9 | 95.2 | 26.6 | 17.3
2007-05-01 | 22:30 | 5.79 | 7.45 | 11.92 | 95.2 | 26.6 | 17.3
2007-05-01 | 22:45 | 5.74 | 7.45 | 11.92 | 95.1 | 26.6 | 17.3
2007-05-01 | 23:00 | 5.68 | 7.45 | 11.93 | 95.1 | 26.6 | 17.3
2007-05-01 | 23:15 | 5.63 | 7.46 | 11.95 | 95.1 | 26.5 | 17.2
2007-05-01 | 23:30 | 5.59 | 7.45 | 11.97 | 95.1 | 26.5 | 17.2
2007-05-01 | 23:45 | 5.55 | 7.45 | 11.99 | 95.2 | 26.5 | 17.2
2007-05-02 | 0:00 | 5.52 | 7.46 | 11.99 | 95.2 | 26.5 | 17.2
2007-05-02 | 0:15 | 5.48 | 7.46 | 11.99 | 95 | 26.4 | 17.2
2007-05-02 | 0:30 | 5.43 | 7.46 | 12.02 | 95.1 | 26.4 | 17.1
2007-05-02 | 0:45 | 5.38 | 7.46 | 12.03 | 95.1 | 26.4 | 17.2
2007-05-02 | 1:01 | 5.32 | 7.46 | 12.06 | 95.2 | 26.4 | 17.1
2007-05-02 | 1:16 | 5.27 | 7.45 | 12.07 | 95.2 | 26.4 | 17.1
2007-05-02 | 1:31 | 5.23 | 7.45 | 12.08 | 95.1 | 26.3 | 17.1
2007-05-02 | 1:46 | 5.18 | 7.45 | 12.1 | 95.1 | 26.3 | 17.1
2007-05-02 | 2:01 | 5.12 | 7.45 | 12.1 | 95 | 26.3 | 17.1
2007-05-02 | 2:16 | 5.06 | 7.44 | 12.14 | 95.2 | 26.3 | 17.1
2007-05-02 | 2:31 | 4.99 | 7.44 | 12.15 | 95.1 | 26.3 | 17.1
2007-05-02 | 2:46 | 4.94 | 7.44 | 12.16 | 95.1 | 26.2 | 17.1
2007-05-02 | 3:01 | 4.89 | 7.44 | 12.18 | 95.1 | 26.2 | 17
2007-05-02 | 3:16 | 4.84 | 7.43 | 12.2 | 95.1 | 26.2 | 17
2007-05-02 | 3:31 | 4.78 | 7.44 | 12.21 | 95.1 | 26.2 | 17
2007-05-02 | 3:46 | 4.73 | 7.44 | 12.25 | 95.2 | 26.2 | 17
2007-05-02 | 4:01 | 4.68 | 7.44 | 12.24 | 95.1 | 26.2 | 17
2007-05-02 | 4:16 | 4.63 | 7.44 | 12.27 | 95.1 | 26.2 | 17
2007-05-02 | 4:31 | 4.57 | 7.43 | 12.29 | 95.1 | 26.1 | 17
2007-05-02 | 4:46 | 4.51 | 7.44 | 12.29 | 95.1 | 26.1 | 16.9
2007-05-02 | 5:01 | 4.46 | 7.43 | 12.32 | 95.1 | 26 | 16.9
2007-05-02 | 5:16 | 4.41 | 7.44 | 12.32 | 95 | 25.9 | 16.9
2007-05-02 | 5:31 | 4.37 | 7.44 | 12.36 | 95.2 | 25.9 | 16.8
2007-05-02 | 5:46 | 4.31 | 7.44 | 12.38 | 95.2 | 25.8 | 16.8
2007-05-02 | 6:01 | 4.25 | 7.44 | 12.39 | 95.1 | 25.9 | 16.8
2007-05-02 | 6:16 | 4.2 | 7.43 | 12.41 | 95.2 | 25.8 | 16.8
2007-05-02 | 6:31 | 4.14 | 7.44 | 12.43 | 95.2 | 25.8 | 16.8
2007-05-02 | 6:46 | 4.09 | 7.44 | 12.45 | 95.2 | 25.8 | 16.7
2007-05-02 | 7:01 | 4.06 | 7.44 | 12.48 | 95.3 | 25.7 | 16.7
2007-05-02 | 7:16 | 4.03 | 7.44 | 12.48 | 95.3 | 25.7 | 16.7
2007-05-02 | 7:31 | 4.01 | 7.44 | 12.5 | 95.4 | 25.7 | 16.7
2007-05-02 | 7:46 | 3.99 | 7.44 | 12.5 | 95.4 | 25.6 | 16.7
2007-05-02 | 8:01 | 3.95 | 7.45 | 12.52 | 95.4 | 25.6 | 16.7
2007-05-02 | 8:16 | 3.93 | 7.45 | 12.54 | 95.5 | 25.6 | 16.6
2007-05-02 | 8:31 | 3.91 | 7.45 | 12.55 | 95.5 | 25.6 | 16.6
2007-05-02 | 8:46 | 3.87 | 7.44 | 12.55 | 95.4 | 25.5 | 16.6
2007-05-02 | 9:01 | 3.84 | 7.45 | 12.57 | 95.5 | 25.5 | 16.6
2007-05-02 | 9:16 | 3.86 | 7.45 | 12.56 | 95.5 | 25.5 | 16.6
2007-05-02 | 9:31 | 3.88 | 7.45 | 12.55 | 95.5 | 25.5 | 16.6
2007-05-02 | 9:46 | 3.96 | 7.45 | 12.57 | 95.8 | 25.5 | 16.6
2007-05-02 | 10:01 | 4 | 7.46 | 12.56 | 95.8 | 25.5 | 16.6
2007-05-02 | 10:16 | 3.97 | 7.45 | 12.56 | 95.7 | 25.5 | 16.6
2007-05-02 | 10:31 | 3.9 | 7.45 | 12.58 | 95.7 | 25.4 | 16.5
2007-05-02 | 10:46 | 3.83 | 7.45 | 12.59 | 95.6 | 25.4 | 16.5
2007-05-02 | 11:01 | 3.8 | 7.45 | 12.61 | 95.7 | 25.6 | 16.6
2007-05-02 | 11:16 | 3.81 | 7.45 | 12.61 | 95.7 | 25.5 | 16.6
2007-05-02 | 11:31 | 3.82 | 7.45 | 12.62 | 95.8 | 25.5 | 16.6
2007-05-02 | 11:46 | 3.96 | 7.46 | 12.59 | 96 | 25.5 | 16.5
2007-05-02 | 12:01 | 4.24 | 7.46 | 12.48 | 95.8 | 25.5 | 16.6
2007-05-02 | 12:16 | 4.57 | 7.46 | 12.39 | 96 | 25.5 | 16.5
2007-05-02 | 12:31 | 4.88 | 7.47 | 12.32 | 96.1 | 25.5 | 16.6
2007-05-02 | 12:46 | 5.13 | 7.48 | 12.23 | 96.1 | 25.5 | 16.6
2007-05-02 | 13:00 | 5.25 | 7.48 | 12.2 | 96.1 | 25.5 | 16.6
2007-05-02 | 13:15 | 5.39 | 7.48 | 12.16 | 96.2 | 25.5 | 16.6
2007-05-02 | 13:30 | 5.55 | 7.48 | 12.1 | 96.1 | 25.5 | 16.6
2007-05-02 | 13:45 | 5.64 | 7.49 | 12.06 | 96 | 25.6 | 16.6
2007-05-02 | 14:00 | 5.78 | 7.48 | 12.05 | 96.2 | 25.6 | 16.6
2007-05-02 | 14:15 | 6.11 | 7.49 | 11.94 | 96.2 | 25.6 | 16.7
2007-05-02 | 14:30 | 6.35 | 7.5 | 11.86 | 96.1 | 25.6 | 16.7
2007-05-02 | 14:45 | 6.77 | 7.49 | 11.75 | 96.2 | 25.7 | 16.7
2007-05-02 | 15:00 | 6.93 | 7.49 | 11.71 | 96.3 | 25.7 | 16.7
2007-05-02 | 15:15 | 7.01 | 7.5 | 11.68 | 96.2 | 25.7 | 16.7
2007-05-02 | 15:30 | 7.02 | 7.5 | 11.67 | 96.2 | 25.7 | 16.7
2007-05-02 | 15:45 | 7.05 | 7.5 | 11.66 | 96.2 | 25.8 | 16.8
2007-05-02 | 16:00 | 7.11 | 7.51 | 11.64 | 96.1 | 25.8 | 16.8
2007-05-02 | 16:15 | 7.16 | 7.51 | 11.62 | 96.1 | 25.8 | 16.8
2007-05-02 | 16:30 | 7.18 | 7.51 | 11.61 | 96.1 | 25.8 | 16.8
2007-05-02 | 16:45 | 7.15 | 7.5 | 11.62 | 96.1 | 25.9 | 16.8
2007-05-02 | 17:00 | 7.05 | 7.5 | 11.64 | 96 | 25.8 | 16.8
2007-05-02 | 17:15 | 6.87 | 7.5 | 11.68 | 95.9 | 25.9 | 16.8
2007-05-02 | 17:30 | 6.78 | 7.5 | 11.74 | 96.1 | 25.9 | 16.8
2007-05-02 | 17:45 | 6.69 | 7.5 | 11.75 | 96.1 | 25.9 | 16.8
2007-05-02 | 18:00 | 6.55 | 7.49 | 11.77 | 95.9 | 25.9 | 16.8
2007-05-02 | 18:15 | 6.47 | 7.49 | 11.81 | 96 | 25.9 | 16.9
2007-05-02 | 18:30 | 6.33 | 7.49 | 11.86 | 96.1 | 26 | 16.9
2007-05-02 | 18:45 | 6.19 | 7.48 | 11.88 | 95.9 | 26 | 16.9
2007-05-02 | 19:00 | 6 | 7.48 | 11.93 | 95.9 | 26 | 16.9
2007-05-02 | 19:15 | 5.77 | 7.47 | 12.01 | 95.9 | 26.1 | 16.9
2007-05-02 | 19:30 | 5.52 | 7.47 | 12.07 | 95.8 | 26.1 | 17
2007-05-02 | 19:45 | 5.27 | 7.46 | 12.16 | 95.9 | 26.1 | 17
2007-05-02 | 20:00 | 5.06 | 7.46 | 12.23 | 95.9 | 26.1 | 17
2007-05-02 | 20:15 | 4.84 | 7.45 | 12.29 | 95.8 | 26.1 | 17
2007-05-02 | 20:30 | 4.62 | 7.45 | 12.36 | 95.9 | 26.1 | 17
2007-05-02 | 20:45 | 4.44 | 7.45 | 12.41 | 95.7 | 26.2 | 17
2007-05-02 | 21:00 | 4.3 | 7.44 | 12.47 | 95.8 | 26.2 | 17
2007-05-02 | 21:15 | 4.15 | 7.44 | 12.51 | 95.8 | 26.2 | 17
2007-05-02 | 21:30 | 3.98 | 7.44 | 12.57 | 95.8 | 26.2 | 17.1
2007-05-02 | 21:45 | 3.83 | 7.43 | 12.62 | 95.8 | 26.2 | 17
2007-05-02 | 22:00 | 3.68 | 7.43 | 12.67 | 95.9 | 26.2 | 17.1
2007-05-02 | 22:15 | 3.53 | 7.43 | 12.72 | 95.8 | 26.3 | 17.1
2007-05-02 | 22:30 | 3.4 | 7.43 | 12.76 | 95.8 | 26.3 | 17.1
2007-05-02 | 22:45 | 3.26 | 7.42 | 12.81 | 95.8 | 26.3 | 17.1
2007-05-02 | 23:00 | 3.14 | 7.42 | 12.86 | 95.9 | 26.4 | 17.2
2007-05-02 | 23:15 | 3.04 | 7.42 | 12.88 | 95.8 | 26.4 | 17.2
2007-05-02 | 23:30 | 2.96 | 7.42 | 12.91 | 95.8 | 26.5 | 17.2
2007-05-02 | 23:45 | 2.89 | 7.42 | 12.93 | 95.8 | 26.5 | 17.2
2007-05-03 | 0:00 | 2.82 | 7.42 | 12.96 | 95.8 | 26.6 | 17.3
2007-05-03 | 0:15 | 2.73 | 7.42 | 12.99 | 95.8 | 26.6 | 17.3
2007-05-03 | 0:30 | 2.65 | 7.42 | 13.02 | 95.8 | 26.6 | 17.3
2007-05-03 | 0:45 | 2.57 | 7.41 | 13.04 | 95.8 | 26.7 | 17.3
2007-05-03 | 1:01 | 2.49 | 7.41 | 13.07 | 95.8 | 26.7 | 17.3
2007-05-03 | 1:16 | 2.42 | 7.41 | 13.11 | 95.9 | 26.7 | 17.4
2007-05-03 | 1:31 | 2.35 | 7.41 | 13.12 | 95.8 | 26.8 | 17.4
2007-05-03 | 1:46 | 2.28 | 7.41 | 13.15 | 95.8 | 26.8 | 17.4
2007-05-03 | 2:01 | 2.21 | 7.41 | 13.15 | 95.6 | 26.8 | 17.4
2007-05-03 | 2:16 | 2.13 | 7.41 | 13.2 | 95.8 | 26.8 | 17.4
2007-05-03 | 2:31 | 2.05 | 7.41 | 13.23 | 95.8 | 26.9 | 17.5
2007-05-03 | 2:46 | 1.98 | 7.41 | 13.25 | 95.7 | 26.9 | 17.5
2007-05-03 | 3:01 | 1.91 | 7.41 | 13.28 | 95.8 | 26.9 | 17.5
2007-05-03 | 3:16 | 1.85 | 7.41 | 13.31 | 95.8 | 27 | 17.5
2007-05-03 | 3:31 | 1.79 | 7.41 | 13.33 | 95.8 | 27 | 17.5
2007-05-03 | 3:46 | 1.75 | 7.41 | 13.33 | 95.7 | 27 | 17.6
2007-05-03 | 4:01 | 1.7 | 7.41 | 13.36 | 95.8 | 27 | 17.6
2007-05-03 | 4:16 | 1.65 | 7.41 | 13.37 | 95.8 | 27.1 | 17.6
2007-05-03 | 4:31 | 1.6 | 7.41 | 13.39 | 95.8 | 27.1 | 17.6
2007-05-03 | 4:46 | 1.55 | 7.41 | 13.42 | 95.8 | 27.1 | 17.6
2007-05-03 | 5:01 | 1.51 | 7.41 | 13.44 | 95.8 | 27.1 | 17.6
2007-05-03 | 5:16 | 1.47 | 7.41 | 13.45 | 95.9 | 27.2 | 17.7
2007-05-03 | 5:31 | 1.43 | 7.41 | 13.46 | 95.8 | 27.2 | 17.7
2007-05-03 | 5:46 | 1.4 | 7.41 | 13.47 | 95.8 | 27.2 | 17.7
2007-05-03 | 6:01 | 1.36 | 7.41 | 13.49 | 95.9 | 27.2 | 17.7
2007-05-03 | 6:16 | 1.34 | 7.41 | 13.49 | 95.8 | 27.3 | 17.7
2007-05-03 | 6:31 | 1.31 | 7.41 | 13.51 | 95.9 | 27.3 | 17.7
2007-05-03 | 6:46 | 1.29 | 7.41 | 13.53 | 95.9 | 27.3 | 17.7
2007-05-03 | 7:01 | 1.28 | 7.41 | 13.54 | 96 | 27.3 | 17.8
2007-05-03 | 7:16 | 1.29 | 7.41 | 13.53 | 96 | 27.3 | 17.8
2007-05-03 | 7:31 | 1.29 | 7.42 | 13.53 | 95.9 | 27.4 | 17.8
2007-05-03 | 7:46 | 1.32 | 7.42 | 13.51 | 95.9 | 27.4 | 17.8
2007-05-03 | 8:01 | 1.37 | 7.43 | 13.52 | 96.1 | 27.4 | 17.8
2007-05-03 | 8:16 | 1.41 | 7.44 | 13.51 | 96.1 | 27.4 | 17.8
2007-05-03 | 8:31 | 1.43 | 7.44 | 13.5 | 96.1 | 27.4 | 17.8
2007-05-03 | 8:46 | 1.48 | 7.44 | 13.49 | 96.2 | 27.5 | 17.8
2007-05-03 | 9:01 | 1.55 | 7.45 | 13.5 | 96.4 | 27.4 | 17.8
2007-05-03 | 9:16 | 1.6 | 7.45 | 13.45 | 96.2 | 27.5 | 17.8
2007-05-03 | 9:31 | 1.65 | 7.45 | 13.44 | 96.3 | 27.5 | 17.9
2007-05-03 | 9:46 | 1.81 | 7.45 | 13.4 | 96.4 | 27.5 | 17.9
2007-05-03 | 10:01 | 1.84 | 7.45 | 13.4 | 96.4 | 27.5 | 17.9
2007-05-03 | 10:16 | 1.85 | 7.46 | 13.37 | 96.3 | 27.5 | 17.9
2007-05-03 | 10:31 | 2.02 | 7.46 | 13.32 | 96.4 | 27.5 | 17.9
2007-05-03 | 10:46 | 2.11 | 7.46 | 13.29 | 96.4 | 27.5 | 17.9
2007-05-03 | 11:01 | 2.19 | 7.46 | 13.28 | 96.6 | 27.6 | 17.9
2007-05-03 | 11:16 | 2.28 | 7.47 | 13.25 | 96.5 | 27.6 | 17.9
2007-05-03 | 11:31 | 2.43 | 7.46 | 13.19 | 96.5 | 27.6 | 17.9
2007-05-03 | 11:46 | 2.74 | 7.47 | 13.08 | 96.5 | 27.6 | 17.9
2007-05-03 | 12:01 | 3.08 | 7.48 | 12.99 | 96.7 | 27.6 | 17.9
2007-05-03 | 12:16 | 3.41 | 7.48 | 12.91 | 96.9 | 27.6 | 17.9
2007-05-03 | 12:31 | 3.67 | 7.48 | 12.78 | 96.6 | 27.6 | 17.9
2007-05-03 | 12:46 | 3.87 | 7.49 | 12.72 | 96.7 | 27.6 | 17.9
2007-05-03 | 13:00 | 4.2 | 7.49 | 12.63 | 96.8 | 27.5 | 17.9
2007-05-03 | 13:15 | 4.48 | 7.49 | 12.52 | 96.8 | 27.5 | 17.9
2007-05-03 | 13:30 | 4.79 | 7.5 | 12.43 | 96.8 | 27.5 | 17.9
2007-05-03 | 13:45 | 4.94 | 7.5 | 12.37 | 96.7 | 27.5 | 17.9
2007-05-03 | 14:00 | 5.13 | 7.5 | 12.31 | 96.7 | 27.5 | 17.9
2007-05-03 | 14:15 | 5.47 | 7.51 | 12.22 | 96.9 | 27.5 | 17.9
2007-05-03 | 14:30 | 5.67 | 7.51 | 12.16 | 96.9 | 27.5 | 17.9
2007-05-03 | 14:45 | 5.91 | 7.52 | 12.09 | 96.9 | 27.5 | 17.9
2007-05-03 | 15:00 | 6.07 | 7.52 | 12.02 | 96.7 | 27.5 | 17.9
2007-05-03 | 15:15 | 6.27 | 7.52 | 11.97 | 96.8 | 27.5 | 17.9
2007-05-03 | 15:30 | 6.43 | 7.52 | 11.91 | 96.7 | 27.5 | 17.8
2007-05-03 | 15:45 | 6.55 | 7.52 | 11.87 | 96.7 | 27.4 | 17.8
2007-05-03 | 16:00 | 6.6 | 7.52 | 11.84 | 96.6 | 27.4 | 17.8
2007-05-03 | 16:15 | 6.67 | 7.53 | 11.84 | 96.7 | 27.4 | 17.8
2007-05-03 | 16:30 | 6.72 | 7.53 | 11.81 | 96.6 | 27.4 | 17.8
2007-05-03 | 16:45 | 6.73 | 7.53 | 11.8 | 96.6 | 27.4 | 17.8
2007-05-03 | 17:00 | 6.74 | 7.52 | 11.8 | 96.6 | 27.4 | 17.8
2007-05-03 | 17:15 | 6.67 | 7.52 | 11.81 | 96.5 | 27.4 | 17.8
2007-05-03 | 17:30 | 6.58 | 7.52 | 11.83 | 96.4 | 27.4 | 17.8
2007-05-03 | 17:45 | 6.48 | 7.52 | 11.86 | 96.4 | 27.4 | 17.8
2007-05-03 | 18:00 | 6.41 | 7.52 | 11.86 | 96.3 | 27.4 | 17.8
2007-05-03 | 18:15 | 6.31 | 7.52 | 11.9 | 96.4 | 27.4 | 17.8
2007-05-03 | 18:30 | 6.17 | 7.51 | 11.94 | 96.3 | 27.4 | 17.8
2007-05-03 | 18:45 | 6.08 | 7.51 | 11.96 | 96.3 | 27.4 | 17.8
2007-05-03 | 19:00 | 5.98 | 7.51 | 12 | 96.3 | 27.4 | 17.8
2007-05-03 | 19:15 | 5.85 | 7.51 | 12.03 | 96.2 | 27.5 | 17.9
2007-05-03 | 19:30 | 5.73 | 7.5 | 12.06 | 96.2 | 27.5 | 17.8
2007-05-03 | 19:45 | 5.61 | 7.5 | 12.09 | 96.2 | 27.5 | 17.9
2007-05-03 | 20:00 | 5.5 | 7.49 | 12.13 | 96.2 | 27.5 | 17.9
2007-05-03 | 20:15 | 5.4 | 7.49 | 12.15 | 96.1 | 27.5 | 17.9
2007-05-03 | 20:30 | 5.27 | 7.49 | 12.19 | 96.1 | 27.5 | 17.9
2007-05-03 | 20:45 | 5.11 | 7.49 | 12.24 | 96.1 | 27.5 | 17.9
2007-05-03 | 21:00 | 4.94 | 7.48 | 12.29 | 96.1 | 27.6 | 17.9
2007-05-03 | 21:15 | 4.77 | 7.47 | 12.35 | 96.1 | 27.5 | 17.9
2007-05-03 | 21:30 | 4.62 | 7.48 | 12.41 | 96.2 | 27.6 | 17.9
2007-05-03 | 21:45 | 4.47 | 7.47 | 12.45 | 96.1 | 27.6 | 17.9
2007-05-03 | 22:00 | 4.35 | 7.47 | 12.5 | 96.2 | 27.6 | 17.9
2007-05-03 | 22:15 | 4.24 | 7.47 | 12.53 | 96.2 | 27.6 | 18
2007-05-03 | 22:30 | 4.13 | 7.46 | 12.56 | 96.1 | 27.6 | 18
2007-05-03 | 22:45 | 4.04 | 7.46 | 12.59 | 96.1 | 27.7 | 18
2007-05-03 | 23:00 | 3.95 | 7.46 | 12.62 | 96.2 | 27.7 | 18
2007-05-03 | 23:15 | 3.86 | 7.46 | 12.64 | 96.1 | 27.7 | 18
2007-05-03 | 23:30 | 3.75 | 7.46 | 12.7 | 96.2 | 27.7 | 18
2007-05-03 | 23:45 | 3.65 | 7.46 | 12.71 | 96.1 | 27.8 | 18
2007-05-04 | 0:00 | 3.57 | 7.46 | 12.74 | 96.1 | 27.8 | 18
2007-05-04 | 0:15 | 3.5 | 7.45 | 12.78 | 96.2 | 27.8 | 18.1
2007-05-04 | 0:30 | 3.43 | 7.46 | 12.8 | 96.2 | 27.8 | 18.1
2007-05-04 | 0:45 | 3.37 | 7.46 | 12.83 | 96.3 | 27.8 | 18.1
2007-05-04 | 1:01 | 3.31 | 7.46 | 12.84 | 96.1 | 27.9 | 18.1
2007-05-04 | 1:16 | 3.23 | 7.46 | 12.86 | 96.1 | 27.9 | 18.1
2007-05-04 | 1:31 | 3.16 | 7.46 | 12.89 | 96.2 | 27.9 | 18.2
2007-05-04 | 1:46 | 3.08 | 7.45 | 12.91 | 96.1 | 27.9 | 18.2
2007-05-04 | 2:01 | 3.01 | 7.45 | 12.95 | 96.2 | 28 | 18.2
2007-05-04 | 2:16 | 2.94 | 7.45 | 12.96 | 96.2 | 28 | 18.2
2007-05-04 | 2:31 | 2.88 | 7.45 | 12.98 | 96.1 | 28 | 18.2
2007-05-04 | 2:46 | 2.82 | 7.45 | 13.01 | 96.2 | 28.1 | 18.2
2007-05-04 | 3:01 | 2.77 | 7.45 | 13.03 | 96.2 | 28.1 | 18.2
2007-05-04 | 3:16 | 2.71 | 7.45 | 13.04 | 96.2 | 28.1 | 18.3
2007-05-04 | 3:31 | 2.66 | 7.45 | 13.05 | 96.1 | 28.1 | 18.3
2007-05-04 | 3:46 | 2.62 | 7.45 | 13.07 | 96.1 | 28.2 | 18.3
2007-05-04 | 4:01 | 2.58 | 7.45 | 13.09 | 96.2 | 28.2 | 18.3
2007-05-04 | 4:16 | 2.55 | 7.44 | 13.12 | 96.3 | 28.2 | 18.3
2007-05-04 | 4:31 | 2.51 | 7.45 | 13.12 | 96.2 | 28.2 | 18.3
2007-05-04 | 4:46 | 2.48 | 7.45 | 13.14 | 96.2 | 28.2 | 18.3
2007-05-04 | 5:01 | 2.44 | 7.45 | 13.14 | 96.2 | 28.2 | 18.4
2007-05-04 | 5:16 | 2.41 | 7.45 | 13.16 | 96.2 | 28.3 | 18.4
2007-05-04 | 5:31 | 2.39 | 7.44 | 13.16 | 96.2 | 28.3 | 18.4
2007-05-04 | 5:46 | 2.36 | 7.45 | 13.18 | 96.2 | 28.3 | 18.4
2007-05-04 | 6:01 | 2.33 | 7.45 | 13.19 | 96.3 | 28.3 | 18.4
2007-05-04 | 6:16 | 2.31 | 7.45 | 13.2 | 96.2 | 28.3 | 18.4
2007-05-04 | 6:31 | 2.28 | 7.45 | 13.21 | 96.3 | 28.3 | 18.4
2007-05-04 | 6:46 | 2.26 | 7.46 | 13.22 | 96.3 | 28.4 | 18.4
2007-05-04 | 7:01 | 2.24 | 7.46 | 13.23 | 96.3 | 28.4 | 18.4
2007-05-04 | 7:16 | 2.23 | 7.46 | 13.25 | 96.4 | 28.4 | 18.4
2007-05-04 | 7:31 | 2.22 | 7.47 | 13.25 | 96.4 | 28.4 | 18.4
2007-05-04 | 7:46 | 2.23 | 7.48 | 13.25 | 96.4 | 28.4 | 18.4
2007-05-04 | 8:01 | 2.24 | 7.48 | 13.25 | 96.5 | 28.4 | 18.4
2007-05-04 | 8:16 | 2.24 | 7.48 | 13.24 | 96.4 | 28.4 | 18.4
2007-05-04 | 8:31 | 2.25 | 7.48 | 13.25 | 96.5 | 28.4 | 18.4
2007-05-04 | 8:46 | 2.26 | 7.48 | 13.24 | 96.4 | 28.4 | 18.5
2007-05-04 | 9:01 | 2.28 | 7.48 | 13.25 | 96.5 | 28.4 | 18.5
2007-05-04 | 9:16 | 2.31 | 7.48 | 13.24 | 96.5 | 28.4 | 18.5
2007-05-04 | 9:31 | 2.38 | 7.48 | 13.23 | 96.6 | 28.4 | 18.5
2007-05-04 | 9:46 | 2.34 | 7.48 | 13.25 | 96.7 | 28.4 | 18.5
2007-05-04 | 10:01 | 2.39 | 7.48 | 13.24 | 96.7 | 28.4 | 18.5
2007-05-04 | 10:16 | 2.52 | 7.48 | 13.19 | 96.7 | 28.4 | 18.5
2007-05-04 | 10:31 | 2.58 | 7.48 | 13.17 | 96.8 | 28.4 | 18.5
2007-05-04 | 10:46 | 2.73 | 7.49 | 13.13 | 96.8 | 28.4 | 18.5
2007-05-04 | 11:01 | 2.88 | 7.49 | 13.07 | 96.8 | 28.4 | 18.5
2007-05-04 | 11:16 | 3.08 | 7.49 | 13.01 | 96.9 | 28.4 | 18.5
2007-05-04 | 11:31 | 3.17 | 7.49 | 12.98 | 96.9 | 28.4 | 18.5
2007-05-04 | 11:46 | 3.36 | 7.5 | 12.93 | 97 | 28.4 | 18.5
2007-05-04 | 12:01 | 3.47 | 7.5 | 12.88 | 96.9 | 28.4 | 18.5
2007-05-04 | 12:16 | 3.65 | 7.5 | 12.85 | 97.1 | 28.4 | 18.5
2007-05-04 | 12:31 | 3.86 | 7.5 | 12.78 | 97.1 | 28.4 | 18.5
2007-05-04 | 12:46 | 4.08 | 7.51 | 12.7 | 97.1 | 28.4 | 18.5
2007-05-04 | 13:00 | 4.31 | 7.51 | 12.64 | 97.2 | 28.4 | 18.5
2007-05-04 | 13:15 | 4.44 | 7.51 | 12.59 | 97.1 | 28.4 | 18.5
2007-05-04 | 13:30 | 4.6 | 7.51 | 12.5 | 96.9 | 28.4 | 18.5
2007-05-04 | 13:45 | 4.66 | 7.51 | 12.48 | 96.9 | 28.4 | 18.5
2007-05-04 | 14:00 | 5.14 | 7.52 | 12.36 | 97.1 | 28.4 | 18.5
2007-05-04 | 14:15 | 5.7 | 7.53 | 12.2 | 97.2 | 28.4 | 18.4
2007-05-04 | 14:30 | 5.9 | 7.54 | 12.13 | 97.2 | 28.4 | 18.4
2007-05-04 | 14:45 | 6.12 | 7.53 | 12.06 | 97.2 | 28.4 | 18.4
2007-05-04 | 15:00 | 6.19 | 7.54 | 12 | 96.9 | 28.4 | 18.5
2007-05-04 | 15:15 | 6.18 | 7.53 | 12.01 | 96.9 | 28.4 | 18.4
2007-05-04 | 15:30 | 6.42 | 7.54 | 11.98 | 97.3 | 28.4 | 18.4
2007-05-04 | 15:45 | 6.54 | 7.54 | 11.94 | 97.3 | 28.4 | 18.4
2007-05-04 | 16:00 | 6.61 | 7.54 | 11.88 | 97 | 28.4 | 18.4
2007-05-04 | 16:15 | 6.68 | 7.54 | 11.89 | 97.2 | 28.4 | 18.5
2007-05-04 | 16:30 | 6.72 | 7.54 | 11.87 | 97.1 | 28.4 | 18.5
2007-05-04 | 16:45 | 6.69 | 7.54 | 11.87 | 97 | 28.4 | 18.5
2007-05-04 | 17:00 | 6.61 | 7.54 | 11.87 | 96.8 | 28.3 | 18.4
2007-05-04 | 17:15 | 6.47 | 7.54 | 11.93 | 96.9 | 28.4 | 18.4
2007-05-04 | 17:30 | 6.46 | 7.54 | 11.93 | 97 | 28.3 | 18.4
2007-05-04 | 17:45 | 6.43 | 7.54 | 11.93 | 96.8 | 28.3 | 18.4
2007-05-04 | 18:00 | 6.45 | 7.53 | 11.92 | 96.8 | 28.3 | 18.4
2007-05-04 | 18:15 | 6.47 | 7.53 | 11.92 | 96.9 | 28.3 | 18.4
2007-05-04 | 18:30 | 6.44 | 7.54 | 11.9 | 96.7 | 28.3 | 18.4
2007-05-04 | 18:45 | 6.46 | 7.53 | 11.91 | 96.8 | 28.3 | 18.4
2007-05-04 | 19:00 | 6.41 | 7.53 | 11.93 | 96.8 | 28.3 | 18.4
2007-05-04 | 19:15 | 6.33 | 7.53 | 11.95 | 96.8 | 28.3 | 18.4
2007-05-04 | 19:30 | 6.22 | 7.53 | 11.97 | 96.7 | 28.3 | 18.4
2007-05-04 | 19:45 | 6.07 | 7.52 | 12.04 | 96.9 | 28.4 | 18.4
2007-05-04 | 20:00 | 5.91 | 7.52 | 12.08 | 96.8 | 28.3 | 18.4
2007-05-04 | 20:15 | 5.74 | 7.52 | 12.13 | 96.8 | 28.4 | 18.4
2007-05-04 | 20:30 | 5.58 | 7.51 | 12.18 | 96.8 | 28.3 | 18.4
2007-05-04 | 20:45 | 5.4 | 7.5 | 12.23 | 96.8 | 28.3 | 18.4
2007-05-04 | 21:00 | 5.23 | 7.5 | 12.28 | 96.7 | 28.4 | 18.4
2007-05-04 | 21:15 | 5.07 | 7.49 | 12.34 | 96.8 | 28.3 | 18.4
2007-05-04 | 21:30 | 4.92 | 7.49 | 12.39 | 96.8 | 28.3 | 18.4
2007-05-04 | 21:45 | 4.77 | 7.49 | 12.44 | 96.8 | 28.3 | 18.4
2007-05-04 | 22:00 | 4.6 | 7.49 | 12.5 | 96.9 | 28.3 | 18.4
2007-05-04 | 22:15 | 4.4 | 7.48 | 12.57 | 96.9 | 28.3 | 18.4
2007-05-04 | 22:30 | 4.18 | 7.48 | 12.64 | 96.9 | 28.3 | 18.4
2007-05-04 | 22:45 | 3.98 | 7.47 | 12.7 | 96.8 | 28.3 | 18.4
2007-05-04 | 23:00 | 3.83 | 7.47 | 12.76 | 96.9 | 28.3 | 18.4
2007-05-04 | 23:15 | 3.72 | 7.47 | 12.79 | 96.8 | 28.3 | 18.4
2007-05-04 | 23:30 | 3.61 | 7.46 | 12.82 | 96.8 | 28.3 | 18.4
2007-05-04 | 23:45 | 3.51 | 7.46 | 12.87 | 96.9 | 28.4 | 18.4
2007-05-05 | 0:00 | 3.4 | 7.46 | 12.91 | 96.9 | 28.4 | 18.4
2007-05-05 | 0:15 | 3.3 | 7.46 | 12.95 | 97 | 28.4 | 18.4
2007-05-05 | 0:30 | 3.2 | 7.46 | 12.93 | 96.6 | 28.4 | 18.5
2007-05-05 | 0:45 | 3.11 | 7.46 | 12.99 | 96.8 | 28.4 | 18.5
2007-05-05 | 1:01 | 3.02 | 7.46 | 13.02 | 96.8 | 28.4 | 18.5
2007-05-05 | 1:16 | 2.94 | 7.46 | 13.07 | 96.9 | 28.5 | 18.5
2007-05-05 | 1:31 | 2.86 | 7.46 | 13.11 | 97 | 28.5 | 18.5
2007-05-05 | 1:46 | 2.78 | 7.46 | 13.07 | 96.5 | 28.5 | 18.5
2007-05-05 | 2:01 | 2.69 | 7.46 | 13.14 | 96.8 | 28.6 | 18.6
2007-05-05 | 2:16 | 2.6 | 7.46 | 13.18 | 96.9 | 28.6 | 18.6
2007-05-05 | 2:31 | 2.53 | 7.46 | 13.22 | 97 | 28.6 | 18.6
2007-05-05 | 2:46 | 2.45 | 7.46 | 13.25 | 97 | 28.6 | 18.6
2007-05-05 | 3:01 | 2.39 | 7.46 | 13.26 | 96.9 | 28.6 | 18.6
2007-05-05 | 3:16 | 2.34 | 7.46 | 13.3 | 97 | 28.7 | 18.6
2007-05-05 | 3:31 | 2.28 | 7.46 | 13.32 | 97 | 28.7 | 18.6
2007-05-05 | 3:46 | 2.22 | 7.46 | 13.34 | 97.1 | 28.7 | 18.7
2007-05-05 | 4:01 | 2.16 | 7.45 | 13.39 | 97.2 | 28.8 | 18.7
2007-05-05 | 4:16 | 2.09 | 7.45 | 13.4 | 97.2 | 28.8 | 18.7
2007-05-05 | 4:31 | 2.02 | 7.45 | 13.43 | 97.1 | 28.8 | 18.7
2007-05-05 | 4:46 | 1.96 | 7.45 | 13.45 | 97.1 | 28.8 | 18.7
2007-05-05 | 5:01 | 1.89 | 7.45 | 13.48 | 97.1 | 28.8 | 18.7
2007-05-05 | 5:16 | 1.82 | 7.45 | 13.49 | 97.1 | 28.9 | 18.8
2007-05-05 | 5:31 | 1.76 | 7.45 | 13.52 | 97.1 | 28.8 | 18.7
2007-05-05 | 5:46 | 1.69 | 7.44 | 13.54 | 97.1 | 28.9 | 18.8
2007-05-05 | 6:01 | 1.63 | 7.44 | 13.56 | 97 | 28.9 | 18.8
2007-05-05 | 6:16 | 1.55 | 7.45 | 13.6 | 97.2 | 28.9 | 18.8
2007-05-05 | 6:31 | 1.5 | 7.45 | 13.63 | 97.2 | 28.9 | 18.8
2007-05-05 | 6:46 | 1.45 | 7.44 | 13.64 | 97.1 | 28.9 | 18.8
2007-05-05 | 7:01 | 1.4 | 7.44 | 13.61 | 96.8 | 28.9 | 18.8
2007-05-05 | 7:16 | 1.36 | 7.45 | 13.68 | 97.2 | 29 | 18.8
2007-05-05 | 7:31 | 1.32 | 7.45 | 13.72 | 97.4 | 29 | 18.8
2007-05-05 | 7:46 | 1.29 | 7.45 | 13.72 | 97.3 | 29 | 18.9
2007-05-05 | 8:01 | 1.29 | 7.45 | 13.74 | 97.4 | 29 | 18.8
2007-05-05 | 8:16 | 1.3 | 7.45 | 13.74 | 97.5 | 29 | 18.8
2007-05-05 | 8:31 | 1.3 | 7.46 | 13.74 | 97.5 | 29 | 18.9
2007-05-05 | 8:46 | 1.31 | 7.47 | 13.76 | 97.6 | 29 | 18.9
2007-05-05 | 9:01 | 1.35 | 7.47 | 13.73 | 97.5 | 29 | 18.9
2007-05-05 | 9:16 | 1.4 | 7.47 | 13.73 | 97.6 | 29 | 18.9
2007-05-05 | 9:31 | 1.46 | 7.47 | 13.7 | 97.6 | 29 | 18.9
2007-05-05 | 9:46 | 1.54 | 7.48 | 13.68 | 97.7 | 29 | 18.9
2007-05-05 | 10:01 | 1.63 | 7.48 | 13.66 | 97.8 | 29 | 18.9
2007-05-05 | 10:16 | 1.73 | 7.48 | 13.62 | 97.8 | 29 | 18.9
2007-05-05 | 10:31 | 1.85 | 7.49 | 13.59 | 97.9 | 29.1 | 18.9
2007-05-05 | 10:46 | 1.99 | 7.49 | 13.53 | 97.8 | 29.1 | 18.9
2007-05-05 | 11:01 | 2.14 | 7.49 | 13.5 | 98 | 29.1 | 18.9
2007-05-05 | 11:16 | 2.28 | 7.49 | 13.46 | 98 | 29.1 | 18.9
2007-05-05 | 11:31 | 2.41 | 7.49 | 13.39 | 97.9 | 29.1 | 18.9
2007-05-05 | 11:46 | 2.59 | 7.49 | 13.37 | 98.2 | 29.1 | 18.9
2007-05-05 | 12:01 | 2.79 | 7.5 | 13.29 | 98.1 | 29.1 | 18.9
2007-05-05 | 12:16 | 3.02 | 7.5 | 13.19 | 98 | 29.1 | 18.9
2007-05-05 | 12:31 | 3.23 | 7.51 | 13.14 | 98.2 | 29.1 | 18.9
2007-05-05 | 12:46 | 3.5 | 7.51 | 13.06 | 98.3 | 29.1 | 18.9
2007-05-05 | 13:00 | 3.81 | 7.52 | 12.94 | 98.3 | 29.1 | 18.9
2007-05-05 | 13:15 | 4.11 | 7.52 | 12.84 | 98.2 | 29.1 | 18.9
2007-05-05 | 13:30 | 4.41 | 7.52 | 12.75 | 98.3 | 29 | 18.9
2007-05-05 | 13:45 | 4.73 | 7.53 | 12.65 | 98.4 | 29 | 18.9
2007-05-05 | 14:00 | 5.04 | 7.53 | 12.55 | 98.4 | 29 | 18.9
2007-05-05 | 14:15 | 5.28 | 7.53 | 12.48 | 98.4 | 29 | 18.8
2007-05-05 | 14:30 | 5.54 | 7.54 | 12.39 | 98.4 | 29 | 18.8
2007-05-05 | 14:45 | 5.82 | 7.53 | 12.3 | 98.3 | 28.9 | 18.8
2007-05-05 | 15:00 | 6.07 | 7.54 | 12.23 | 98.4 | 28.9 | 18.8
2007-05-05 | 15:15 | 6.3 | 7.55 | 12.15 | 98.4 | 28.9 | 18.8
2007-05-05 | 15:30 | 6.53 | 7.56 | 12.09 | 98.5 | 28.9 | 18.8
2007-05-05 | 15:45 | 6.71 | 7.55 | 12.02 | 98.3 | 28.8 | 18.7
2007-05-05 | 16:00 | 6.83 | 7.56 | 11.97 | 98.2 | 28.9 | 18.8
2007-05-05 | 16:15 | 6.92 | 7.56 | 11.94 | 98.1 | 28.9 | 18.8
2007-05-05 | 16:30 | 7.05 | 7.56 | 11.91 | 98.2 | 28.9 | 18.8
2007-05-05 | 16:45 | 7.19 | 7.56 | 11.86 | 98.2 | 28.9 | 18.8
2007-05-05 | 17:00 | 7.29 | 7.56 | 11.83 | 98.2 | 28.9 | 18.8
2007-05-05 | 17:15 | 7.35 | 7.56 | 11.81 | 98.1 | 28.9 | 18.8
2007-05-05 | 17:30 | 7.4 | 7.56 | 11.79 | 98.1 | 28.9 | 18.8
2007-05-05 | 17:45 | 7.41 | 7.56 | 11.79 | 98.1 | 28.9 | 18.8
2007-05-05 | 18:00 | 7.39 | 7.56 | 11.79 | 98.1 | 28.8 | 18.7
2007-05-05 | 18:15 | 7.34 | 7.56 | 11.8 | 98.1 | 28.8 | 18.7
2007-05-05 | 18:30 | 7.3 | 7.56 | 11.8 | 97.9 | 28.8 | 18.7
2007-05-05 | 18:45 | 7.22 | 7.55 | 11.81 | 97.8 | 28.8 | 18.7
2007-05-05 | 19:00 | 7.17 | 7.55 | 11.83 | 97.8 | 28.8 | 18.7
2007-05-05 | 19:15 | 7.1 | 7.54 | 11.84 | 97.8 | 28.8 | 18.7
2007-05-05 | 19:30 | 7.02 | 7.54 | 11.86 | 97.7 | 28.8 | 18.7
2007-05-05 | 19:45 | 6.91 | 7.54 | 11.89 | 97.8 | 28.8 | 18.7
2007-05-05 | 20:00 | 6.77 | 7.53 | 11.93 | 97.7 | 28.8 | 18.7
2007-05-05 | 20:15 | 6.6 | 7.53 | 11.98 | 97.7 | 28.8 | 18.7
2007-05-05 | 20:30 | 6.44 | 7.53 | 12.02 | 97.6 | 28.8 | 18.7
2007-05-05 | 20:45 | 6.28 | 7.52 | 12.06 | 97.6 | 28.8 | 18.7
2007-05-05 | 21:00 | 6.13 | 7.52 | 12.13 | 97.7 | 28.8 | 18.7
2007-05-05 | 21:15 | 5.99 | 7.52 | 12.18 | 97.8 | 28.8 | 18.7
2007-05-05 | 21:30 | 5.84 | 7.52 | 12.22 | 97.8 | 28.8 | 18.7
2007-05-05 | 21:45 | 5.69 | 7.51 | 12.27 | 97.8 | 28.8 | 18.7
2007-05-05 | 22:00 | 5.55 | 7.51 | 12.31 | 97.8 | 28.7 | 18.7
2007-05-05 | 22:15 | 5.42 | 7.5 | 12.35 | 97.8 | 28.7 | 18.7
2007-05-05 | 22:30 | 5.3 | 7.5 | 12.39 | 97.8 | 28.7 | 18.7
2007-05-05 | 22:45 | 5.19 | 7.49 | 12.44 | 97.9 | 28.8 | 18.7
2007-05-05 | 23:00 | 5.08 | 7.49 | 12.47 | 97.8 | 28.7 | 18.7
2007-05-05 | 23:15 | 4.96 | 7.49 | 12.5 | 97.8 | 28.7 | 18.7
2007-05-05 | 23:30 | 4.85 | 7.48 | 12.53 | 97.7 | 28.7 | 18.7
2007-05-05 | 23:45 | 4.73 | 7.48 | 12.57 | 97.7 | 28.7 | 18.7
2007-05-06 | 0:00 | 4.61 | 7.48 | 12.63 | 97.9 | 28.8 | 18.7
2007-05-06 | 0:15 | 4.49 | 7.48 | 12.66 | 97.8 | 28.8 | 18.7
2007-05-06 | 0:30 | 4.37 | 7.48 | 12.72 | 97.9 | 28.8 | 18.7
2007-05-06 | 0:45 | 4.25 | 7.48 | 12.74 | 97.9 | 28.7 | 18.7
2007-05-06 | 1:01 | 4.15 | 7.48 | 12.77 | 97.8 | 28.8 | 18.7
2007-05-06 | 1:16 | 4.06 | 7.48 | 12.8 | 97.8 | 28.8 | 18.7
2007-05-06 | 1:31 | 3.97 | 7.47 | 12.83 | 97.8 | 28.8 | 18.7
2007-05-06 | 1:46 | 3.89 | 7.47 | 12.86 | 97.8 | 28.8 | 18.7
2007-05-06 | 2:01 | 3.82 | 7.47 | 12.88 | 97.8 | 28.8 | 18.7
2007-05-06 | 2:16 | 3.76 | 7.47 | 12.9 | 97.8 | 28.8 | 18.7
2007-05-06 | 2:31 | 3.69 | 7.47 | 12.93 | 97.9 | 28.9 | 18.8
2007-05-06 | 2:46 | 3.62 | 7.47 | 12.95 | 97.8 | 28.8 | 18.7
2007-05-06 | 3:01 | 3.55 | 7.47 | 12.99 | 97.9 | 28.9 | 18.8
2007-05-06 | 3:16 | 3.49 | 7.46 | 13 | 97.8 | 28.9 | 18.8
2007-05-06 | 3:31 | 3.43 | 7.46 | 13.02 | 97.8 | 28.9 | 18.8
2007-05-06 | 3:46 | 3.36 | 7.46 | 13.05 | 97.8 | 28.9 | 18.8
2007-05-06 | 4:01 | 3.3 | 7.46 | 13.06 | 97.8 | 28.9 | 18.8
2007-05-06 | 4:16 | 3.25 | 7.46 | 13.08 | 97.8 | 28.9 | 18.8
2007-05-06 | 4:31 | 3.21 | 7.46 | 13.11 | 97.9 | 29 | 18.8
2007-05-06 | 4:46 | 3.16 | 7.45 | 13.14 | 98 | 29 | 18.8
2007-05-06 | 5:01 | 3.11 | 7.45 | 13.13 | 97.8 | 29 | 18.8
2007-05-06 | 5:16 | 3.07 | 7.46 | 13.15 | 97.8 | 29 | 18.8
2007-05-06 | 5:31 | 3.04 | 7.45 | 13.16 | 97.8 | 29 | 18.9
2007-05-06 | 5:46 | 3 | 7.45 | 13.17 | 97.8 | 29 | 18.9
2007-05-06 | 6:01 | 2.97 | 7.45 | 13.2 | 98 | 29.1 | 18.9
2007-05-06 | 6:16 | 2.95 | 7.45 | 13.2 | 97.9 | 29.1 | 18.9
2007-05-06 | 6:31 | 2.94 | 7.45 | 13.2 | 97.9 | 29 | 18.9
2007-05-06 | 6:46 | 2.93 | 7.45 | 13.19 | 97.8 | 29.1 | 18.9
2007-05-06 | 7:01 | 2.92 | 7.45 | 13.24 | 98.2 | 29.1 | 18.9
2007-05-06 | 7:16 | 2.93 | 7.45 | 13.22 | 98 | 29.1 | 18.9
2007-05-06 | 7:31 | 2.96 | 7.46 | 13.22 | 98.1 | 29.1 | 18.9
2007-05-06 | 7:46 | 2.99 | 7.46 | 13.19 | 98 | 29.1 | 18.9
2007-05-06 | 8:01 | 3.01 | 7.47 | 13.22 | 98.2 | 29.1 | 18.9
2007-05-06 | 8:16 | 3.09 | 7.47 | 13.19 | 98.2 | 29.1 | 18.9
2007-05-06 | 8:31 | 3.13 | 7.48 | 13.19 | 98.3 | 29.1 | 18.9
2007-05-06 | 8:46 | 3.21 | 7.49 | 13.16 | 98.3 | 29.1 | 18.9
2007-05-06 | 9:01 | 3.34 | 7.5 | 13.12 | 98.3 | 29.1 | 18.9
2007-05-06 | 9:16 | 3.41 | 7.5 | 13.1 | 98.4 | 29.1 | 18.9
2007-05-06 | 9:31 | 3.5 | 7.5 | 13.06 | 98.4 | 29.1 | 18.9
2007-05-06 | 9:46 | 3.58 | 7.5 | 13.04 | 98.4 | 29.1 | 18.9
2007-05-06 | 10:01 | 3.76 | 7.51 | 12.98 | 98.4 | 29.1 | 18.9
2007-05-06 | 10:16 | 3.91 | 7.52 | 12.93 | 98.4 | 29.1 | 18.9
2007-05-06 | 10:31 | 4.12 | 7.51 | 12.86 | 98.4 | 29.1 | 18.9
2007-05-06 | 10:46 | 4.35 | 7.52 | 12.8 | 98.5 | 29.1 | 18.9
2007-05-06 | 11:01 | 4.58 | 7.53 | 12.74 | 98.7 | 29.1 | 18.9
2007-05-06 | 11:16 | 4.82 | 7.53 | 12.64 | 98.5 | 29.1 | 18.9
2007-05-06 | 11:31 | 5.1 | 7.54 | 12.57 | 98.6 | 29.1 | 18.9
2007-05-06 | 11:46 | 5.35 | 7.54 | 12.48 | 98.6 | 29.1 | 18.9
2007-05-06 | 12:01 | 5.65 | 7.54 | 12.41 | 98.8 | 29.1 | 18.9
2007-05-06 | 12:16 | 5.97 | 7.55 | 12.31 | 98.9 | 29.1 | 18.9
2007-05-06 | 12:31 | 6.25 | 7.55 | 12.23 | 98.9 | 29.1 | 18.9
2007-05-06 | 12:46 | 6.52 | 7.56 | 12.14 | 98.8 | 29.1 | 18.9
2007-05-06 | 13:00 | 6.96 | 7.56 | 12.03 | 99 | 29 | 18.9
2007-05-06 | 13:15 | 7.32 | 7.56 | 11.92 | 99 | 29 | 18.9
2007-05-06 | 13:30 | 7.69 | 7.57 | 11.81 | 98.9 | 29 | 18.9
2007-05-06 | 13:45 | 8.05 | 7.57 | 11.71 | 99 | 29 | 18.9
2007-05-06 | 14:00 | 8.42 | 7.58 | 11.61 | 99 | 29 | 18.9
2007-05-06 | 14:15 | 8.77 | 7.58 | 11.51 | 99 | 29 | 18.9
2007-05-06 | 14:30 | 9.11 | 7.59 | 11.39 | 98.8 | 29 | 18.9
2007-05-06 | 14:45 | 9.4 | 7.6 | 11.33 | 99 | 29 | 18.9
2007-05-06 | 15:00 | 9.7 | 7.6 | 11.25 | 99 | 29 | 18.9
2007-05-06 | 15:15 | 9.93 | 7.6 | 11.18 | 98.9 | 29 | 18.9
2007-05-06 | 15:30 | 10.14 | 7.6 | 11.13 | 98.9 | 29 | 18.8
2007-05-06 | 15:45 | 10.16 | 7.6 | 11.09 | 98.6 | 29 | 18.8
2007-05-06 | 16:00 | 10.34 | 7.6 | 11.06 | 98.8 | 29 | 18.8
2007-05-06 | 16:15 | 10.51 | 7.6 | 11.01 | 98.7 | 29 | 18.8
2007-05-06 | 16:30 | 10.7 | 7.61 | 10.96 | 98.7 | 29 | 18.8
2007-05-06 | 16:45 | 10.86 | 7.59 | 10.92 | 98.7 | 29 | 18.8
2007-05-06 | 17:00 | 11 | 7.6 | 10.89 | 98.8 | 28.9 | 18.8
2007-05-06 | 17:15 | 11.12 | 7.61 | 10.86 | 98.7 | 28.9 | 18.8
2007-05-06 | 17:30 | 11.19 | 7.6 | 10.85 | 98.8 | 28.9 | 18.8
2007-05-06 | 17:45 | 11.23 | 7.61 | 10.82 | 98.7 | 28.9 | 18.8
2007-05-06 | 18:00 | 11.22 | 7.61 | 10.82 | 98.6 | 28.8 | 18.7
2007-05-06 | 18:15 | 11.15 | 7.6 | 10.83 | 98.5 | 28.8 | 18.7
2007-05-06 | 18:30 | 11.03 | 7.6 | 10.85 | 98.4 | 28.8 | 18.7
2007-05-06 | 18:45 | 10.93 | 7.59 | 10.86 | 98.3 | 28.7 | 18.7
2007-05-06 | 19:00 | 10.84 | 7.59 | 10.88 | 98.3 | 28.7 | 18.6
2007-05-06 | 19:15 | 10.72 | 7.58 | 10.92 | 98.4 | 28.6 | 18.6
2007-05-06 | 19:30 | 10.59 | 7.59 | 10.95 | 98.3 | 28.5 | 18.5
2007-05-06 | 19:45 | 10.42 | 7.57 | 10.99 | 98.3 | 28.5 | 18.5
2007-05-06 | 20:00 | 10.23 | 7.57 | 11.04 | 98.4 | 28.3 | 18.4
2007-05-06 | 20:15 | 10.04 | 7.56 | 11.1 | 98.4 | 28.2 | 18.3
2007-05-06 | 20:30 | 9.84 | 7.55 | 11.13 | 98.2 | 28.2 | 18.3
2007-05-06 | 20:45 | 9.65 | 7.55 | 11.19 | 98.3 | 28 | 18.2
2007-05-06 | 21:00 | 9.46 | 7.54 | 11.25 | 98.4 | 27.9 | 18.1
2007-05-06 | 21:15 | 9.28 | 7.53 | 11.29 | 98.3 | 27.7 | 18
2007-05-06 | 21:30 | 9.09 | 7.52 | 11.33 | 98.2 | 27.6 | 18
2007-05-06 | 21:45 | 8.91 | 7.51 | 11.38 | 98.2 | 27.5 | 17.9
2007-05-06 | 22:00 | 8.73 | 7.51 | 11.43 | 98.3 | 27.4 | 17.8
2007-05-06 | 22:15 | 8.56 | 7.5 | 11.48 | 98.2 | 27.3 | 17.7
2007-05-06 | 22:30 | 8.39 | 7.5 | 11.52 | 98.2 | 27.2 | 17.7
2007-05-06 | 22:45 | 8.21 | 7.5 | 11.58 | 98.3 | 27.1 | 17.6
2007-05-06 | 23:00 | 8.04 | 7.48 | 11.64 | 98.4 | 27 | 17.6
2007-05-06 | 23:15 | 7.87 | 7.48 | 11.68 | 98.3 | 26.9 | 17.5
2007-05-06 | 23:30 | 7.71 | 7.47 | 11.73 | 98.3 | 26.9 | 17.5
2007-05-06 | 23:45 | 7.55 | 7.47 | 11.77 | 98.3 | 26.8 | 17.4
2007-05-07 | 0:00 | 7.4 | 7.47 | 11.81 | 98.3 | 26.8 | 17.4
2007-05-07 | 0:15 | 7.25 | 7.47 | 11.86 | 98.3 | 26.8 | 17.4
2007-05-07 | 0:30 | 7.1 | 7.46 | 11.9 | 98.3 | 26.8 | 17.4
2007-05-07 | 0:45 | 6.94 | 7.46 | 11.95 | 98.3 | 26.8 | 17.4
2007-05-07 | 1:01 | 6.79 | 7.44 | 11.97 | 98.1 | 26.7 | 17.4
2007-05-07 | 1:16 | 6.63 | 7.45 | 12.04 | 98.3 | 26.7 | 17.3
2007-05-07 | 1:31 | 6.48 | 7.45 | 12.09 | 98.3 | 26.7 | 17.3
2007-05-07 | 1:46 | 6.33 | 7.44 | 12.14 | 98.3 | 26.7 | 17.4
2007-05-07 | 2:01 | 6.17 | 7.45 | 12.18 | 98.3 | 26.7 | 17.4
2007-05-07 | 2:16 | 6.01 | 7.45 | 12.21 | 98.1 | 26.7 | 17.4
2007-05-07 | 2:31 | 5.85 | 7.44 | 12.27 | 98.2 | 26.7 | 17.4
2007-05-07 | 2:46 | 5.69 | 7.44 | 12.3 | 98 | 26.7 | 17.4
2007-05-07 | 3:01 | 5.53 | 7.44 | 12.37 | 98.2 | 26.8 | 17.4
2007-05-07 | 3:16 | 5.38 | 7.43 | 12.42 | 98.2 | 26.8 | 17.4
2007-05-07 | 3:31 | 5.24 | 7.43 | 12.46 | 98.2 | 26.8 | 17.4
2007-05-07 | 3:46 | 5.1 | 7.43 | 12.5 | 98.2 | 26.9 | 17.5
2007-05-07 | 4:01 | 4.97 | 7.43 | 12.55 | 98.2 | 27 | 17.5
2007-05-07 | 4:16 | 4.86 | 7.43 | 12.58 | 98.1 | 27 | 17.6
2007-05-07 | 4:31 | 4.75 | 7.43 | 12.61 | 98.1 | 27.1 | 17.6
2007-05-07 | 4:46 | 4.65 | 7.43 | 12.65 | 98.2 | 27.1 | 17.6
2007-05-07 | 5:01 | 4.55 | 7.43 | 12.69 | 98.2 | 27.1 | 17.6
2007-05-07 | 5:16 | 4.45 | 7.43 | 12.72 | 98.2 | 27.1 | 17.6
2007-05-07 | 5:31 | 4.37 | 7.43 | 12.74 | 98.1 | 27.2 | 17.7
2007-05-07 | 5:46 | 4.29 | 7.42 | 12.77 | 98.2 | 27.4 | 17.8
2007-05-07 | 6:01 | 4.22 | 7.42 | 12.81 | 98.3 | 27.2 | 17.7
2007-05-07 | 6:16 | 4.15 | 7.43 | 12.82 | 98.2 | 27.2 | 17.7
2007-05-07 | 6:31 | 4.09 | 7.43 | 12.85 | 98.3 | 27.3 | 17.7
2007-05-07 | 6:46 | 4.03 | 7.43 | 12.87 | 98.3 | 27.3 | 17.7
2007-05-07 | 7:01 | 3.98 | 7.43 | 12.88 | 98.2 | 27.3 | 17.7
2007-05-07 | 7:16 | 3.94 | 7.43 | 12.91 | 98.3 | 27.4 | 17.8
2007-05-07 | 7:31 | 3.91 | 7.44 | 12.91 | 98.3 | 27.3 | 17.8
2007-05-07 | 7:46 | 3.9 | 7.44 | 12.88 | 98 | 27.4 | 17.8
2007-05-07 | 8:01 | 3.9 | 7.44 | 12.9 | 98.1 | 27.5 | 17.9
2007-05-07 | 8:16 | 3.9 | 7.45 | 12.9 | 98.2 | 27.6 | 17.9
2007-05-07 | 8:31 | 3.92 | 7.45 | 12.91 | 98.3 | 27.6 | 18
2007-05-07 | 8:46 | 3.94 | 7.47 | 12.92 | 98.4 | 27.7 | 18
2007-05-07 | 9:01 | 3.99 | 7.47 | 12.91 | 98.4 | 27.7 | 18
2007-05-07 | 9:16 | 4.06 | 7.48 | 12.88 | 98.4 | 27.9 | 18.2
2007-05-07 | 9:31 | 4.15 | 7.48 | 12.86 | 98.5 | 28 | 18.2
2007-05-07 | 9:46 | 4.25 | 7.49 | 12.83 | 98.5 | 28 | 18.2
2007-05-07 | 10:01 | 4.38 | 7.49 | 12.79 | 98.6 | 28.1 | 18.3
2007-05-07 | 10:16 | 4.53 | 7.5 | 12.75 | 98.6 | 28.2 | 18.3
2007-05-07 | 10:31 | 4.68 | 7.5 | 12.72 | 98.7 | 28.2 | 18.3
2007-05-07 | 10:46 | 4.82 | 7.5 | 12.64 | 98.5 | 28.2 | 18.3
2007-05-07 | 11:01 | 4.99 | 7.5 | 12.6 | 98.7 | 28.2 | 18.3
2007-05-07 | 11:16 | 5.15 | 7.51 | 12.55 | 98.6 | 28.2 | 18.4
2007-05-07 | 11:31 | 5.32 | 7.51 | 12.5 | 98.7 | 28.3 | 18.4
2007-05-07 | 11:46 | 5.53 | 7.51 | 12.43 | 98.7 | 28.3 | 18.4
2007-05-07 | 12:01 | 5.77 | 7.51 | 12.37 | 98.8 | 28.3 | 18.4
2007-05-07 | 12:16 | 6.05 | 7.52 | 12.3 | 98.9 | 28.3 | 18.4
2007-05-07 | 12:31 | 6.33 | 7.53 | 12.21 | 98.9 | 28.4 | 18.4
2007-05-07 | 12:46 | 6.63 | 7.52 | 12.12 | 98.9 | 28.5 | 18.5
2007-05-07 | 13:00 | 6.93 | 7.53 | 12.03 | 98.9 | 27.7 | 18
2007-05-07 | 13:15 | 7.22 | 7.53 | 11.93 | 98.8 | 27.7 | 18
2007-05-07 | 13:30 | 7.54 | 7.54 | 11.85 | 98.9 | 27.7 | 18
2007-05-07 | 13:45 | 7.85 | 7.54 | 11.75 | 98.9 | 27.8 | 18
2007-05-07 | 14:00 | 8.17 | 7.55 | 11.68 | 99 | 27.7 | 18
2007-05-07 | 14:15 | 8.48 | 7.55 | 11.58 | 98.9 | 27.8 | 18.1
2007-05-07 | 14:30 | 8.8 | 7.56 | 11.48 | 98.9 | 27.8 | 18.1
2007-05-07 | 14:45 | 9.12 | 7.56 | 11.4 | 98.9 | 27.8 | 18.1
2007-05-07 | 15:00 | 9.42 | 7.57 | 11.29 | 98.6 | 27.8 | 18.1
2007-05-07 | 15:15 | 9.72 | 7.58 | 11.22 | 98.8 | 27.8 | 18.1
2007-05-07 | 15:30 | 9.98 | 7.58 | 11.13 | 98.6 | 27.9 | 18.1
2007-05-07 | 15:45 | 10.21 | 7.57 | 11.08 | 98.7 | 27.8 | 18.1
2007-05-07 | 16:00 | 10.41 | 7.59 | 11.02 | 98.6 | 27.8 | 18.1
2007-05-07 | 16:15 | 10.6 | 7.58 | 10.97 | 98.5 | 27.8 | 18.1
2007-05-07 | 16:30 | 10.76 | 7.58 | 10.93 | 98.5 | 27.7 | 18
2007-05-07 | 16:45 | 10.89 | 7.59 | 10.89 | 98.5 | 27.7 | 18
2007-05-07 | 17:00 | 10.99 | 7.58 | 10.86 | 98.5 | 27.7 | 18
2007-05-07 | 17:15 | 11.08 | 7.58 | 10.84 | 98.5 | 27.7 | 18
2007-05-07 | 17:30 | 11.14 | 7.59 | 10.81 | 98.4 | 27.7 | 18
2007-05-07 | 17:45 | 11.18 | 7.58 | 10.8 | 98.3 | 27.7 | 18
2007-05-07 | 18:00 | 11.2 | 7.57 | 10.79 | 98.3 | 27.7 | 18
2007-05-07 | 18:15 | 11.19 | 7.58 | 10.77 | 98.1 | 27.7 | 18
2007-05-07 | 18:30 | 11.16 | 7.56 | 10.78 | 98.1 | 27.6 | 17.9
2007-05-07 | 18:45 | 11.1 | 7.57 | 10.78 | 97.9 | 27.6 | 17.9
2007-05-07 | 19:00 | 10.99 | 7.56 | 10.81 | 98 | 27.5 | 17.9
2007-05-07 | 19:15 | 10.85 | 7.55 | 10.85 | 98 | 27.5 | 17.9
2007-05-07 | 19:30 | 10.66 | 7.54 | 10.88 | 97.9 | 27.4 | 17.8
2007-05-07 | 19:45 | 10.44 | 7.53 | 10.94 | 97.9 | 27.4 | 17.8
2007-05-07 | 20:00 | 10.2 | 7.52 | 11 | 97.9 | 27.3 | 17.7
2007-05-07 | 20:15 | 9.96 | 7.5 | 11.06 | 97.9 | 27.2 | 17.7
2007-05-07 | 20:30 | 9.73 | 7.5 | 11.12 | 97.9 | 27.1 | 17.6
2007-05-07 | 20:45 | 9.51 | 7.5 | 11.18 | 97.9 | 27 | 17.5
2007-05-07 | 21:00 | 9.3 | 7.49 | 11.26 | 98.1 | 26.9 | 17.5
2007-05-07 | 21:15 | 9.1 | 7.48 | 11.29 | 97.9 | 26.8 | 17.4
2007-05-07 | 21:30 | 8.9 | 7.48 | 11.34 | 97.8 | 26.6 | 17.3
2007-05-07 | 21:45 | 8.7 | 7.47 | 11.4 | 97.9 | 26.6 | 17.3
2007-05-07 | 22:00 | 8.51 | 7.46 | 11.44 | 97.8 | 26.5 | 17.2
2007-05-07 | 22:15 | 8.32 | 7.46 | 11.5 | 97.9 | 26.4 | 17.1
2007-05-07 | 22:30 | 8.12 | 7.45 | 11.57 | 98 | 26.3 | 17.1
2007-05-07 | 22:45 | 7.94 | 7.45 | 11.62 | 98 | 26.2 | 17
2007-05-07 | 23:00 | 7.77 | 7.45 | 11.66 | 97.9 | 26.1 | 17
2007-05-07 | 23:15 | 7.6 | 7.44 | 11.7 | 97.8 | 26 | 16.9
2007-05-07 | 23:30 | 7.43 | 7.44 | 11.75 | 97.8 | 26 | 16.9
2007-05-07 | 23:45 | 7.26 | 7.43 | 11.81 | 97.9 | 25.9 | 16.8
2007-05-08 | 0:00 | 7.1 | 7.43 | 11.85 | 97.8 | 25.9 | 16.8
2007-05-08 | 0:15 | 6.93 | 7.44 | 11.91 | 97.9 | 25.8 | 16.8
2007-05-08 | 0:30 | 6.77 | 7.43 | 11.94 | 97.8 | 25.8 | 16.8
2007-05-08 | 0:45 | 6.6 | 7.43 | 11.99 | 97.8 | 25.8 | 16.8
2007-05-08 | 1:01 | 6.44 | 7.42 | 12.03 | 97.7 | 25.8 | 16.8
2007-05-08 | 1:16 | 6.3 | 7.42 | 12.08 | 97.7 | 25.8 | 16.7
2007-05-08 | 1:31 | 6.16 | 7.42 | 12.11 | 97.7 | 25.8 | 16.7
2007-05-08 | 1:46 | 6.03 | 7.42 | 12.15 | 97.7 | 25.8 | 16.7
2007-05-08 | 2:01 | 5.89 | 7.42 | 12.19 | 97.7 | 25.8 | 16.8
2007-05-08 | 2:16 | 5.76 | 7.42 | 12.23 | 97.7 | 25.8 | 16.8
2007-05-08 | 2:31 | 5.63 | 7.42 | 12.27 | 97.7 | 25.8 | 16.8
2007-05-08 | 2:46 | 5.52 | 7.42 | 12.31 | 97.7 | 25.8 | 16.8
2007-05-08 | 3:01 | 5.41 | 7.42 | 12.35 | 97.7 | 25.8 | 16.8
2007-05-08 | 3:16 | 5.3 | 7.42 | 12.38 | 97.7 | 25.8 | 16.8
2007-05-08 | 3:31 | 5.2 | 7.42 | 12.4 | 97.6 | 25.8 | 16.8
2007-05-08 | 3:46 | 5.1 | 7.42 | 12.44 | 97.7 | 25.9 | 16.8
2007-05-08 | 4:01 | 5.01 | 7.42 | 12.46 | 97.6 | 25.9 | 16.8
2007-05-08 | 4:16 | 4.93 | 7.41 | 12.48 | 97.5 | 25.9 | 16.9
2007-05-08 | 4:31 | 4.85 | 7.41 | 12.51 | 97.5 | 25.9 | 16.9
2007-05-08 | 4:46 | 4.77 | 7.41 | 12.52 | 97.5 | 25.9 | 16.9
2007-05-08 | 5:01 | 4.7 | 7.42 | 12.53 | 97.4 | 26 | 16.9
2007-05-08 | 5:16 | 4.63 | 7.42 | 12.57 | 97.5 | 26 | 16.9
2007-05-08 | 5:31 | 4.57 | 7.41 | 12.58 | 97.4 | 26 | 16.9
2007-05-08 | 5:46 | 4.51 | 7.41 | 12.6 | 97.4 | 26 | 16.9
2007-05-08 | 6:01 | 4.45 | 7.41 | 12.62 | 97.4 | 26.1 | 16.9
2007-05-08 | 6:16 | 4.39 | 7.42 | 12.64 | 97.4 | 26.1 | 17
2007-05-08 | 6:31 | 4.35 | 7.42 | 12.66 | 97.4 | 26.1 | 17
2007-05-08 | 6:46 | 4.31 | 7.42 | 12.67 | 97.4 | 26.2 | 17
2007-05-08 | 7:01 | 4.27 | 7.42 | 12.68 | 97.5 | 26.1 | 17
2007-05-08 | 7:16 | 4.25 | 7.43 | 12.69 | 97.4 | 26.2 | 17
2007-05-08 | 7:31 | 4.23 | 7.43 | 12.68 | 97.3 | 26.2 | 17
2007-05-08 | 7:46 | 4.22 | 7.44 | 12.72 | 97.6 | 26.2 | 17.1
2007-05-08 | 8:01 | 4.22 | 7.43 | 12.7 | 97.4 | 26.3 | 17.1
2007-05-08 | 8:16 | 4.22 | 7.44 | 12.7 | 97.4 | 26.3 | 17.1
2007-05-08 | 8:31 | 4.24 | 7.44 | 12.7 | 97.5 | 26.3 | 17.1
2007-05-08 | 8:46 | 4.27 | 7.45 | 12.7 | 97.6 | 26.4 | 17.1
2007-05-08 | 9:01 | 4.32 | 7.46 | 12.67 | 97.5 | 26.4 | 17.1
2007-05-08 | 9:16 | 4.4 | 7.46 | 12.65 | 97.5 | 26.4 | 17.2
2007-05-08 | 9:31 | 4.49 | 7.48 | 12.62 | 97.5 | 26.4 | 17.2
2007-05-08 | 9:46 | 4.58 | 7.48 | 12.59 | 97.5 | 26.4 | 17.2
2007-05-08 | 10:01 | 4.73 | 7.48 | 12.48 | 97.1 | 26.4 | 17.2
2007-05-08 | 10:16 | 4.86 | 7.48 | 12.49 | 97.5 | 26.5 | 17.2
2007-05-08 | 10:31 | 5.02 | 7.49 | 12.46 | 97.6 | 26.5 | 17.2
2007-05-08 | 10:46 | 5.21 | 7.5 | 12.4 | 97.6 | 26.5 | 17.2
2007-05-08 | 11:01 | 5.42 | 7.5 | 12.35 | 97.7 | 26.5 | 17.2
2007-05-08 | 11:16 | 5.7 | 7.51 | 12.25 | 97.6 | 26.5 | 17.2
2007-05-08 | 11:31 | 5.98 | 7.51 | 12.17 | 97.8 | 26.5 | 17.2
2007-05-08 | 11:46 | 6.21 | 7.51 | 12.09 | 97.7 | 26.5 | 17.2
2007-05-08 | 12:01 | 6.48 | 7.52 | 12.02 | 97.7 | 26.5 | 17.2
2007-05-08 | 12:16 | 6.73 | 7.53 | 11.96 | 97.9 | 26.5 | 17.3
2007-05-08 | 12:31 | 7.13 | 7.53 | 11.85 | 97.9 | 26.5 | 17.3
2007-05-08 | 12:46 | 7.48 | 7.53 | 11.75 | 98 | 26.6 | 17.3
2007-05-08 | 13:00 | 7.81 | 7.54 | 11.65 | 97.9 | 26.6 | 17.3
2007-05-08 | 13:15 | 8.13 | 7.54 | 11.56 | 97.9 | 26.6 | 17.3
2007-05-08 | 13:30 | 8.53 | 7.54 | 11.45 | 97.9 | 26.6 | 17.3
2007-05-08 | 13:45 | 8.91 | 7.55 | 11.35 | 97.9 | 26.6 | 17.3
2007-05-08 | 14:00 | 9.27 | 7.55 | 11.26 | 98 | 26.6 | 17.3
2007-05-08 | 14:15 | 9.62 | 7.55 | 11.17 | 98 | 26.6 | 17.3
2007-05-08 | 14:30 | 9.95 | 7.56 | 11.07 | 98 | 26.7 | 17.3
2007-05-08 | 14:45 | 10.25 | 7.56 | 10.99 | 97.9 | 26.7 | 17.3
2007-05-08 | 15:00 | 10.56 | 7.56 | 10.91 | 97.9 | 26.7 | 17.3
2007-05-08 | 15:15 | 10.85 | 7.57 | 10.83 | 97.8 | 26.7 | 17.3
2007-05-08 | 15:30 | 11.1 | 7.57 | 10.76 | 97.8 | 26.7 | 17.3
2007-05-08 | 15:45 | 11.33 | 7.58 | 10.71 | 97.8 | 26.7 | 17.4
2007-05-08 | 16:00 | 11.55 | 7.58 | 10.64 | 97.7 | 26.7 | 17.4
2007-05-08 | 16:15 | 11.75 | 7.58 | 10.59 | 97.7 | 26.7 | 17.4
2007-05-08 | 16:30 | 11.93 | 7.59 | 10.54 | 97.6 | 26.7 | 17.3
2007-05-08 | 16:45 | 12.06 | 7.58 | 10.52 | 97.7 | 26.6 | 17.3
2007-05-08 | 17:00 | 12.12 | 7.58 | 10.49 | 97.6 | 26.6 | 17.3
2007-05-08 | 17:15 | 12.15 | 7.58 | 10.5 | 97.8 | 26.6 | 17.3
2007-05-08 | 17:30 | 12.1 | 7.58 | 10.51 | 97.7 | 26.5 | 17.2
2007-05-08 | 17:45 | 12.04 | 7.58 | 10.51 | 97.6 | 26.5 | 17.2
2007-05-08 | 18:00 | 11.91 | 7.58 | 10.54 | 97.6 | 26.4 | 17.2
2007-05-08 | 18:15 | 11.74 | 7.57 | 10.57 | 97.5 | 26.4 | 17.1
2007-05-08 | 18:30 | 11.61 | 7.56 | 10.6 | 97.4 | 26.3 | 17.1
2007-05-08 | 18:45 | 11.48 | 7.55 | 10.62 | 97.4 | 26.1 | 17
2007-05-08 | 19:00 | 11.35 | 7.55 | 10.66 | 97.4 | 26 | 16.9
2007-05-08 | 19:15 | 11.19 | 7.55 | 10.7 | 97.4 | 25.9 | 16.8
2007-05-08 | 19:30 | 10.99 | 7.55 | 10.75 | 97.4 | 25.8 | 16.7
2007-05-08 | 19:45 | 10.77 | 7.54 | 10.79 | 97.3 | 25.6 | 16.6
2007-05-08 | 20:00 | 10.54 | 7.52 | 10.86 | 97.4 | 25.4 | 16.5
2007-05-08 | 20:15 | 10.31 | 7.53 | 10.91 | 97.3 | 25.3 | 16.4
2007-05-08 | 20:30 | 10.08 | 7.51 | 10.97 | 97.3 | 25.2 | 16.3
2007-05-08 | 20:45 | 9.84 | 7.5 | 11.06 | 97.6 | 25 | 16.2
2007-05-08 | 21:00 | 9.61 | 7.49 | 11.1 | 97.4 | 24.8 | 16.1
2007-05-08 | 21:15 | 9.38 | 7.48 | 11.18 | 97.6 | 24.7 | 16
2007-05-08 | 21:30 | 9.15 | 7.47 | 11.23 | 97.5 | 24.6 | 16
2007-05-08 | 21:45 | 8.93 | 7.47 | 11.3 | 97.6 | 24.5 | 15.9
2007-05-08 | 22:00 | 8.72 | 7.46 | 11.35 | 97.5 | 24.4 | 15.8
2007-05-08 | 22:15 | 8.52 | 7.46 | 11.4 | 97.5 | 24.3 | 15.8
2007-05-08 | 22:30 | 8.34 | 7.46 | 11.46 | 97.5 | 24.2 | 15.8
2007-05-08 | 22:45 | 8.15 | 7.44 | 11.51 | 97.6 | 24.2 | 15.7
2007-05-08 | 23:00 | 7.95 | 7.44 | 11.55 | 97.4 | 24.1 | 15.7
2007-05-08 | 23:15 | 7.75 | 7.43 | 11.62 | 97.5 | 24 | 15.6
2007-05-08 | 23:30 | 7.55 | 7.43 | 11.68 | 97.5 | 24 | 15.6
2007-05-08 | 23:45 | 7.37 | 7.42 | 11.73 | 97.5 | 24 | 15.6
2007-05-09 | 0:00 | 7.19 | 7.42 | 11.78 | 97.5 | 23.9 | 15.5
2007-05-09 | 0:15 | 7.02 | 7.41 | 11.82 | 97.4 | 23.9 | 15.5
2007-05-09 | 0:30 | 6.85 | 7.4 | 11.89 | 97.6 | 23.8 | 15.5
2007-05-09 | 0:45 | 6.7 | 7.4 | 11.94 | 97.6 | 23.8 | 15.5
2007-05-09 | 1:01 | 6.54 | 7.4 | 11.97 | 97.5 | 23.8 | 15.5
2007-05-09 | 1:16 | 6.4 | 7.4 | 12.03 | 97.6 | 23.8 | 15.4
2007-05-09 | 1:31 | 6.25 | 7.4 | 12.07 | 97.5 | 23.7 | 15.4
2007-05-09 | 1:46 | 6.11 | 7.39 | 12.13 | 97.7 | 23.8 | 15.5
2007-05-09 | 2:01 | 5.97 | 7.39 | 12.15 | 97.5 | 23.7 | 15.4
2007-05-09 | 2:16 | 5.84 | 7.4 | 12.2 | 97.6 | 23.7 | 15.4
2007-05-09 | 2:31 | 5.71 | 7.39 | 12.24 | 97.6 | 23.8 | 15.4
2007-05-09 | 2:46 | 5.59 | 7.38 | 12.28 | 97.6 | 23.8 | 15.4
2007-05-09 | 3:01 | 5.46 | 7.39 | 12.31 | 97.5 | 23.8 | 15.5
2007-05-09 | 3:16 | 5.34 | 7.39 | 12.34 | 97.5 | 23.8 | 15.5
2007-05-09 | 3:31 | 5.22 | 7.39 | 12.4 | 97.6 | 23.8 | 15.5
2007-05-09 | 3:46 | 5.1 | 7.38 | 12.42 | 97.5 | 23.8 | 15.5
2007-05-09 | 4:01 | 4.99 | 7.38 | 12.44 | 97.4 | 23.8 | 15.5
2007-05-09 | 4:16 | 4.87 | 7.38 | 12.49 | 97.5 | 23.9 | 15.5
2007-05-09 | 4:31 | 4.76 | 7.38 | 12.47 | 97.1 | 23.9 | 15.5
2007-05-09 | 4:46 | 4.65 | 7.37 | 12.54 | 97.3 | 23.9 | 15.5
2007-05-09 | 5:01 | 4.54 | 7.37 | 12.59 | 97.4 | 23.9 | 15.5
2007-05-09 | 5:16 | 4.44 | 7.37 | 12.62 | 97.4 | 23.9 | 15.6
2007-05-09 | 5:31 | 4.35 | 7.37 | 12.67 | 97.5 | 24 | 15.6
2007-05-09 | 5:46 | 4.26 | 7.37 | 12.69 | 97.4 | 24 | 15.6
2007-05-09 | 6:01 | 4.17 | 7.37 | 12.72 | 97.4 | 24 | 15.6
2007-05-09 | 6:16 | 4.09 | 7.37 | 12.76 | 97.6 | 24 | 15.6
2007-05-09 | 6:31 | 4.02 | 7.37 | 12.78 | 97.6 | 24.1 | 15.6
2007-05-09 | 6:46 | 3.95 | 7.38 | 12.8 | 97.5 | 24.1 | 15.7
2007-05-09 | 7:01 | 3.9 | 7.38 | 12.82 | 97.5 | 24.1 | 15.7
2007-05-09 | 7:16 | 3.85 | 7.38 | 12.84 | 97.6 | 24.1 | 15.7
2007-05-09 | 7:31 | 3.81 | 7.39 | 12.86 | 97.6 | 24.2 | 15.7
2007-05-09 | 7:46 | 3.79 | 7.38 | 12.87 | 97.6 | 24.2 | 15.7
2007-05-09 | 8:01 | 3.77 | 7.39 | 12.89 | 97.7 | 24.2 | 15.7
2007-05-09 | 8:16 | 3.77 | 7.4 | 12.87 | 97.6 | 24.2 | 15.8
2007-05-09 | 8:31 | 3.78 | 7.42 | 12.88 | 97.7 | 24.3 | 15.8
2007-05-09 | 8:46 | 3.81 | 7.42 | 12.87 | 97.7 | 24.3 | 15.8
2007-05-09 | 9:01 | 3.85 | 7.42 | 12.87 | 97.8 | 24.3 | 15.8
2007-05-09 | 9:16 | 3.92 | 7.43 | 12.85 | 97.8 | 24.3 | 15.8
2007-05-09 | 9:31 | 3.99 | 7.44 | 12.83 | 97.9 | 24.4 | 15.8
2007-05-09 | 9:46 | 4.08 | 7.44 | 12.81 | 97.9 | 24.4 | 15.8
2007-05-09 | 10:01 | 4.18 | 7.44 | 12.76 | 97.8 | 24.3 | 15.8
2007-05-09 | 10:16 | 4.31 | 7.45 | 12.73 | 97.9 | 24.4 | 15.8
2007-05-09 | 10:31 | 4.44 | 7.44 | 12.69 | 97.9 | 24.4 | 15.9
2007-05-09 | 10:46 | 4.59 | 7.45 | 12.64 | 97.9 | 24.5 | 15.9
2007-05-09 | 11:01 | 4.77 | 7.45 | 12.59 | 98 | 24.5 | 15.9
2007-05-09 | 11:16 | 4.96 | 7.46 | 12.52 | 98 | 24.5 | 15.9
2007-05-09 | 11:31 | 5.18 | 7.47 | 12.46 | 98 | 24.5 | 15.9
2007-05-09 | 11:46 | 5.41 | 7.47 | 12.38 | 97.9 | 24.6 | 16
2007-05-09 | 12:01 | 5.65 | 7.47 | 12.31 | 98 | 24.6 | 16
2007-05-09 | 12:16 | 5.9 | 7.48 | 12.27 | 98.3 | 24.6 | 16
2007-05-09 | 12:31 | 6.16 | 7.47 | 12.17 | 98.2 | 24.6 | 16
2007-05-09 | 12:46 | 6.45 | 7.48 | 12.07 | 98.1 | 24.6 | 16
2007-05-09 | 13:00 | 6.75 | 7.48 | 11.98 | 98 | 24.7 | 16
2007-05-09 | 13:15 | 7.06 | 7.49 | 11.9 | 98.2 | 24.7 | 16.1
2007-05-09 | 13:30 | 7.39 | 7.49 | 11.77 | 97.9 | 24.7 | 16.1
2007-05-09 | 13:45 | 7.74 | 7.5 | 11.7 | 98.2 | 24.7 | 16.1
2007-05-09 | 14:00 | 8.08 | 7.51 | 11.6 | 98.2 | 24.8 | 16.1
2007-05-09 | 14:15 | 8.41 | 7.51 | 11.51 | 98.1 | 24.8 | 16.1
2007-05-09 | 14:30 | 8.73 | 7.52 | 11.41 | 98.1 | 24.8 | 16.1
2007-05-09 | 14:45 | 9.05 | 7.52 | 11.32 | 98 | 24.8 | 16.1
2007-05-09 | 15:00 | 9.35 | 7.53 | 11.25 | 98.1 | 24.8 | 16.1
2007-05-09 | 15:15 | 9.62 | 7.52 | 11.16 | 98 | 24.9 | 16.2
2007-05-09 | 15:30 | 9.88 | 7.53 | 11.09 | 98 | 24.9 | 16.2
2007-05-09 | 15:45 | 10.14 | 7.53 | 11.03 | 98 | 24.9 | 16.2
2007-05-09 | 16:00 | 10.39 | 7.53 | 10.96 | 98 | 24.9 | 16.2
2007-05-09 | 16:15 | 10.57 | 7.53 | 10.9 | 97.9 | 24.9 | 16.2
2007-05-09 | 16:30 | 10.75 | 7.54 | 10.86 | 97.9 | 24.9 | 16.2
2007-05-09 | 16:45 | 10.78 | 7.53 | 10.83 | 97.7 | 24.9 | 16.2
2007-05-09 | 17:00 | 10.83 | 7.54 | 10.84 | 97.9 | 24.9 | 16.2
2007-05-09 | 17:15 | 10.84 | 7.53 | 10.82 | 97.8 | 24.9 | 16.2
2007-05-09 | 17:30 | 10.82 | 7.53 | 10.82 | 97.7 | 24.9 | 16.2
2007-05-09 | 17:45 | 10.8 | 7.53 | 10.82 | 97.7 | 24.9 | 16.2
2007-05-09 | 18:00 | 10.81 | 7.52 | 10.82 | 97.7 | 24.9 | 16.2
2007-05-09 | 18:15 | 10.78 | 7.53 | 10.83 | 97.7 | 24.9 | 16.2
2007-05-09 | 18:30 | 10.76 | 7.52 | 10.83 | 97.7 | 24.8 | 16.2
2007-05-09 | 18:45 | 10.7 | 7.52 | 10.85 | 97.7 | 24.8 | 16.1
2007-05-09 | 19:00 | 10.52 | 7.52 | 10.89 | 97.7 | 24.8 | 16.1
2007-05-09 | 19:15 | 10.3 | 7.51 | 10.96 | 97.7 | 24.7 | 16
2007-05-09 | 19:30 | 10.07 | 7.5 | 11.01 | 97.7 | 24.6 | 16
2007-05-09 | 19:45 | 9.84 | 7.49 | 11.08 | 97.8 | 24.5 | 15.9
2007-05-09 | 20:00 | 9.63 | 7.49 | 11.12 | 97.7 | 24.4 | 15.8
2007-05-09 | 20:15 | 9.42 | 7.48 | 11.18 | 97.7 | 24.3 | 15.8
2007-05-09 | 20:30 | 9.22 | 7.47 | 11.24 | 97.7 | 24.2 | 15.7
2007-05-09 | 20:45 | 9.05 | 7.46 | 11.28 | 97.7 | 24.1 | 15.7
2007-05-09 | 21:00 | 8.89 | 7.47 | 11.32 | 97.7 | 24 | 15.6
2007-05-09 | 21:15 | 8.7 | 7.45 | 11.38 | 97.7 | 24 | 15.6
2007-05-09 | 21:30 | 8.51 | 7.44 | 11.44 | 97.8 | 23.9 | 15.5
2007-05-09 | 21:45 | 8.33 | 7.44 | 11.5 | 97.8 | 23.8 | 15.5
2007-05-09 | 22:00 | 8.15 | 7.43 | 11.55 | 97.9 | 23.8 | 15.5
2007-05-09 | 22:15 | 7.97 | 7.43 | 11.59 | 97.8 | 23.7 | 15.4
2007-05-09 | 22:30 | 7.8 | 7.42 | 11.65 | 97.9 | 23.6 | 15.4
2007-05-09 | 22:45 | 7.62 | 7.42 | 11.7 | 97.8 | 23.6 | 15.3
2007-05-09 | 23:00 | 7.45 | 7.42 | 11.74 | 97.8 | 23.6 | 15.3
2007-05-09 | 23:15 | 7.27 | 7.41 | 11.78 | 97.7 | 23.5 | 15.3
2007-05-09 | 23:30 | 7.1 | 7.41 | 11.85 | 97.8 | 23.5 | 15.3
2007-05-09 | 23:45 | 6.93 | 7.4 | 11.89 | 97.7 | 23.5 | 15.3
2007-05-10 | 0:00 | 6.76 | 7.4 | 11.94 | 97.7 | 23.5 | 15.3
2007-05-10 | 0:15 | 6.6 | 7.4 | 11.99 | 97.8 | 23.5 | 15.3
2007-05-10 | 0:30 | 6.43 | 7.39 | 12.03 | 97.7 | 23.5 | 15.3
2007-05-10 | 0:45 | 6.27 | 7.39 | 12.08 | 97.7 | 23.5 | 15.3
2007-05-10 | 1:01 | 6.11 | 7.39 | 12.14 | 97.8 | 23.5 | 15.3
2007-05-10 | 1:16 | 5.96 | 7.39 | 12.17 | 97.7 | 23.5 | 15.3
2007-05-10 | 1:31 | 5.81 | 7.39 | 12.21 | 97.6 | 23.5 | 15.3
2007-05-10 | 1:46 | 5.67 | 7.39 | 12.25 | 97.6 | 23.5 | 15.3
2007-05-10 | 2:01 | 5.53 | 7.39 | 12.31 | 97.7 | 23.5 | 15.3
2007-05-10 | 2:16 | 5.4 | 7.38 | 12.34 | 97.6 | 23.6 | 15.3
2007-05-10 | 2:31 | 5.27 | 7.38 | 12.38 | 97.6 | 23.6 | 15.3
2007-05-10 | 2:46 | 5.14 | 7.38 | 12.41 | 97.5 | 23.6 | 15.3
2007-05-10 | 3:01 | 5.01 | 7.38 | 12.46 | 97.6 | 23.6 | 15.3
2007-05-10 | 3:16 | 4.9 | 7.38 | 12.48 | 97.5 | 23.6 | 15.4
2007-05-10 | 3:31 | 4.79 | 7.38 | 12.52 | 97.5 | 23.6 | 15.4
2007-05-10 | 3:46 | 4.68 | 7.38 | 12.56 | 97.5 | 23.7 | 15.4
2007-05-10 | 4:01 | 4.58 | 7.37 | 12.59 | 97.5 | 23.7 | 15.4
2007-05-10 | 4:16 | 4.48 | 7.37 | 12.61 | 97.4 | 23.7 | 15.4
2007-05-10 | 4:31 | 4.39 | 7.37 | 12.67 | 97.6 | 23.8 | 15.5
2007-05-10 | 4:46 | 4.3 | 7.38 | 12.68 | 97.5 | 23.8 | 15.5
2007-05-10 | 5:01 | 4.21 | 7.38 | 12.72 | 97.6 | 23.8 | 15.5
2007-05-10 | 5:16 | 4.13 | 7.38 | 12.74 | 97.6 | 23.9 | 15.5
2007-05-10 | 5:31 | 4.06 | 7.37 | 12.77 | 97.6 | 23.9 | 15.5
2007-05-10 | 5:46 | 3.99 | 7.38 | 12.79 | 97.6 | 23.9 | 15.5
2007-05-10 | 6:01 | 3.92 | 7.38 | 12.82 | 97.6 | 23.9 | 15.6
2007-05-10 | 6:16 | 3.86 | 7.38 | 12.83 | 97.5 | 24 | 15.6
2007-05-10 | 6:31 | 3.79 | 7.38 | 12.87 | 97.6 | 24 | 15.6
2007-05-10 | 6:46 | 3.73 | 7.38 | 12.89 | 97.7 | 24 | 15.6
2007-05-10 | 7:01 | 3.68 | 7.39 | 12.91 | 97.7 | 24.1 | 15.6
2007-05-10 | 7:16 | 3.64 | 7.38 | 12.92 | 97.6 | 24.1 | 15.7
2007-05-10 | 7:31 | 3.62 | 7.39 | 12.94 | 97.7 | 24.1 | 15.7
2007-05-10 | 7:46 | 3.6 | 7.4 | 12.95 | 97.8 | 24.2 | 15.7
2007-05-10 | 8:01 | 3.6 | 7.4 | 12.96 | 97.8 | 24.2 | 15.7
2007-05-10 | 8:16 | 3.61 | 7.41 | 12.97 | 97.9 | 24.2 | 15.7
2007-05-10 | 8:31 | 3.64 | 7.42 | 12.94 | 97.8 | 24.3 | 15.8
2007-05-10 | 8:46 | 3.68 | 7.42 | 12.93 | 97.8 | 24.2 | 15.8
2007-05-10 | 9:01 | 3.74 | 7.42 | 12.92 | 97.9 | 24.2 | 15.8
2007-05-10 | 9:16 | 3.82 | 7.43 | 12.89 | 97.9 | 24.3 | 15.8
2007-05-10 | 9:31 | 3.9 | 7.43 | 12.86 | 97.9 | 24.4 | 15.8
2007-05-10 | 9:46 | 4 | 7.43 | 12.82 | 97.8 | 24.3 | 15.8
2007-05-10 | 10:01 | 4.11 | 7.44 | 12.8 | 97.9 | 24.4 | 15.8
2007-05-10 | 10:16 | 4.24 | 7.45 | 12.76 | 98 | 24.3 | 15.8
2007-05-10 | 10:31 | 4.38 | 7.45 | 12.72 | 98 | 24.4 | 15.8
2007-05-10 | 10:46 | 4.55 | 7.46 | 12.67 | 98 | 24.4 | 15.9
2007-05-10 | 11:01 | 4.73 | 7.45 | 12.61 | 98 | 24.4 | 15.9
2007-05-10 | 11:16 | 4.93 | 7.46 | 12.55 | 98.1 | 24.4 | 15.9
2007-05-10 | 11:31 | 5.14 | 7.46 | 12.48 | 98.1 | 24.5 | 15.9
2007-05-10 | 11:46 | 5.37 | 7.47 | 12.42 | 98.1 | 24.5 | 15.9
2007-05-10 | 12:01 | 5.61 | 7.47 | 12.34 | 98.1 | 24.5 | 15.9
2007-05-10 | 12:16 | 5.87 | 7.47 | 12.26 | 98.2 | 24.5 | 15.9
2007-05-10 | 12:31 | 6.14 | 7.48 | 12.18 | 98.2 | 24.5 | 15.9
2007-05-10 | 12:46 | 6.43 | 7.48 | 12.09 | 98.2 | 24.5 | 16
2007-05-10 | 13:00 | 6.74 | 7.48 | 12 | 98.2 | 24.6 | 16
2007-05-10 | 13:15 | 7.05 | 7.5 | 11.92 | 98.3 | 24.6 | 16
2007-05-10 | 13:30 | 7.33 | 7.5 | 11.83 | 98.2 | 24.6 | 16
2007-05-10 | 13:45 | 7.6 | 7.5 | 11.75 | 98.2 | 24.6 | 16
2007-05-10 | 14:00 | 7.9 | 7.51 | 11.67 | 98.3 | 24.6 | 16
2007-05-10 | 14:15 | 8.26 | 7.51 | 11.57 | 98.3 | 24.6 | 16
2007-05-10 | 14:30 | 8.59 | 7.51 | 11.47 | 98.2 | 24.7 | 16
2007-05-10 | 14:45 | 8.88 | 7.51 | 11.39 | 98.2 | 24.7 | 16
2007-05-10 | 15:00 | 9.19 | 7.52 | 11.3 | 98.2 | 24.7 | 16
2007-05-10 | 15:15 | 9.47 | 7.52 | 11.23 | 98.2 | 24.7 | 16.1
2007-05-10 | 15:30 | 9.74 | 7.52 | 11.16 | 98.3 | 24.7 | 16.1
2007-05-10 | 15:45 | 10.02 | 7.54 | 11.08 | 98.2 | 24.7 | 16.1
2007-05-10 | 16:00 | 10.21 | 7.54 | 11.02 | 98.1 | 24.7 | 16.1
2007-05-10 | 16:15 | 10.41 | 7.54 | 10.98 | 98.2 | 24.8 | 16.1
2007-05-10 | 16:30 | 10.6 | 7.54 | 10.92 | 98.1 | 24.8 | 16.1
2007-05-10 | 16:45 | 10.72 | 7.54 | 10.84 | 97.7 | 24.8 | 16.1
2007-05-10 | 17:00 | 10.83 | 7.54 | 10.86 | 98.1 | 24.7 | 16.1
2007-05-10 | 17:15 | 10.89 | 7.54 | 10.84 | 98.1 | 24.8 | 16.1
2007-05-10 | 17:30 | 10.92 | 7.55 | 10.83 | 98.1 | 24.7 | 16.1
2007-05-10 | 17:45 | 10.89 | 7.56 | 10.84 | 98.1 | 24.8 | 16.1
2007-05-10 | 18:00 | 10.84 | 7.54 | 10.85 | 98 | 24.7 | 16.1
2007-05-10 | 18:15 | 10.77 | 7.55 | 10.87 | 98 | 24.7 | 16
2007-05-10 | 18:30 | 10.65 | 7.55 | 10.9 | 98 | 24.6 | 16
2007-05-10 | 18:45 | 10.52 | 7.53 | 10.92 | 98 | 24.6 | 16
2007-05-10 | 19:00 | 10.37 | 7.53 | 10.97 | 98 | 24.5 | 15.9
2007-05-10 | 19:15 | 10.2 | 7.53 | 11.01 | 98 | 24.5 | 15.9
2007-05-10 | 19:30 | 10 | 7.51 | 11.05 | 97.9 | 24.4 | 15.9
2007-05-10 | 19:45 | 9.79 | 7.5 | 11.11 | 97.9 | 24.4 | 15.8
2007-05-10 | 20:00 | 9.57 | 7.5 | 11.16 | 97.9 | 24.3 | 15.8
2007-05-10 | 20:15 | 9.37 | 7.49 | 11.23 | 98 | 24.2 | 15.7
2007-05-10 | 20:30 | 9.17 | 7.48 | 11.29 | 98 | 24.1 | 15.7
2007-05-10 | 20:45 | 8.97 | 7.47 | 11.33 | 98 | 24 | 15.6
2007-05-10 | 21:00 | 8.77 | 7.47 | 11.4 | 98 | 23.9 | 15.6
2007-05-10 | 21:15 | 8.57 | 7.46 | 11.47 | 98.2 | 23.9 | 15.5
2007-05-10 | 21:30 | 8.35 | 7.44 | 11.52 | 98.1 | 23.8 | 15.5
2007-05-10 | 21:45 | 8.14 | 7.44 | 11.58 | 98.1 | 23.7 | 15.4
2007-05-10 | 22:00 | 7.91 | 7.44 | 11.66 | 98.2 | 23.7 | 15.4
2007-05-10 | 22:15 | 7.7 | 7.42 | 11.7 | 98.1 | 23.6 | 15.3
2007-05-10 | 22:30 | 7.5 | 7.43 | 11.77 | 98.1 | 23.5 | 15.3
2007-05-10 | 22:45 | 7.3 | 7.42 | 11.83 | 98.2 | 23.4 | 15.2
2007-05-10 | 23:00 | 7.11 | 7.4 | 11.88 | 98.1 | 23.4 | 15.2
2007-05-10 | 23:15 | 6.93 | 7.4 | 11.93 | 98.1 | 23.4 | 15.2
2007-05-10 | 23:30 | 6.74 | 7.39 | 11.99 | 98.1 | 23.4 | 15.2
2007-05-10 | 23:45 | 6.57 | 7.39 | 11.98 | 97.6 | 23.3 | 15.2
2007-05-11 | 0:00 | 6.4 | 7.39 | 12.07 | 98 | 23.3 | 15.1
2007-05-11 | 0:15 | 6.23 | 7.39 | 12.14 | 98.1 | 23.1 | 15
2007-05-11 | 0:30 | 6.06 | 7.38 | 12.19 | 98.1 | 23.1 | 15
2007-05-11 | 0:45 | 5.89 | 7.38 | 12.25 | 98.1 | 23.2 | 15.1
2007-05-11 | 1:01 | 5.73 | 7.38 | 12.28 | 98 | 23.2 | 15.1
2007-05-11 | 1:16 | 5.57 | 7.37 | 12.35 | 98.1 | 23.2 | 15
2007-05-11 | 1:31 | 5.42 | 7.37 | 12.39 | 98 | 23.2 | 15.1
2007-05-11 | 1:46 | 5.28 | 7.37 | 12.43 | 98 | 23.2 | 15.1
2007-05-11 | 2:01 | 5.13 | 7.37 | 12.47 | 98 | 23.2 | 15.1
2007-05-11 | 2:16 | 5 | 7.37 | 12.51 | 97.9 | 23.2 | 15.1
2007-05-11 | 2:31 | 4.87 | 7.37 | 12.57 | 98.1 | 23.2 | 15.1
2007-05-11 | 2:46 | 4.75 | 7.36 | 12.6 | 98 | 23.2 | 15.1
2007-05-11 | 3:01 | 4.64 | 7.37 | 12.64 | 98 | 23.2 | 15.1
2007-05-11 | 3:16 | 4.54 | 7.37 | 12.67 | 98 | 23.2 | 15.1
2007-05-11 | 3:31 | 4.44 | 7.37 | 12.71 | 98.1 | 23.3 | 15.1
2007-05-11 | 3:46 | 4.35 | 7.37 | 12.74 | 98.1 | 23.3 | 15.1
2007-05-11 | 4:01 | 4.25 | 7.37 | 12.77 | 98 | 23.3 | 15.1
2007-05-11 | 4:16 | 4.15 | 7.37 | 12.8 | 98 | 23.3 | 15.2
2007-05-11 | 4:31 | 4.06 | 7.37 | 12.83 | 98 | 23.4 | 15.2
2007-05-11 | 4:46 | 3.98 | 7.37 | 12.86 | 98.1 | 23.4 | 15.2
2007-05-11 | 5:01 | 3.9 | 7.37 | 12.9 | 98.1 | 23.4 | 15.2
2007-05-11 | 5:16 | 3.82 | 7.37 | 12.92 | 98.1 | 23.5 | 15.2
2007-05-11 | 5:31 | 3.76 | 7.37 | 12.94 | 98.1 | 23.5 | 15.3
2007-05-11 | 5:46 | 3.69 | 7.37 | 12.95 | 98 | 23.5 | 15.3
2007-05-11 | 6:01 | 3.62 | 7.37 | 12.96 | 97.9 | 23.5 | 15.3
2007-05-11 | 6:16 | 3.57 | 7.36 | 12.99 | 98 | 23.6 | 15.3
2007-05-11 | 6:31 | 3.52 | 7.37 | 13.02 | 98.1 | 23.6 | 15.3
2007-05-11 | 6:46 | 3.49 | 7.37 | 13.03 | 98 | 23.6 | 15.3
2007-05-11 | 7:01 | 3.46 | 7.37 | 13.02 | 97.9 | 23.5 | 15.3
2007-05-11 | 7:16 | 3.44 | 7.37 | 13.06 | 98.2 | 23.6 | 15.4
2007-05-11 | 7:31 | 3.44 | 7.38 | 13.05 | 98.1 | 23.6 | 15.4
2007-05-11 | 7:46 | 3.46 | 7.39 | 13.05 | 98.2 | 23.7 | 15.4
2007-05-11 | 8:01 | 3.47 | 7.4 | 13.05 | 98.2 | 23.7 | 15.4
2007-05-11 | 8:16 | 3.51 | 7.41 | 13.04 | 98.2 | 23.7 | 15.4
2007-05-11 | 8:31 | 3.54 | 7.42 | 13.04 | 98.3 | 23.7 | 15.4
2007-05-11 | 8:46 | 3.58 | 7.42 | 13.03 | 98.3 | 23.8 | 15.5
2007-05-11 | 9:01 | 3.63 | 7.43 | 13.01 | 98.3 | 23.8 | 15.5
2007-05-11 | 9:16 | 3.7 | 7.43 | 12.99 | 98.3 | 23.8 | 15.5
2007-05-11 | 9:31 | 3.8 | 7.44 | 12.96 | 98.3 | 23.9 | 15.5
2007-05-11 | 9:46 | 3.91 | 7.45 | 12.92 | 98.3 | 23.8 | 15.5
2007-05-11 | 10:01 | 4.04 | 7.45 | 12.89 | 98.4 | 23.8 | 15.5
2007-05-11 | 10:16 | 4.16 | 7.46 | 12.84 | 98.4 | 23.9 | 15.5
2007-05-11 | 10:31 | 4.26 | 7.46 | 12.81 | 98.4 | 23.9 | 15.5
2007-05-11 | 10:46 | 4.4 | 7.46 | 12.77 | 98.5 | 23.9 | 15.5
2007-05-11 | 11:01 | 4.56 | 7.46 | 12.72 | 98.4 | 24 | 15.6
2007-05-11 | 11:16 | 4.73 | 7.46 | 12.66 | 98.4 | 23.9 | 15.6
2007-05-11 | 11:31 | 4.91 | 7.47 | 12.6 | 98.5 | 23.9 | 15.6
2007-05-11 | 11:46 | 5.12 | 7.47 | 12.54 | 98.5 | 24 | 15.6
2007-05-11 | 12:01 | 5.34 | 7.48 | 12.47 | 98.5 | 24 | 15.6
2007-05-11 | 12:16 | 5.57 | 7.48 | 12.4 | 98.5 | 24 | 15.6
2007-05-11 | 12:31 | 5.81 | 7.48 | 12.32 | 98.5 | 24 | 15.6
2007-05-11 | 12:46 | 5.99 | 7.48 | 12.25 | 98.4 | 24 | 15.6
2007-05-11 | 13:00 | 6.18 | 7.48 | 12.19 | 98.3 | 24 | 15.6
2007-05-11 | 13:15 | 6.36 | 7.49 | 12.14 | 98.4 | 24.1 | 15.6
2007-05-11 | 13:30 | 6.61 | 7.49 | 12.08 | 98.5 | 24.1 | 15.7
2007-05-11 | 13:45 | 6.86 | 7.49 | 11.99 | 98.4 | 24.1 | 15.7
2007-05-11 | 14:00 | 7.13 | 7.5 | 11.93 | 98.6 | 24.1 | 15.7
2007-05-11 | 14:15 | 7.3 | 7.5 | 11.86 | 98.4 | 24.1 | 15.7
2007-05-11 | 14:30 | 7.5 | 7.5 | 11.81 | 98.5 | 24.2 | 15.7
2007-05-11 | 14:45 | 7.66 | 7.5 | 11.75 | 98.4 | 24.2 | 15.7
2007-05-11 | 15:00 | 7.8 | 7.51 | 11.71 | 98.4 | 24.2 | 15.7
2007-05-11 | 15:15 | 7.87 | 7.51 | 11.68 | 98.3 | 24.2 | 15.7
2007-05-11 | 15:30 | 7.98 | 7.51 | 11.64 | 98.3 | 24.2 | 15.7
2007-05-11 | 15:45 | 8.14 | 7.51 | 11.6 | 98.2 | 24.2 | 15.7
2007-05-11 | 16:00 | 8.3 | 7.51 | 11.55 | 98.2 | 24.2 | 15.7
2007-05-11 | 16:15 | 8.45 | 7.52 | 11.5 | 98.2 | 24.2 | 15.7
2007-05-11 | 16:30 | 8.59 | 7.52 | 11.46 | 98.2 | 24.2 | 15.8
2007-05-11 | 16:45 | 8.72 | 7.52 | 11.43 | 98.2 | 24.2 | 15.8
2007-05-11 | 17:00 | 8.84 | 7.52 | 11.41 | 98.3 | 24.2 | 15.8
2007-05-11 | 17:15 | 8.96 | 7.52 | 11.36 | 98.2 | 24.2 | 15.7
2007-05-11 | 17:30 | 9.02 | 7.52 | 11.35 | 98.3 | 24.2 | 15.8
2007-05-11 | 17:45 | 9.06 | 7.52 | 11.32 | 98.1 | 24.3 | 15.8
2007-05-11 | 18:00 | 9.08 | 7.52 | 11.32 | 98.1 | 24.2 | 15.7
2007-05-11 | 18:15 | 9.06 | 7.52 | 11.32 | 98 | 24.2 | 15.7
2007-05-11 | 18:30 | 9.02 | 7.52 | 11.34 | 98.2 | 24.2 | 15.7
2007-05-11 | 18:45 | 8.98 | 7.52 | 11.36 | 98.2 | 24.2 | 15.7
2007-05-11 | 19:00 | 8.92 | 7.51 | 11.36 | 98.1 | 24.2 | 15.7
2007-05-11 | 19:15 | 8.84 | 7.51 | 11.39 | 98.1 | 24.2 | 15.7
2007-05-11 | 19:30 | 8.75 | 7.5 | 11.41 | 98.1 | 24.1 | 15.7
2007-05-11 | 19:45 | 8.65 | 7.49 | 11.43 | 98 | 24.2 | 15.7
2007-05-11 | 20:00 | 8.54 | 7.48 | 11.47 | 98.1 | 24.2 | 15.7
2007-05-11 | 20:15 | 8.43 | 7.48 | 11.5 | 98.2 | 24.1 | 15.7
2007-05-11 | 20:30 | 8.31 | 7.47 | 11.54 | 98.1 | 24.1 | 15.7
2007-05-11 | 20:45 | 8.19 | 7.47 | 11.55 | 97.9 | 24.1 | 15.6
2007-05-11 | 21:00 | 8.06 | 7.46 | 11.59 | 98 | 24.1 | 15.6
2007-05-11 | 21:15 | 7.93 | 7.45 | 11.65 | 98.2 | 24 | 15.6
2007-05-11 | 21:30 | 7.8 | 7.44 | 11.67 | 98 | 24 | 15.6
2007-05-11 | 21:45 | 7.66 | 7.45 | 11.72 | 98.1 | 24 | 15.6
2007-05-11 | 22:00 | 7.52 | 7.44 | 11.75 | 98.1 | 24 | 15.6
2007-05-11 | 22:15 | 7.38 | 7.43 | 11.8 | 98.1 | 23.9 | 15.6
2007-05-11 | 22:30 | 7.25 | 7.43 | 11.83 | 98 | 23.9 | 15.6
2007-05-11 | 22:45 | 7.12 | 7.42 | 11.87 | 98.1 | 23.9 | 15.5
2007-05-11 | 23:00 | 6.99 | 7.42 | 11.9 | 98 | 23.9 | 15.6
2007-05-11 | 23:15 | 6.86 | 7.42 | 11.95 | 98.1 | 23.9 | 15.5
2007-05-11 | 23:30 | 6.73 | 7.41 | 11.99 | 98.1 | 23.8 | 15.5
2007-05-11 | 23:45 | 6.6 | 7.41 | 12.02 | 98.1 | 23.8 | 15.5
2007-05-12 | 0:00 | 6.48 | 7.41 | 12.05 | 98 | 23.9 | 15.5
2007-05-12 | 0:15 | 6.36 | 7.41 | 12.1 | 98.1 | 23.8 | 15.5
2007-05-12 | 0:30 | 6.25 | 7.41 | 12.12 | 97.9 | 23.8 | 15.5
2007-05-12 | 0:45 | 6.13 | 7.41 | 12.16 | 98 | 23.9 | 15.5
2007-05-12 | 1:01 | 6.03 | 7.41 | 12.19 | 98 | 23.9 | 15.5
2007-05-12 | 1:16 | 5.92 | 7.41 | 12.23 | 98 | 23.9 | 15.5
2007-05-12 | 1:31 | 5.81 | 7.41 | 12.26 | 98 | 23.9 | 15.5
2007-05-12 | 1:46 | 5.71 | 7.41 | 12.3 | 98 | 23.9 | 15.5
2007-05-12 | 2:01 | 5.6 | 7.41 | 12.33 | 98 | 23.9 | 15.5
2007-05-12 | 2:16 | 5.49 | 7.4 | 12.35 | 98 | 23.9 | 15.5
2007-05-12 | 2:31 | 5.37 | 7.41 | 12.38 | 97.9 | 24 | 15.6
2007-05-12 | 2:46 | 5.26 | 7.4 | 12.42 | 97.9 | 24 | 15.6
2007-05-12 | 3:01 | 5.16 | 7.4 | 12.46 | 98 | 24 | 15.6
2007-05-12 | 3:16 | 5.05 | 7.4 | 12.5 | 98 | 24 | 15.6
2007-05-12 | 3:31 | 4.96 | 7.4 | 12.53 | 98 | 23.9 | 15.6
2007-05-12 | 3:46 | 4.87 | 7.4 | 12.56 | 98 | 24 | 15.6
2007-05-12 | 4:01 | 4.79 | 7.4 | 12.59 | 98 | 24 | 15.6
2007-05-12 | 4:16 | 4.7 | 7.4 | 12.61 | 98 | 24.3 | 15.8
2007-05-12 | 4:31 | 4.63 | 7.4 | 12.62 | 97.8 | 24.6 | 16
2007-05-12 | 4:46 | 4.55 | 7.39 | 12.66 | 98 | 31.1 | 20.2
2007-05-12 | 5:01 | 4.48 | 7.4 | 12.67 | 97.9 | 27.4 | 17.8
2007-05-12 | 5:16 | 4.42 | 7.4 | 12.69 | 97.9 | 35.3 | 23
2007-05-12 | 5:31 | 4.36 | 7.39 | 12.71 | 97.9 | 30.9 | 20.1
2007-05-12 | 5:46 | 4.3 | 7.4 | 12.73 | 97.9 | 32.4 | 21.1
2007-05-12 | 6:01 | 4.25 | 7.4 | 12.75 | 97.9 | 33 | 21.4
2007-05-12 | 6:16 | 4.21 | 7.39 | 12.77 | 98 | 27.2 | 17.7
2007-05-12 | 6:31 | 4.18 | 7.4 | 12.78 | 97.9 | 26.4 | 17.2
2007-05-12 | 6:46 | 4.15 | 7.4 | 12.78 | 97.9 | 26.6 | 17.3
2007-05-12 | 7:01 | 4.13 | 7.41 | 12.8 | 98 | 31.8 | 20.7
2007-05-12 | 7:16 | 4.13 | 7.41 | 12.82 | 98.1 | 32 | 20.8
2007-05-12 | 7:31 | 4.14 | 7.43 | 12.8 | 98 | 30.7 | 19.9
2007-05-12 | 7:46 | 4.15 | 7.42 | 12.8 | 98 | 32.4 | 21.1
2007-05-12 | 8:01 | 4.17 | 7.43 | 12.79 | 98 | 30.3 | 19.7
2007-05-12 | 8:16 | 4.21 | 7.44 | 12.8 | 98.1 | 31.6 | 20.6
2007-05-12 | 8:31 | 4.25 | 7.45 | 12.78 | 98.1 | 31.6 | 20.6
2007-05-12 | 8:46 | 4.31 | 7.46 | 12.77 | 98.2 | 31.9 | 20.8
2007-05-12 | 9:01 | 4.38 | 7.46 | 12.74 | 98.1 | 31.8 | 20.7
2007-05-12 | 9:16 | 4.4 | 7.46 | 12.72 | 98.1 | 32.2 | 20.9
2007-05-12 | 9:31 | 4.44 | 7.46 | 12.72 | 98.2 | 32.3 | 21
2007-05-12 | 9:46 | 4.5 | 7.46 | 12.71 | 98.2 | 32.6 | 21.2
2007-05-12 | 10:01 | 4.6 | 7.47 | 12.67 | 98.2 | 33.8 | 22
2007-05-12 | 10:16 | 4.71 | 7.47 | 12.65 | 98.3 | 33.9 | 22
2007-05-12 | 10:31 | 4.83 | 7.48 | 12.61 | 98.3 | 33.2 | 21.6
2007-05-12 | 10:46 | 4.98 | 7.48 | 12.57 | 98.4 | 32.9 | 21.4
2007-05-12 | 11:01 | 5.19 | 7.48 | 12.5 | 98.3 | 32.8 | 21.3
2007-05-12 | 11:16 | 5.39 | 7.49 | 12.44 | 98.4 | 32.9 | 21.4
2007-05-12 | 11:31 | 5.54 | 7.49 | 12.38 | 98.3 | 33.3 | 21.7
2007-05-12 | 11:46 | 5.62 | 7.5 | 12.35 | 98.2 | 33.1 | 21.5
2007-05-12 | 12:01 | 5.64 | 7.48 | 12.32 | 98.1 | 32.9 | 21.4
2007-05-12 | 12:16 | 5.85 | 7.5 | 12.31 | 98.5 | 33.3 | 21.7
2007-05-12 | 12:31 | 6.02 | 7.5 | 12.23 | 98.3 | 32.6 | 21.2
2007-05-12 | 12:46 | 6.31 | 7.5 | 12.15 | 98.4 | 32.6 | 21.2
2007-05-12 | 13:00 | 6.41 | 7.5 | 12.11 | 98.3 | 33 | 21.4
2007-05-12 | 13:15 | 6.63 | 7.51 | 12.07 | 98.5 | 33 | 21.5
2007-05-12 | 13:30 | 6.87 | 7.51 | 12 | 98.6 | 32.9 | 21.4
2007-05-12 | 13:45 | 7 | 7.52 | 11.94 | 98.4 | 31.2 | 20.3
2007-05-12 | 14:00 | 7.16 | 7.52 | 11.91 | 98.5 | 31.6 | 20.6
2007-05-12 | 14:15 | 7.31 | 7.52 | 11.87 | 98.5 | 31.5 | 20.4
2007-05-12 | 14:30 | 7.38 | 7.52 | 11.84 | 98.5 | 31.9 | 20.7
2007-05-12 | 14:45 | 7.45 | 7.52 | 11.81 | 98.4 | 32.5 | 21.1
2007-05-12 | 15:00 | 7.63 | 7.52 | 11.77 | 98.5 | 32.1 | 20.9
2007-05-12 | 15:15 | 7.82 | 7.53 | 11.74 | 98.7 | 32.2 | 20.9
2007-05-12 | 15:30 | 7.97 | 7.52 | 11.7 | 98.8 | 32.5 | 21.1
2007-05-12 | 15:45 | 7.99 | 7.52 | 11.69 | 98.7 | 32.4 | 21
2007-05-12 | 16:00 | 7.94 | 7.52 | 11.71 | 98.7 | 32.6 | 21.2
2007-05-12 | 16:15 | 7.89 | 7.51 | 11.71 | 98.6 | 32.7 | 21.3
2007-05-12 | 16:30 | 7.86 | 7.51 | 11.73 | 98.7 | 32.6 | 21.2
2007-05-12 | 16:45 | 7.75 | 7.51 | 11.75 | 98.6 | 32.8 | 21.3
2007-05-12 | 17:00 | 7.72 | 7.51 | 11.77 | 98.7 | 33.1 | 21.5
2007-05-12 | 17:15 | 7.79 | 7.52 | 11.75 | 98.7 | 33.3 | 21.6
2007-05-12 | 17:30 | 7.96 | 7.52 | 11.7 | 98.7 | 33.4 | 21.7
2007-05-12 | 17:45 | 8.1 | 7.52 | 11.67 | 98.8 | 33.4 | 21.7
2007-05-12 | 18:00 | 8.22 | 7.52 | 11.55 | 98 | 32.9 | 21.4
2007-05-12 | 18:15 | 8.37 | 7.51 | 11.54 | 98.4 | 32.9 | 21.4
2007-05-12 | 18:30 | 8.48 | 7.51 | 11.52 | 98.4 | 33 | 21.5
2007-05-12 | 18:45 | 8.49 | 7.51 | 11.53 | 98.5 | 34.1 | 22.1
2007-05-12 | 19:00 | 8.4 | 7.51 | 11.55 | 98.4 | 31.6 | 20.5
2007-05-12 | 19:15 | 8.26 | 7.51 | 11.6 | 98.6 | 31.4 | 20.4
2007-05-12 | 19:30 | 8.08 | 7.5 | 11.65 | 98.6 | 31.3 | 20.3
2007-05-12 | 19:45 | 7.89 | 7.5 | 11.72 | 98.7 | 30.9 | 20.1
2007-05-12 | 20:00 | 7.7 | 7.49 | 11.78 | 98.7 | 29.9 | 19.4
2007-05-12 | 20:15 | 7.52 | 7.48 | 11.81 | 98.5 | 31.7 | 20.6
2007-05-12 | 20:30 | 7.33 | 7.49 | 11.88 | 98.7 | 31.7 | 20.6
2007-05-12 | 20:45 | 7.15 | 7.48 | 11.92 | 98.6 | 33.7 | 21.9
2007-05-12 | 21:00 | 6.97 | 7.48 | 11.98 | 98.6 | 30 | 19.5
2007-05-12 | 21:15 | 6.79 | 7.47 | 12.04 | 98.7 | 30.4 | 19.8
2007-05-12 | 21:30 | 6.6 | 7.46 | 12.08 | 98.6 | 27.8 | 18
2007-05-12 | 21:45 | 6.41 | 7.45 | 12.15 | 98.6 | 31.6 | 20.5
2007-05-12 | 22:00 | 6.21 | 7.44 | 12.21 | 98.6 | 24.2 | 15.7
2007-05-12 | 22:15 | 6.01 | 7.44 | 12.28 | 98.7 | 24.1 | 15.7
2007-05-12 | 22:30 | 5.81 | 7.43 | 12.34 | 98.6 | 24.1 | 15.6
2007-05-12 | 22:45 | 5.61 | 7.41 | 12.4 | 98.7 | 24 | 15.6
2007-05-12 | 23:00 | 5.41 | 7.42 | 12.46 | 98.6 | 24.1 | 15.6
2007-05-12 | 23:15 | 5.22 | 7.41 | 12.52 | 98.6 | 24 | 15.6
2007-05-12 | 23:30 | 5.04 | 7.4 | 12.58 | 98.6 | 24 | 15.6
2007-05-12 | 23:45 | 4.87 | 7.4 | 12.65 | 98.7 | 24 | 15.6
2007-05-13 | 0:00 | 4.71 | 7.4 | 12.69 | 98.6 | 23.9 | 15.5
2007-05-13 | 0:15 | 4.56 | 7.4 | 12.75 | 98.7 | 23.9 | 15.6
2007-05-13 | 0:30 | 4.42 | 7.4 | 12.79 | 98.6 | 24 | 15.6
2007-05-13 | 0:45 | 4.31 | 7.4 | 12.83 | 98.7 | 24 | 15.6
2007-05-13 | 1:01 | 4.25 | 7.41 | 12.85 | 98.6 | 23.9 | 15.5
2007-05-13 | 1:16 | 4.17 | 7.42 | 12.87 | 98.6 | 24 | 15.6
2007-05-13 | 1:31 | 4.07 | 7.42 | 12.9 | 98.6 | 24 | 15.6
2007-05-13 | 1:46 | 3.96 | 7.41 | 12.94 | 98.6 | 24.1 | 15.6
2007-05-13 | 2:01 | 3.84 | 7.41 | 12.97 | 98.5 | 24.1 | 15.7
2007-05-13 | 2:16 | 3.72 | 7.41 | 13.03 | 98.7 | 24.1 | 15.7
2007-05-13 | 2:31 | 3.61 | 7.41 | 13.06 | 98.6 | 24.1 | 15.7
2007-05-13 | 2:46 | 3.49 | 7.4 | 13.12 | 98.7 | 24.1 | 15.7
2007-05-13 | 3:01 | 3.37 | 7.4 | 13.14 | 98.6 | 24.2 | 15.7
2007-05-13 | 3:16 | 3.26 | 7.39 | 13.16 | 98.5 | 24.2 | 15.7
2007-05-13 | 3:31 | 3.14 | 7.39 | 13.22 | 98.6 | 24.2 | 15.8
2007-05-13 | 3:46 | 3.03 | 7.39 | 13.27 | 98.6 | 24.3 | 15.8
2007-05-13 | 4:01 | 2.93 | 7.38 | 13.31 | 98.7 | 24.3 | 15.8
2007-05-13 | 4:16 | 2.83 | 7.39 | 13.32 | 98.5 | 24.3 | 15.8
2007-05-13 | 4:31 | 2.74 | 7.38 | 13.37 | 98.6 | 24.4 | 15.8
2007-05-13 | 4:46 | 2.66 | 7.38 | 13.39 | 98.5 | 24.4 | 15.8
2007-05-13 | 5:01 | 2.59 | 7.38 | 13.41 | 98.5 | 24.4 | 15.9
2007-05-13 | 5:16 | 2.53 | 7.38 | 13.43 | 98.5 | 24.5 | 15.9
2007-05-13 | 5:31 | 2.47 | 7.38 | 13.45 | 98.5 | 24.5 | 15.9
2007-05-13 | 5:46 | 2.41 | 7.38 | 13.48 | 98.6 | 24.6 | 16
2007-05-13 | 6:01 | 2.36 | 7.38 | 13.51 | 98.6 | 24.6 | 16
2007-05-13 | 6:16 | 2.32 | 7.38 | 13.52 | 98.6 | 24.6 | 16
2007-05-13 | 6:31 | 2.28 | 7.38 | 13.55 | 98.7 | 24.6 | 16
2007-05-13 | 6:46 | 2.25 | 7.38 | 13.55 | 98.7 | 24.6 | 16
2007-05-13 | 7:01 | 2.23 | 7.38 | 13.57 | 98.7 | 24.6 | 16
2007-05-13 | 7:16 | 2.2 | 7.39 | 13.58 | 98.7 | 24.6 | 16
2007-05-13 | 7:31 | 2.19 | 7.39 | 13.6 | 98.9 | 24.7 | 16.1
2007-05-13 | 7:46 | 2.2 | 7.4 | 13.59 | 98.8 | 24.7 | 16.1
2007-05-13 | 8:01 | 2.22 | 7.41 | 13.6 | 98.9 | 24.8 | 16.1
2007-05-13 | 8:16 | 2.24 | 7.42 | 13.59 | 98.9 | 24.7 | 16.1
2007-05-13 | 8:31 | 2.27 | 7.42 | 13.57 | 98.9 | 24.8 | 16.1
2007-05-13 | 8:46 | 2.32 | 7.42 | 13.57 | 99 | 24.8 | 16.1
2007-05-13 | 9:01 | 2.38 | 7.43 | 13.56 | 99 | 24.8 | 16.1
2007-05-13 | 9:16 | 2.45 | 7.43 | 13.52 | 99 | 24.8 | 16.1
2007-05-13 | 9:31 | 2.54 | 7.43 | 13.5 | 99.1 | 24.8 | 16.1
2007-05-13 | 9:46 | 2.64 | 7.43 | 13.45 | 99 | 24.8 | 16.1
2007-05-13 | 10:01 | 2.75 | 7.44 | 13.41 | 99 | 24.8 | 16.1
2007-05-13 | 10:16 | 2.89 | 7.44 | 13.37 | 99 | 24.8 | 16.1
2007-05-13 | 10:31 | 3.03 | 7.44 | 13.32 | 99.1 | 24.9 | 16.2
2007-05-13 | 10:46 | 3.19 | 7.42 | 13.27 | 99.1 | 24.9 | 16.2
2007-05-13 | 11:01 | 3.38 | 7.43 | 13.2 | 99 | 24.9 | 16.2
2007-05-13 | 11:16 | 3.58 | 7.45 | 13.14 | 99.1 | 25 | 16.2
2007-05-13 | 11:31 | 3.8 | 7.46 | 13.05 | 99.1 | 25 | 16.2
2007-05-13 | 11:46 | 4.02 | 7.47 | 13 | 99.2 | 25 | 16.2
2007-05-13 | 12:01 | 4.26 | 7.47 | 12.92 | 99.2 | 25 | 16.2
2007-05-13 | 12:16 | 4.51 | 7.46 | 12.83 | 99.2 | 25 | 16.2
2007-05-13 | 12:31 | 4.78 | 7.48 | 12.75 | 99.3 | 25 | 16.2
2007-05-13 | 12:46 | 5.08 | 7.47 | 12.65 | 99.2 | 25 | 16.3
2007-05-13 | 13:00 | 5.38 | 7.48 | 12.56 | 99.3 | 25 | 16.3
2007-05-13 | 13:15 | 5.7 | 7.49 | 12.46 | 99.3 | 25 | 16.3
2007-05-13 | 13:30 | 6.03 | 7.49 | 12.37 | 99.4 | 25 | 16.3
2007-05-13 | 13:45 | 6.35 | 7.49 | 12.27 | 99.4 | 25.1 | 16.3
2007-05-13 | 14:00 | 6.67 | 7.5 | 12.16 | 99.3 | 25.1 | 16.3
2007-05-13 | 14:15 | 6.98 | 7.51 | 12.08 | 99.4 | 25.1 | 16.3
2007-05-13 | 14:30 | 7.28 | 7.51 | 12 | 99.5 | 25 | 16.3
2007-05-13 | 14:45 | 7.58 | 7.51 | 11.9 | 99.5 | 25 | 16.3
2007-05-13 | 15:00 | 7.86 | 7.52 | 11.81 | 99.4 | 25.1 | 16.3
2007-05-13 | 15:15 | 8.12 | 7.52 | 11.74 | 99.4 | 25 | 16.3
2007-05-13 | 15:30 | 8.35 | 7.52 | 11.67 | 99.4 | 25 | 16.3
2007-05-13 | 15:45 | 8.57 | 7.52 | 11.59 | 99.2 | 25.1 | 16.3
2007-05-13 | 16:00 | 8.77 | 7.53 | 11.56 | 99.4 | 25 | 16.3
2007-05-13 | 16:15 | 8.95 | 7.52 | 11.49 | 99.3 | 25.1 | 16.3
2007-05-13 | 16:30 | 9.1 | 7.53 | 11.46 | 99.4 | 25.1 | 16.3
2007-05-13 | 16:45 | 9.22 | 7.52 | 11.42 | 99.3 | 25.1 | 16.3
2007-05-13 | 17:00 | 9.32 | 7.52 | 11.38 | 99.2 | 25.1 | 16.3
2007-05-13 | 17:15 | 9.38 | 7.53 | 11.36 | 99.2 | 25.1 | 16.3
2007-05-13 | 17:30 | 9.43 | 7.54 | 11.35 | 99.2 | 25.1 | 16.3
2007-05-13 | 17:45 | 9.44 | 7.53 | 11.34 | 99.1 | 25.2 | 16.4
2007-05-13 | 18:00 | 9.44 | 7.53 | 11.33 | 99 | 25.2 | 16.4
2007-05-13 | 18:15 | 9.4 | 7.54 | 11.34 | 99.1 | 25.3 | 16.4
2007-05-13 | 18:30 | 9.34 | 7.53 | 11.35 | 99 | 25.3 | 16.4
2007-05-13 | 18:45 | 9.24 | 7.53 | 11.4 | 99.1 | 25.2 | 16.4
2007-05-13 | 19:00 | 9.12 | 7.53 | 11.44 | 99.2 | 25.2 | 16.4
2007-05-13 | 19:15 | 8.96 | 7.52 | 11.47 | 99.1 | 25.2 | 16.4
2007-05-13 | 19:30 | 8.79 | 7.51 | 11.52 | 99.1 | 25.2 | 16.4
2007-05-13 | 19:45 | 8.6 | 7.5 | 11.57 | 99.1 | 25.1 | 16.3
2007-05-13 | 20:00 | 8.42 | 7.5 | 11.62 | 99.1 | 25.1 | 16.3
2007-05-13 | 20:15 | 8.23 | 7.49 | 11.66 | 99 | 25 | 16.3
2007-05-13 | 20:30 | 8.05 | 7.49 | 11.72 | 99 | 25 | 16.3
2007-05-13 | 20:45 | 7.87 | 7.49 | 11.77 | 99.1 | 25 | 16.2
2007-05-13 | 21:00 | 7.68 | 7.46 | 11.82 | 99.1 | 25 | 16.2
2007-05-13 | 21:15 | 7.48 | 7.46 | 11.89 | 99.1 | 24.9 | 16.2
2007-05-13 | 21:30 | 7.29 | 7.45 | 11.94 | 99.1 | 24.8 | 16.1
2007-05-13 | 21:45 | 7.1 | 7.45 | 12.01 | 99.2 | 24.8 | 16.1
2007-05-13 | 22:00 | 6.91 | 7.45 | 12.06 | 99.2 | 24.8 | 16.1
2007-05-13 | 22:15 | 6.73 | 7.44 | 12.12 | 99.2 | 24.7 | 16.1
2007-05-13 | 22:30 | 6.55 | 7.44 | 12.18 | 99.2 | 24.7 | 16.1
2007-05-13 | 22:45 | 6.38 | 7.43 | 12.24 | 99.2 | 24.7 | 16
2007-05-13 | 23:00 | 6.21 | 7.43 | 12.28 | 99.1 | 24.8 | 16.1
2007-05-13 | 23:15 | 6.03 | 7.42 | 12.33 | 99.1 | 24.8 | 16.1
2007-05-13 | 23:30 | 5.87 | 7.42 | 12.37 | 99 | 24.8 | 16.1
2007-05-13 | 23:45 | 5.71 | 7.42 | 12.44 | 99.2 | 24.7 | 16.1
2007-05-14 | 0:00 | 5.55 | 7.42 | 12.47 | 99 | 24.7 | 16.1
2007-05-14 | 0:15 | 5.4 | 7.41 | 12.52 | 99 | 24.7 | 16.1
2007-05-14 | 0:30 | 5.25 | 7.41 | 12.56 | 99 | 24.8 | 16.1
2007-05-14 | 0:45 | 5.11 | 7.41 | 12.63 | 99.2 | 24.7 | 16.1
2007-05-14 | 1:01 | 4.97 | 7.41 | 12.65 | 99 | 24.7 | 16.1
2007-05-14 | 1:16 | 4.83 | 7.4 | 12.69 | 98.9 | 24.8 | 16.1
2007-05-14 | 1:31 | 4.7 | 7.41 | 12.72 | 98.8 | 24.8 | 16.1
2007-05-14 | 1:46 | 4.56 | 7.4 | 12.78 | 98.9 | 24.8 | 16.1
2007-05-14 | 2:01 | 4.43 | 7.4 | 12.83 | 98.9 | 24.9 | 16.2
2007-05-14 | 2:16 | 4.29 | 7.39 | 12.86 | 98.8 | 24.9 | 16.2
2007-05-14 | 2:31 | 4.15 | 7.4 | 12.92 | 98.9 | 24.9 | 16.2
2007-05-14 | 2:46 | 4.02 | 7.4 | 12.98 | 99 | 24.9 | 16.2
2007-05-14 | 3:01 | 3.89 | 7.4 | 13 | 98.9 | 24.9 | 16.2
2007-05-14 | 3:16 | 3.77 | 7.39 | 13.04 | 98.9 | 24.9 | 16.2
2007-05-14 | 3:31 | 3.65 | 7.4 | 13.09 | 98.9 | 24.9 | 16.2
2007-05-14 | 3:46 | 3.54 | 7.39 | 13.12 | 98.9 | 24.9 | 16.2
2007-05-14 | 4:01 | 3.43 | 7.4 | 13.16 | 98.9 | 24.9 | 16.2
2007-05-14 | 4:16 | 3.33 | 7.39 | 13.2 | 98.9 | 24.9 | 16.2
2007-05-14 | 4:31 | 3.23 | 7.39 | 13.23 | 98.9 | 24.9 | 16.2
2007-05-14 | 4:46 | 3.14 | 7.4 | 13.27 | 98.9 | 24.9 | 16.2
2007-05-14 | 5:01 | 3.05 | 7.39 | 13.3 | 98.9 | 24.9 | 16.2
2007-05-14 | 5:16 | 2.97 | 7.39 | 13.33 | 99 | 24.9 | 16.2
2007-05-14 | 5:31 | 2.89 | 7.39 | 13.35 | 98.9 | 24.9 | 16.2
2007-05-14 | 5:46 | 2.81 | 7.39 | 13.39 | 99 | 25 | 16.2
2007-05-14 | 6:01 | 2.74 | 7.38 | 13.42 | 99 | 24.9 | 16.2
2007-05-14 | 6:16 | 2.67 | 7.39 | 13.44 | 98.9 | 25 | 16.2
2007-05-14 | 6:31 | 2.61 | 7.38 | 13.47 | 99 | 25 | 16.2
2007-05-14 | 6:46 | 2.56 | 7.39 | 13.47 | 98.9 | 25 | 16.2
2007-05-14 | 7:01 | 2.52 | 7.39 | 13.5 | 99 | 25 | 16.2
2007-05-14 | 7:16 | 2.48 | 7.39 | 13.53 | 99.1 | 25 | 16.3
2007-05-14 | 7:31 | 2.46 | 7.4 | 13.54 | 99.1 | 25 | 16.3
2007-05-14 | 7:46 | 2.45 | 7.4 | 13.54 | 99.1 | 25 | 16.3
2007-05-14 | 8:01 | 2.46 | 7.41 | 13.55 | 99.2 | 25 | 16.3
2007-05-14 | 8:16 | 2.48 | 7.41 | 13.53 | 99.1 | 25.1 | 16.3
2007-05-14 | 8:31 | 2.51 | 7.42 | 13.53 | 99.2 | 25 | 16.3
2007-05-14 | 8:46 | 2.56 | 7.43 | 13.51 | 99.2 | 25 | 16.3
2007-05-14 | 9:01 | 2.62 | 7.43 | 13.49 | 99.2 | 25 | 16.3
2007-05-14 | 9:16 | 2.7 | 7.44 | 13.41 | 98.9 | 25 | 16.3
2007-05-14 | 9:31 | 2.8 | 7.44 | 13.43 | 99.2 | 25 | 16.3
2007-05-14 | 9:46 | 2.9 | 7.44 | 13.4 | 99.3 | 25 | 16.3
2007-05-14 | 10:01 | 3.02 | 7.44 | 13.35 | 99.2 | 25 | 16.3
2007-05-14 | 10:16 | 3.16 | 7.45 | 13.32 | 99.3 | 25 | 16.3
2007-05-14 | 10:31 | 3.31 | 7.45 | 13.28 | 99.4 | 25.1 | 16.3
2007-05-14 | 10:46 | 3.47 | 7.46 | 13.2 | 99.3 | 25.1 | 16.3
2007-05-14 | 11:01 | 3.67 | 7.45 | 13.13 | 99.3 | 25.1 | 16.3
2007-05-14 | 11:16 | 3.87 | 7.46 | 13.07 | 99.4 | 25.1 | 16.3
2007-05-14 | 11:31 | 4.1 | 7.46 | 12.99 | 99.3 | 25.1 | 16.3
2007-05-14 | 11:46 | 4.33 | 7.47 | 12.92 | 99.4 | 25.1 | 16.3
2007-05-14 | 12:01 | 4.58 | 7.47 | 12.85 | 99.5 | 25.1 | 16.3
2007-05-14 | 12:16 | 4.83 | 7.47 | 12.78 | 99.6 | 25.1 | 16.3
2007-05-14 | 12:31 | 5.11 | 7.48 | 12.66 | 99.4 | 25.1 | 16.3
2007-05-14 | 12:46 | 5.4 | 7.49 | 12.59 | 99.6 | 25.1 | 16.3
2007-05-14 | 13:00 | 5.71 | 7.48 | 12.48 | 99.6 | 25.1 | 16.3
2007-05-14 | 13:15 | 6.04 | 7.49 | 12.37 | 99.5 | 25.1 | 16.3
2007-05-14 | 13:30 | 6.37 | 7.5 | 12.28 | 99.6 | 25.1 | 16.3
2007-05-14 | 13:45 | 6.7 | 7.5 | 12.2 | 99.8 | 25.1 | 16.3
2007-05-14 | 14:00 | 7.03 | 7.5 | 12.1 | 99.7 | 25.1 | 16.3
2007-05-14 | 14:15 | 7.35 | 7.51 | 11.97 | 99.5 | 25.1 | 16.3
2007-05-14 | 14:30 | 7.63 | 7.51 | 11.88 | 99.4 | 25.1 | 16.3
2007-05-14 | 14:45 | 7.92 | 7.51 | 11.81 | 99.5 | 25.1 | 16.3
2007-05-14 | 15:00 | 8.23 | 7.51 | 11.73 | 99.6 | 25.2 | 16.4
2007-05-14 | 15:15 | 8.51 | 7.52 | 11.64 | 99.5 | 25.2 | 16.4
2007-05-14 | 15:30 | 8.79 | 7.52 | 11.56 | 99.5 | 25.2 | 16.4
2007-05-14 | 15:45 | 9.03 | 7.53 | 11.48 | 99.4 | 25.2 | 16.4
2007-05-14 | 16:00 | 9.25 | 7.53 | 11.44 | 99.6 | 25.2 | 16.4
2007-05-14 | 16:15 | 9.45 | 7.53 | 11.36 | 99.3 | 25.2 | 16.4
2007-05-14 | 16:30 | 9.62 | 7.53 | 11.32 | 99.3 | 25.2 | 16.4
2007-05-14 | 16:45 | 9.76 | 7.53 | 11.29 | 99.4 | 25.3 | 16.4
2007-05-14 | 17:00 | 9.86 | 7.54 | 11.26 | 99.4 | 25.3 | 16.4
2007-05-14 | 17:15 | 9.93 | 7.54 | 11.22 | 99.2 | 25.2 | 16.4
2007-05-14 | 17:30 | 9.99 | 7.54 | 11.21 | 99.3 | 25.3 | 16.4
2007-05-14 | 17:45 | 10.04 | 7.54 | 11.19 | 99.2 | 25.3 | 16.4
2007-05-14 | 18:00 | 10.05 | 7.53 | 11.18 | 99.2 | 25.3 | 16.5
2007-05-14 | 18:15 | 10.04 | 7.53 | 11.17 | 99.1 | 25.4 | 16.5
2007-05-14 | 18:30 | 9.99 | 7.53 | 11.18 | 99 | 25.3 | 16.5
2007-05-14 | 18:45 | 9.91 | 7.52 | 11.2 | 99 | 25.3 | 16.4
2007-05-14 | 19:00 | 9.81 | 7.53 | 11.23 | 99 | 25.3 | 16.4
2007-05-14 | 19:15 | 9.67 | 7.51 | 11.26 | 99 | 25.3 | 16.5
2007-05-14 | 19:30 | 9.52 | 7.51 | 11.3 | 99 | 25.3 | 16.5
2007-05-14 | 19:45 | 9.36 | 7.51 | 11.36 | 99.1 | 25.3 | 16.5
2007-05-14 | 20:00 | 9.19 | 7.48 | 11.38 | 98.9 | 25.4 | 16.5
2007-05-14 | 20:15 | 9.03 | 7.49 | 11.43 | 98.9 | 25.4 | 16.5
2007-05-14 | 20:30 | 8.87 | 7.47 | 11.48 | 99 | 25.3 | 16.5
2007-05-14 | 20:45 | 8.7 | 7.47 | 11.53 | 99 | 25.3 | 16.4
2007-05-14 | 21:00 | 8.53 | 7.47 | 11.58 | 99 | 25.3 | 16.4
2007-05-14 | 21:15 | 8.36 | 7.45 | 11.61 | 98.9 | 25.2 | 16.4
2007-05-14 | 21:30 | 8.2 | 7.46 | 11.67 | 99 | 25.2 | 16.3
2007-05-14 | 21:45 | 8.03 | 7.45 | 11.72 | 99 | 25.1 | 16.3
2007-05-14 | 22:00 | 7.87 | 7.45 | 11.78 | 99.1 | 25.1 | 16.3
2007-05-14 | 22:15 | 7.69 | 7.44 | 11.82 | 99.1 | 25 | 16.3
2007-05-14 | 22:30 | 7.51 | 7.43 | 11.88 | 99.1 | 25 | 16.2
2007-05-14 | 22:45 | 7.34 | 7.43 | 11.94 | 99.2 | 24.9 | 16.2
2007-05-14 | 23:00 | 7.17 | 7.42 | 11.98 | 99.1 | 24.9 | 16.2
2007-05-14 | 23:15 | 7.01 | 7.42 | 12.01 | 98.9 | 24.8 | 16.1
2007-05-14 | 23:30 | 6.84 | 7.42 | 12.07 | 99 | 24.8 | 16.1
2007-05-14 | 23:45 | 6.69 | 7.42 | 12.1 | 98.9 | 24.8 | 16.1
2007-05-15 | 0:00 | 6.54 | 7.42 | 12.16 | 99 | 24.8 | 16.1
2007-05-15 | 0:15 | 6.4 | 7.42 | 12.2 | 99 | 24.8 | 16.1
2007-05-15 | 0:30 | 6.25 | 7.41 | 12.24 | 98.9 | 24.8 | 16.1
2007-05-15 | 0:45 | 6.12 | 7.41 | 12.29 | 99 | 24.8 | 16.1
2007-05-15 | 1:01 | 5.98 | 7.41 | 12.33 | 99 | 24.8 | 16.1
2007-05-15 | 1:16 | 5.85 | 7.41 | 12.37 | 99 | 24.8 | 16.1
2007-05-15 | 1:31 | 5.72 | 7.4 | 12.4 | 98.9 | 24.8 | 16.1
2007-05-15 | 1:46 | 5.59 | 7.41 | 12.44 | 98.9 | 24.8 | 16.1
2007-05-15 | 2:01 | 5.46 | 7.4 | 12.48 | 98.9 | 24.8 | 16.1
2007-05-15 | 2:16 | 5.33 | 7.4 | 12.52 | 98.8 | 24.8 | 16.1
2007-05-15 | 2:31 | 5.21 | 7.4 | 12.56 | 98.8 | 24.8 | 16.1
2007-05-15 | 2:46 | 5.09 | 7.4 | 12.59 | 98.8 | 24.8 | 16.1
2007-05-15 | 3:01 | 4.97 | 7.4 | 12.58 | 98.4 | 24.8 | 16.1
2007-05-15 | 3:16 | 4.85 | 7.41 | 12.65 | 98.7 | 24.8 | 16.1
2007-05-15 | 3:31 | 4.74 | 7.4 | 12.69 | 98.7 | 24.9 | 16.2
2007-05-15 | 3:46 | 4.64 | 7.4 | 12.73 | 98.8 | 24.9 | 16.2
2007-05-15 | 4:01 | 4.54 | 7.4 | 12.77 | 98.8 | 24.9 | 16.2
2007-05-15 | 4:16 | 4.45 | 7.4 | 12.8 | 98.8 | 24.9 | 16.2
2007-05-15 | 4:31 | 4.36 | 7.39 | 12.82 | 98.7 | 24.9 | 16.2
2007-05-15 | 4:46 | 4.28 | 7.39 | 12.82 | 98.5 | 24.9 | 16.2
2007-05-15 | 5:01 | 4.2 | 7.39 | 12.88 | 98.7 | 25 | 16.2
2007-05-15 | 5:16 | 4.12 | 7.39 | 12.89 | 98.7 | 25 | 16.2
2007-05-15 | 5:31 | 4.05 | 7.39 | 12.91 | 98.7 | 25 | 16.2
2007-05-15 | 5:46 | 3.98 | 7.39 | 12.94 | 98.7 | 25 | 16.3
2007-05-15 | 6:01 | 3.92 | 7.39 | 12.97 | 98.7 | 25 | 16.3
2007-05-15 | 6:16 | 3.86 | 7.39 | 12.98 | 98.7 | 25 | 16.3
2007-05-15 | 6:31 | 3.81 | 7.39 | 13 | 98.7 | 25.1 | 16.3
2007-05-15 | 6:46 | 3.76 | 7.4 | 13.02 | 98.7 | 25.1 | 16.3
2007-05-15 | 7:01 | 3.73 | 7.42 | 13.04 | 98.8 | 24.4 | 15.8
2007-05-15 | 7:16 | 3.71 | 7.41 | 13.05 | 98.8 | 24.4 | 15.9
2007-05-15 | 7:31 | 3.69 | 7.4 | 13.06 | 98.8 | 24.4 | 15.9
2007-05-15 | 7:46 | 3.68 | 7.41 | 13.06 | 98.8 | 24.4 | 15.9
2007-05-15 | 8:01 | 3.68 | 7.42 | 13.06 | 98.8 | 24.4 | 15.9
2007-05-15 | 8:16 | 3.69 | 7.43 | 13.07 | 98.9 | 24.4 | 15.9
2007-05-15 | 8:31 | 3.71 | 7.43 | 13.05 | 98.8 | 24.4 | 15.9
2007-05-15 | 8:46 | 3.74 | 7.44 | 13.06 | 98.9 | 24.4 | 15.9
2007-05-15 | 9:01 | 3.79 | 7.44 | 13.04 | 98.9 | 24.4 | 15.9
2007-05-15 | 9:16 | 3.86 | 7.44 | 13.01 | 98.9 | 24.5 | 15.9
2007-05-15 | 9:31 | 3.94 | 7.45 | 12.98 | 98.9 | 24.5 | 15.9
2007-05-15 | 9:46 | 4.05 | 7.45 | 12.95 | 98.9 | 24.5 | 15.9
2007-05-15 | 10:01 | 4.16 | 7.46 | 12.92 | 99 | 24.5 | 15.9
2007-05-15 | 10:16 | 4.3 | 7.46 | 12.87 | 98.9 | 24.5 | 15.9
2007-05-15 | 10:31 | 4.45 | 7.46 | 12.85 | 99.1 | 24.5 | 15.9
2007-05-15 | 10:46 | 4.61 | 7.47 | 12.76 | 98.9 | 24.6 | 16
2007-05-15 | 11:01 | 4.8 | 7.47 | 12.73 | 99.2 | 24.5 | 16
2007-05-15 | 11:16 | 5.02 | 7.48 | 12.64 | 99 | 24.6 | 16
2007-05-15 | 11:31 | 5.25 | 7.48 | 12.57 | 99.1 | 24.6 | 16
2007-05-15 | 11:46 | 5.49 | 7.48 | 12.51 | 99.2 | 24.6 | 16
2007-05-15 | 12:01 | 5.74 | 7.48 | 12.42 | 99.1 | 24.6 | 16
2007-05-15 | 12:16 | 6.01 | 7.49 | 12.34 | 99.2 | 24.6 | 16
2007-05-15 | 12:31 | 6.28 | 7.5 | 12.27 | 99.3 | 24.6 | 16
2007-05-15 | 12:46 | 6.57 | 7.5 | 12.17 | 99.2 | 24.6 | 16
2007-05-15 | 13:00 | 6.87 | 7.5 | 12.08 | 99.2 | 24.6 | 16
2007-05-15 | 13:15 | 7.21 | 7.5 | 11.98 | 99.2 | 24.8 | 16.1
2007-05-15 | 13:30 | 7.56 | 7.51 | 11.88 | 99.2 | 24.8 | 16.1
2007-05-15 | 13:45 | 7.92 | 7.51 | 11.79 | 99.3 | 24.8 | 16.1
2007-05-15 | 14:00 | 8.27 | 7.52 | 11.68 | 99.3 | 24.8 | 16.1
2007-05-15 | 14:15 | 8.61 | 7.52 | 11.57 | 99.2 | 24.8 | 16.1
2007-05-15 | 14:30 | 8.93 | 7.52 | 11.47 | 99.1 | 24.8 | 16.1
2007-05-15 | 14:45 | 9.24 | 7.53 | 11.4 | 99.2 | 24.8 | 16.1
2007-05-15 | 15:00 | 9.53 | 7.53 | 11.3 | 99 | 24.8 | 16.1
2007-05-15 | 15:15 | 9.8 | 7.53 | 11.24 | 99.1 | 24.8 | 16.2
2007-05-15 | 15:30 | 10.07 | 7.54 | 11.16 | 99 | 24.9 | 16.2
2007-05-15 | 15:45 | 10.31 | 7.54 | 11.1 | 99.1 | 24.9 | 16.2
2007-05-15 | 16:00 | 10.55 | 7.54 | 11.03 | 99 | 24.9 | 16.2
2007-05-15 | 16:15 | 10.77 | 7.55 | 10.97 | 99 | 24.9 | 16.2
2007-05-15 | 16:30 | 10.95 | 7.55 | 10.92 | 98.9 | 24.9 | 16.2
2007-05-15 | 16:45 | 11.08 | 7.54 | 10.89 | 98.9 | 24.9 | 16.2
2007-05-15 | 17:00 | 11.16 | 7.55 | 10.86 | 98.9 | 25 | 16.2
2007-05-15 | 17:15 | 11.25 | 7.55 | 10.81 | 98.6 | 25 | 16.2
2007-05-15 | 17:30 | 11.34 | 7.55 | 10.82 | 98.9 | 25 | 16.2
2007-05-15 | 17:45 | 11.4 | 7.54 | 10.8 | 98.8 | 24.9 | 16.2
2007-05-15 | 18:00 | 11.43 | 7.55 | 10.79 | 98.8 | 24.9 | 16.2
2007-05-15 | 18:15 | 11.44 | 7.55 | 10.78 | 98.8 | 24.8 | 16.1
2007-05-15 | 18:30 | 11.4 | 7.54 | 10.77 | 98.6 | 24.8 | 16.1
2007-05-15 | 18:45 | 11.34 | 7.54 | 10.79 | 98.6 | 24.8 | 16.1
2007-05-15 | 19:00 | 11.24 | 7.53 | 10.83 | 98.7 | 24.7 | 16.1
2007-05-15 | 19:15 | 11.08 | 7.52 | 10.86 | 98.7 | 24.7 | 16
2007-05-15 | 19:30 | 10.89 | 7.53 | 10.91 | 98.7 | 24.6 | 16
2007-05-15 | 19:45 | 10.69 | 7.51 | 10.96 | 98.7 | 24.6 | 16
2007-05-15 | 20:00 | 10.48 | 7.5 | 11.01 | 98.6 | 24.5 | 16
2007-05-15 | 20:15 | 10.31 | 7.5 | 11.06 | 98.7 | 24.4 | 15.9
2007-05-15 | 20:30 | 10.13 | 7.48 | 11.11 | 98.7 | 24.3 | 15.8
2007-05-15 | 20:45 | 9.97 | 7.47 | 11.15 | 98.7 | 24.3 | 15.8
2007-05-15 | 21:00 | 9.79 | 7.47 | 11.2 | 98.7 | 24.2 | 15.8
2007-05-15 | 21:15 | 9.61 | 7.46 | 11.26 | 98.9 | 24.1 | 15.7
2007-05-15 | 21:30 | 9.42 | 7.46 | 11.3 | 98.8 | 24 | 15.6
2007-05-15 | 21:45 | 9.23 | 7.46 | 11.36 | 98.8 | 24 | 15.6
2007-05-15 | 22:00 | 9.05 | 7.45 | 11.42 | 98.9 | 23.9 | 15.5
2007-05-15 | 22:15 | 8.87 | 7.45 | 11.46 | 98.9 | 23.8 | 15.5
2007-05-15 | 22:30 | 8.69 | 7.44 | 11.51 | 98.8 | 23.6 | 15.4
2007-05-15 | 22:45 | 8.52 | 7.44 | 11.56 | 98.9 | 23.5 | 15.3
2007-05-15 | 23:00 | 8.34 | 7.43 | 11.62 | 99 | 23.4 | 15.2
2007-05-15 | 23:15 | 8.16 | 7.42 | 11.65 | 98.8 | 23.3 | 15.2
2007-05-15 | 23:30 | 7.98 | 7.42 | 11.69 | 98.7 | 23.3 | 15.1
2007-05-15 | 23:45 | 7.8 | 7.42 | 11.76 | 98.8 | 23.1 | 15
2007-05-16 | 0:00 | 7.62 | 7.41 | 11.81 | 98.8 | 23 | 15
2007-05-16 | 0:15 | 7.44 | 7.4 | 11.85 | 98.7 | 22.9 | 14.9
2007-05-16 | 0:30 | 7.27 | 7.41 | 11.85 | 98.3 | 22.7 | 14.8
2007-05-16 | 0:45 | 7.1 | 7.4 | 11.94 | 98.6 | 22.6 | 14.7
2007-05-16 | 1:01 | 6.93 | 7.39 | 12.01 | 98.8 | 22.7 | 14.7
2007-05-16 | 1:16 | 6.78 | 7.39 | 12.04 | 98.7 | 22.5 | 14.6
2007-05-16 | 1:31 | 6.63 | 7.38 | 12.08 | 98.6 | 22.4 | 14.6
2007-05-16 | 1:46 | 6.48 | 7.39 | 12.15 | 98.8 | 22.2 | 14.5
2007-05-16 | 2:01 | 6.35 | 7.38 | 12.18 | 98.7 | 22.1 | 14.4
2007-05-16 | 2:16 | 6.21 | 7.38 | 12.22 | 98.7 | 22.1 | 14.4
2007-05-16 | 2:31 | 6.09 | 7.38 | 12.25 | 98.7 | 22 | 14.3
2007-05-16 | 2:46 | 5.97 | 7.38 | 12.29 | 98.6 | 22.1 | 14.4
2007-05-16 | 3:01 | 5.86 | 7.38 | 12.32 | 98.7 | 22.1 | 14.4
2007-05-16 | 3:16 | 5.76 | 7.38 | 12.36 | 98.7 | 22.1 | 14.4
2007-05-16 | 3:31 | 5.67 | 7.38 | 12.39 | 98.7 | 22.2 | 14.4
2007-05-16 | 3:46 | 5.59 | 7.38 | 12.39 | 98.5 | 22.3 | 14.5
2007-05-16 | 4:01 | 5.51 | 7.38 | 12.43 | 98.6 | 22.3 | 14.5
2007-05-16 | 4:16 | 5.44 | 7.38 | 12.45 | 98.6 | 22.4 | 14.5
2007-05-16 | 4:31 | 5.37 | 7.38 | 12.48 | 98.6 | 22.4 | 14.6
2007-05-16 | 4:46 | 5.3 | 7.37 | 12.5 | 98.7 | 22.5 | 14.6
2007-05-16 | 5:01 | 5.22 | 7.38 | 12.53 | 98.7 | 22.5 | 14.7
2007-05-16 | 5:16 | 5.15 | 7.38 | 12.56 | 98.7 | 22.5 | 14.6
2007-05-16 | 5:31 | 5.08 | 7.39 | 12.57 | 98.7 | 22.5 | 14.6
2007-05-16 | 5:46 | 5.02 | 7.38 | 12.57 | 98.4 | 22.5 | 14.6
2007-05-16 | 6:01 | 4.96 | 7.38 | 12.61 | 98.6 | 22.5 | 14.6
2007-05-16 | 6:16 | 4.91 | 7.38 | 12.63 | 98.6 | 22.6 | 14.7
2007-05-16 | 6:31 | 4.86 | 7.38 | 12.65 | 98.7 | 22.5 | 14.7
2007-05-16 | 6:46 | 4.83 | 7.38 | 12.64 | 98.6 | 22.6 | 14.7
2007-05-16 | 7:01 | 4.8 | 7.39 | 12.68 | 98.8 | 22.5 | 14.6
2007-05-16 | 7:16 | 4.78 | 7.39 | 12.68 | 98.7 | 22.5 | 14.7
2007-05-16 | 7:31 | 4.78 | 7.4 | 12.68 | 98.8 | 22.6 | 14.7
2007-05-16 | 7:46 | 4.79 | 7.4 | 12.68 | 98.8 | 22.6 | 14.7
2007-05-16 | 8:01 | 4.79 | 7.42 | 12.69 | 98.8 | 22.6 | 14.7
2007-05-16 | 8:16 | 4.8 | 7.43 | 12.69 | 98.9 | 22.6 | 14.7
2007-05-16 | 8:31 | 4.81 | 7.43 | 12.69 | 98.9 | 22.6 | 14.7
2007-05-16 | 8:46 | 4.84 | 7.44 | 12.68 | 98.9 | 22.6 | 14.7
2007-05-16 | 9:01 | 4.88 | 7.44 | 12.67 | 98.9 | 22.5 | 14.7
2007-05-16 | 9:16 | 4.95 | 7.45 | 12.65 | 98.9 | 22.6 | 14.7
2007-05-16 | 9:31 | 5.06 | 7.45 | 12.63 | 99.1 | 22.6 | 14.7
2007-05-16 | 9:46 | 5.16 | 7.46 | 12.56 | 98.8 | 22.6 | 14.7
2007-05-16 | 10:01 | 5.28 | 7.46 | 12.54 | 98.9 | 22.6 | 14.7
2007-05-16 | 10:16 | 5.41 | 7.46 | 12.5 | 98.9 | 22.6 | 14.7
2007-05-16 | 10:31 | 5.55 | 7.47 | 12.47 | 99 | 22.6 | 14.7
2007-05-16 | 10:46 | 5.71 | 7.47 | 12.44 | 99.2 | 22.6 | 14.7
2007-05-16 | 11:01 | 5.87 | 7.47 | 12.37 | 99.1 | 22.7 | 14.7
2007-05-16 | 11:16 | 6.04 | 7.48 | 12.34 | 99.2 | 22.6 | 14.7
2007-05-16 | 11:31 | 6.21 | 7.48 | 12.26 | 99 | 22.6 | 14.7
2007-05-16 | 11:46 | 6.38 | 7.48 | 12.22 | 99.1 | 22.7 | 14.7
2007-05-16 | 12:01 | 6.55 | 7.49 | 12.18 | 99.2 | 22.6 | 14.7
2007-05-16 | 12:16 | 6.7 | 7.49 | 12.12 | 99.1 | 22.6 | 14.7
2007-05-16 | 12:31 | 6.88 | 7.49 | 12.05 | 99 | 22.6 | 14.7
2007-05-16 | 12:46 | 7.08 | 7.49 | 12.02 | 99.2 | 22.6 | 14.7
2007-05-16 | 13:00 | 7.33 | 7.49 | 11.92 | 99 | 22.6 | 14.7
2007-05-16 | 13:15 | 7.63 | 7.5 | 11.86 | 99.2 | 22.6 | 14.7
2007-05-16 | 13:30 | 7.95 | 7.5 | 11.78 | 99.3 | 22.6 | 14.7
2007-05-16 | 13:45 | 8.26 | 7.51 | 11.7 | 99.4 | 22.6 | 14.7
2007-05-16 | 14:00 | 8.55 | 7.51 | 11.6 | 99.2 | 22.5 | 14.6
2007-05-16 | 14:15 | 8.85 | 7.52 | 11.52 | 99.3 | 22.5 | 14.6
2007-05-16 | 14:30 | 9.13 | 7.52 | 11.45 | 99.3 | 22.5 | 14.6
2007-05-16 | 14:45 | 9.4 | 7.52 | 11.37 | 99.3 | 22.5 | 14.6
2007-05-16 | 15:00 | 9.68 | 7.52 | 11.27 | 99.1 | 22.5 | 14.6
2007-05-16 | 15:15 | 9.94 | 7.53 | 11.21 | 99.2 | 22.5 | 14.6
2007-05-16 | 15:30 | 10.2 | 7.53 | 11.14 | 99.2 | 22.6 | 14.7
2007-05-16 | 15:45 | 10.43 | 7.52 | 11.07 | 99.1 | 22.6 | 14.7
2007-05-16 | 16:00 | 10.61 | 7.54 | 11.02 | 99 | 22.6 | 14.7
2007-05-16 | 16:15 | 10.79 | 7.53 | 10.98 | 99 | 22.5 | 14.6
2007-05-16 | 16:30 | 10.96 | 7.54 | 10.94 | 99.1 | 22.5 | 14.6
2007-05-16 | 16:45 | 11.04 | 7.53 | 10.92 | 99.1 | 22.5 | 14.6
2007-05-16 | 17:00 | 11.06 | 7.54 | 10.9 | 99 | 22.5 | 14.6
2007-05-16 | 17:15 | 11.03 | 7.54 | 10.92 | 99.1 | 22.5 | 14.6
2007-05-16 | 17:30 | 11.05 | 7.54 | 10.91 | 99 | 22.4 | 14.6
2007-05-16 | 17:45 | 11.13 | 7.54 | 10.87 | 98.8 | 22.3 | 14.5
2007-05-16 | 18:00 | 11.16 | 7.54 | 10.86 | 98.9 | 22.3 | 14.5
2007-05-16 | 18:15 | 11.11 | 7.53 | 10.86 | 98.8 | 22.2 | 14.4
2007-05-16 | 18:30 | 11.03 | 7.53 | 10.88 | 98.7 | 22.2 | 14.4
2007-05-16 | 18:45 | 10.9 | 7.53 | 10.93 | 98.9 | 22.2 | 14.4
2007-05-16 | 19:00 | 10.75 | 7.52 | 10.95 | 98.7 | 22 | 14.3
2007-05-16 | 19:15 | 10.58 | 7.51 | 10.98 | 98.6 | 21.9 | 14.2
2007-05-16 | 19:30 | 10.39 | 7.51 | 11.04 | 98.7 | 21.8 | 14.1
2007-05-16 | 19:45 | 10.21 | 7.5 | 11.09 | 98.7 | 21.7 | 14.1
2007-05-16 | 20:00 | 10.04 | 7.5 | 11.13 | 98.7 | 21.5 | 14
2007-05-16 | 20:15 | 9.86 | 7.49 | 11.19 | 98.8 | 21.5 | 14
2007-05-16 | 20:30 | 9.68 | 7.49 | 11.23 | 98.7 | 21.4 | 13.9
2007-05-16 | 20:45 | 9.49 | 7.48 | 11.27 | 98.6 | 21.2 | 13.8
2007-05-16 | 21:00 | 9.29 | 7.46 | 11.32 | 98.6 | 21.1 | 13.7
2007-05-16 | 21:15 | 9.09 | 7.46 | 11.39 | 98.8 | 20.8 | 13.5
2007-05-16 | 21:30 | 8.89 | 7.45 | 11.44 | 98.7 | 20.6 | 13.4
2007-05-16 | 21:45 | 8.68 | 7.44 | 11.5 | 98.7 | 20.5 | 13.3
2007-05-16 | 22:00 | 8.48 | 7.43 | 11.56 | 98.7 | 20.3 | 13.2
2007-05-16 | 22:15 | 8.28 | 7.42 | 11.63 | 98.9 | 20.2 | 13.1
2007-05-16 | 22:30 | 8.08 | 7.42 | 11.68 | 98.8 | 20.1 | 13.1
2007-05-16 | 22:45 | 7.88 | 7.41 | 11.73 | 98.8 | 20 | 13
2007-05-16 | 23:00 | 7.69 | 7.41 | 11.8 | 98.8 | 20 | 13
2007-05-16 | 23:15 | 7.5 | 7.41 | 11.84 | 98.7 | 19.9 | 13
2007-05-16 | 23:30 | 7.33 | 7.4 | 11.89 | 98.8 | 19.9 | 12.9
2007-05-16 | 23:45 | 7.15 | 7.39 | 11.95 | 98.8 | 19.9 | 12.9
2007-05-17 | 0:00 | 6.99 | 7.39 | 12.01 | 98.9 | 19.8 | 12.9
2007-05-17 | 0:15 | 6.82 | 7.39 | 12.05 | 98.8 | 19.8 | 12.9
2007-05-17 | 0:30 | 6.67 | 7.39 | 12.09 | 98.8 | 19.8 | 12.9
2007-05-17 | 0:45 | 6.52 | 7.38 | 12.13 | 98.8 | 19.8 | 12.9
2007-05-17 | 1:01 | 6.38 | 7.38 | 12.17 | 98.7 | 19.7 | 12.8
2007-05-17 | 1:16 | 6.24 | 7.38 | 12.22 | 98.8 | 19.8 | 12.9
2007-05-17 | 1:31 | 6.1 | 7.38 | 12.26 | 98.7 | 19.8 | 12.9
2007-05-17 | 1:46 | 5.96 | 7.38 | 12.3 | 98.7 | 19.8 | 12.9
2007-05-17 | 2:01 | 5.82 | 7.38 | 12.34 | 98.7 | 19.7 | 12.8
2007-05-17 | 2:16 | 5.68 | 7.37 | 12.38 | 98.6 | 20.1 | 13.1
2007-05-17 | 2:31 | 5.55 | 7.36 | 12.43 | 98.7 | 20.1 | 13
2007-05-17 | 2:46 | 5.41 | 7.36 | 12.47 | 98.7 | 20.1 | 13.1
2007-05-17 | 3:01 | 5.29 | 7.36 | 12.51 | 98.7 | 20.1 | 13.1
2007-05-17 | 3:16 | 5.17 | 7.36 | 12.49 | 98.3 | 20.1 | 13.1
2007-05-17 | 3:31 | 5.06 | 7.36 | 12.56 | 98.5 | 20 | 13
2007-05-17 | 3:46 | 4.95 | 7.36 | 12.61 | 98.6 | 19.8 | 12.9
2007-05-17 | 4:01 | 4.85 | 7.36 | 12.64 | 98.6 | 19.8 | 12.9
2007-05-17 | 4:16 | 4.75 | 7.36 | 12.67 | 98.6 | 19.7 | 12.8
2007-05-17 | 4:31 | 4.67 | 7.35 | 12.7 | 98.6 | 19.7 | 12.8
2007-05-17 | 4:46 | 4.58 | 7.35 | 12.73 | 98.6 | 19.8 | 12.9
2007-05-17 | 5:01 | 4.5 | 7.36 | 12.69 | 98.1 | 19.9 | 12.9
2007-05-17 | 5:16 | 4.43 | 7.36 | 12.77 | 98.5 | 19.9 | 12.9
2007-05-17 | 5:31 | 4.36 | 7.35 | 12.8 | 98.5 | 19.9 | 13
2007-05-17 | 5:46 | 4.29 | 7.36 | 12.82 | 98.6 | 20 | 13
2007-05-17 | 6:01 | 4.22 | 7.36 | 12.82 | 98.4 | 20.1 | 13.1
2007-05-17 | 6:16 | 4.17 | 7.36 | 12.86 | 98.5 | 20.2 | 13.2
2007-05-17 | 6:31 | 4.13 | 7.37 | 12.86 | 98.4 | 20.2 | 13.1
2007-05-17 | 6:46 | 4.09 | 7.37 | 12.88 | 98.5 | 20 | 13
2007-05-17 | 7:01 | 4.06 | 7.37 | 12.91 | 98.6 | 20 | 13
2007-05-17 | 7:16 | 4.05 | 7.38 | 12.92 | 98.6 | 20 | 13
2007-05-17 | 7:31 | 4.04 | 7.37 | 12.92 | 98.6 | 20 | 13
2007-05-17 | 7:46 | 4.04 | 7.39 | 12.91 | 98.6 | 20 | 13
2007-05-17 | 8:01 | 4.04 | 7.39 | 12.91 | 98.6 | 20 | 13
2007-05-17 | 8:16 | 4.06 | 7.39 | 12.91 | 98.7 | 20 | 13
2007-05-17 | 8:31 | 4.09 | 7.41 | 12.9 | 98.7 | 20 | 13
2007-05-17 | 8:46 | 4.13 | 7.42 | 12.9 | 98.7 | 20.1 | 13
2007-05-17 | 9:01 | 4.19 | 7.42 | 12.88 | 98.8 | 20 | 13
2007-05-17 | 9:16 | 4.27 | 7.42 | 12.86 | 98.8 | 19.9 | 12.9
2007-05-17 | 9:31 | 4.36 | 7.43 | 12.82 | 98.7 | 19.8 | 12.9
2007-05-17 | 9:46 | 4.47 | 7.43 | 12.79 | 98.8 | 19.8 | 12.9
2007-05-17 | 10:01 | 4.58 | 7.43 | 12.76 | 98.8 | 19.8 | 12.9
2007-05-17 | 10:16 | 4.72 | 7.44 | 12.7 | 98.7 | 19.8 | 12.9
2007-05-17 | 10:31 | 4.87 | 7.44 | 12.65 | 98.7 | 19.8 | 12.9
2007-05-17 | 10:46 | 5.04 | 7.44 | 12.62 | 98.9 | 19.8 | 12.9
2007-05-17 | 11:01 | 5.23 | 7.45 | 12.55 | 98.8 | 19.8 | 12.9
2007-05-17 | 11:16 | 5.45 | 7.46 | 12.49 | 98.9 | 19.8 | 12.9
2007-05-17 | 11:31 | 5.67 | 7.45 | 12.42 | 98.9 | 19.8 | 12.9
2007-05-17 | 11:46 | 5.92 | 7.46 | 12.34 | 98.9 | 19.8 | 12.9
2007-05-17 | 12:01 | 6.17 | 7.46 | 12.27 | 99 | 19.7 | 12.8
2007-05-17 | 12:16 | 6.45 | 7.46 | 12.2 | 99.1 | 19.7 | 12.8
2007-05-17 | 12:31 | 6.75 | 7.47 | 12.1 | 99.1 | 19.7 | 12.8
2007-05-17 | 12:46 | 7.04 | 7.48 | 11.99 | 98.9 | 19.7 | 12.8
2007-05-17 | 13:00 | 7.35 | 7.48 | 11.94 | 99.2 | 19.7 | 12.8
2007-05-17 | 13:15 | 7.65 | 7.48 | 11.84 | 99.1 | 19.7 | 12.8
2007-05-17 | 13:30 | 7.91 | 7.49 | 11.77 | 99.1 | 19.7 | 12.8
2007-05-17 | 13:45 | 8.19 | 7.49 | 11.7 | 99.3 | 19.6 | 12.8
2007-05-17 | 14:00 | 8.51 | 7.49 | 11.61 | 99.2 | 19.8 | 12.8
2007-05-17 | 14:15 | 8.9 | 7.5 | 11.49 | 99.2 | 19.8 | 12.8
2007-05-17 | 14:30 | 9.22 | 7.5 | 11.41 | 99.2 | 19.9 | 12.9
2007-05-17 | 14:45 | 9.52 | 7.5 | 11.29 | 98.9 | 20 | 13
2007-05-17 | 15:00 | 9.77 | 7.51 | 11.26 | 99.2 | 20 | 13
2007-05-17 | 15:15 | 10.01 | 7.51 | 11.19 | 99.1 | 20 | 13
2007-05-17 | 15:30 | 10.21 | 7.51 | 11.13 | 99.1 | 20 | 13
2007-05-17 | 15:45 | 10.38 | 7.52 | 11.1 | 99.2 | 20.1 | 13
2007-05-17 | 16:00 | 10.54 | 7.52 | 11.07 | 99.3 | 20 | 13
2007-05-17 | 16:15 | 10.68 | 7.52 | 11.01 | 99.1 | 20 | 13
2007-05-17 | 16:30 | 10.79 | 7.52 | 10.92 | 98.6 | 20 | 13
2007-05-17 | 16:45 | 10.89 | 7.52 | 10.92 | 98.7 | 20 | 13
2007-05-17 | 17:00 | 10.95 | 7.52 | 10.94 | 99.1 | 20 | 13
2007-05-17 | 17:15 | 10.98 | 7.52 | 10.92 | 99 | 20 | 13
2007-05-17 | 17:30 | 10.99 | 7.52 | 10.92 | 99 | 20 | 13
2007-05-17 | 17:45 | 10.96 | 7.52 | 10.93 | 99 | 20.1 | 13
2007-05-17 | 18:00 | 10.9 | 7.51 | 10.92 | 98.8 | 20.1 | 13.1
2007-05-17 | 18:15 | 10.82 | 7.51 | 10.94 | 98.8 | 20.1 | 13
2007-05-17 | 18:30 | 10.69 | 7.51 | 10.98 | 98.9 | 20 | 13
2007-05-17 | 18:45 | 10.54 | 7.5 | 11 | 98.7 | 19.9 | 12.9
2007-05-17 | 19:00 | 10.37 | 7.5 | 11.05 | 98.8 | 19.9 | 12.9
2007-05-17 | 19:15 | 10.19 | 7.49 | 11.09 | 98.6 | 19.8 | 12.9
2007-05-17 | 19:30 | 10 | 7.5 | 11.15 | 98.8 | 19.7 | 12.8
2007-05-17 | 19:45 | 9.8 | 7.49 | 11.21 | 98.8 | 19.6 | 12.7
2007-05-17 | 20:00 | 9.61 | 7.48 | 11.25 | 98.8 | 19.6 | 12.7
2007-05-17 | 20:15 | 9.43 | 7.47 | 11.3 | 98.7 | 19.5 | 12.7
2007-05-17 | 20:30 | 9.23 | 7.47 | 11.37 | 98.9 | 19.4 | 12.6
2007-05-17 | 20:45 | 9.04 | 7.46 | 11.41 | 98.8 | 19.4 | 12.6
2007-05-17 | 21:00 | 8.85 | 7.46 | 11.47 | 98.9 | 19.2 | 12.5
2007-05-17 | 21:15 | 8.66 | 7.45 | 11.52 | 98.8 | 19.1 | 12.4
2007-05-17 | 21:30 | 8.47 | 7.45 | 11.58 | 98.9 | 19.1 | 12.4
2007-05-17 | 21:45 | 8.27 | 7.44 | 11.63 | 98.8 | 19.2 | 12.5
2007-05-17 | 22:00 | 8.06 | 7.42 | 11.69 | 98.9 | 19.3 | 12.6
2007-05-17 | 22:15 | 7.85 | 7.42 | 11.75 | 98.8 | 19.2 | 12.5
2007-05-17 | 22:30 | 7.65 | 7.42 | 11.8 | 98.8 | 19.3 | 12.6
2007-05-17 | 22:45 | 7.46 | 7.41 | 11.87 | 98.9 | 19.4 | 12.6
2007-05-17 | 23:00 | 7.27 | 7.4 | 11.95 | 99.1 | 19.4 | 12.6
2007-05-17 | 23:15 | 7.08 | 7.4 | 11.99 | 99 | 19.2 | 12.5
2007-05-17 | 23:30 | 6.9 | 7.39 | 12.03 | 98.8 | 19.2 | 12.5
2007-05-17 | 23:45 | 6.73 | 7.39 | 12.07 | 98.8 | 19.1 | 12.4
2007-05-18 | 0:00 | 6.57 | 7.39 | 12.13 | 98.9 | 19 | 12.3
2007-05-18 | 0:15 | 6.41 | 7.39 | 12.18 | 98.8 | 18.9 | 12.3
2007-05-18 | 0:30 | 6.26 | 7.38 | 12.24 | 99 | 18.9 | 12.3
2007-05-18 | 0:45 | 6.12 | 7.37 | 12.26 | 98.8 | 18.8 | 12.2
2007-05-18 | 1:01 | 5.98 | 7.37 | 12.31 | 98.8 | 18.8 | 12.2
2007-05-18 | 1:16 | 5.84 | 7.38 | 12.34 | 98.7 | 18.8 | 12.2
2007-05-18 | 1:31 | 5.71 | 7.37 | 12.38 | 98.7 | 18.8 | 12.2
2007-05-18 | 1:46 | 5.57 | 7.37 | 12.43 | 98.7 | 18.8 | 12.2
2007-05-18 | 2:01 | 5.44 | 7.36 | 12.46 | 98.7 | 18.7 | 12.1
2007-05-18 | 2:16 | 5.32 | 7.37 | 12.51 | 98.7 | 18.7 | 12.1
2007-05-18 | 2:31 | 5.2 | 7.37 | 12.54 | 98.7 | 18.7 | 12.1
2007-05-18 | 2:46 | 5.09 | 7.36 | 12.58 | 98.8 | 18.6 | 12.1
2007-05-18 | 3:01 | 4.99 | 7.36 | 12.61 | 98.7 | 18.5 | 12.1
2007-05-18 | 3:16 | 4.89 | 7.36 | 12.65 | 98.8 | 18.5 | 12
2007-05-18 | 3:31 | 4.8 | 7.36 | 12.68 | 98.7 | 18.5 | 12
2007-05-18 | 3:46 | 4.72 | 7.36 | 12.7 | 98.7 | 18.6 | 12.1
2007-05-18 | 4:01 | 4.63 | 7.37 | 12.72 | 98.7 | 18.6 | 12.1
2007-05-18 | 4:16 | 4.56 | 7.36 | 12.76 | 98.7 | 18.5 | 12
2007-05-18 | 4:31 | 4.49 | 7.35 | 12.77 | 98.7 | 18.5 | 12
2007-05-18 | 4:46 | 4.43 | 7.36 | 12.8 | 98.7 | 18.6 | 12.1
2007-05-18 | 5:01 | 4.37 | 7.36 | 12.82 | 98.7 | 18.5 | 12.1
2007-05-18 | 5:16 | 4.31 | 7.35 | 12.83 | 98.7 | 18.5 | 12
2007-05-18 | 5:31 | 4.25 | 7.36 | 12.85 | 98.7 | 18.5 | 12
2007-05-18 | 5:46 | 4.19 | 7.36 | 12.86 | 98.6 | 18.5 | 12
2007-05-18 | 6:01 | 4.14 | 7.35 | 12.89 | 98.7 | 18.5 | 12
2007-05-18 | 6:16 | 4.09 | 7.36 | 12.9 | 98.7 | 18.5 | 12
2007-05-18 | 6:31 | 4.05 | 7.35 | 12.92 | 98.7 | 18.5 | 12
2007-05-18 | 6:46 | 4 | 7.36 | 12.92 | 98.6 | 18.6 | 12.1
2007-05-18 | 7:01 | 3.98 | 7.36 | 12.94 | 98.7 | 18.7 | 12.1
2007-05-18 | 7:16 | 3.96 | 7.37 | 12.95 | 98.7 | 18.7 | 12.1
2007-05-18 | 7:31 | 3.95 | 7.38 | 12.97 | 98.8 | 18.6 | 12.1
2007-05-18 | 7:46 | 3.94 | 7.38 | 12.96 | 98.7 | 18.6 | 12.1
2007-05-18 | 8:01 | 3.93 | 7.39 | 12.96 | 98.7 | 18.7 | 12.1
2007-05-18 | 8:16 | 3.94 | 7.39 | 12.95 | 98.7 | 18.6 | 12.1
2007-05-18 | 8:31 | 3.93 | 7.39 | 12.98 | 98.8 | 18.6 | 12.1
2007-05-18 | 8:46 | 3.94 | 7.39 | 12.98 | 98.8 | 18.6 | 12.1
2007-05-18 | 9:01 | 3.96 | 7.4 | 12.98 | 98.9 | 18.6 | 12.1
2007-05-18 | 9:16 | 3.98 | 7.41 | 12.95 | 98.8 | 18.6 | 12.1
2007-05-18 | 9:31 | 4.02 | 7.41 | 12.94 | 98.8 | 18.7 | 12.1
2007-05-18 | 9:46 | 4.09 | 7.42 | 12.92 | 98.8 | 18.7 | 12.1
2007-05-18 | 10:01 | 4.17 | 7.42 | 12.9 | 98.8 | 18.6 | 12.1
2007-05-18 | 10:16 | 4.24 | 7.43 | 12.88 | 98.8 | 18.7 | 12.1
2007-05-18 | 10:31 | 4.34 | 7.42 | 12.83 | 98.7 | 18.6 | 12.1
2007-05-18 | 10:46 | 4.44 | 7.43 | 12.81 | 98.8 | 18.7 | 12.1
2007-05-18 | 11:01 | 4.58 | 7.43 | 12.75 | 98.8 | 18.8 | 12.2
2007-05-18 | 11:16 | 4.74 | 7.44 | 12.7 | 98.8 | 18.9 | 12.3
2007-05-18 | 11:31 | 4.95 | 7.44 | 12.67 | 99 | 18.9 | 12.3
2007-05-18 | 11:46 | 5.17 | 7.44 | 12.59 | 99 | 18.9 | 12.3
2007-05-18 | 12:01 | 5.43 | 7.45 | 12.51 | 99 | 18.9 | 12.3
2007-05-18 | 12:16 | 5.7 | 7.45 | 12.42 | 99 | 18.9 | 12.3
2007-05-18 | 12:31 | 5.99 | 7.45 | 12.35 | 99.2 | 18.9 | 12.3
2007-05-18 | 12:46 | 6.27 | 7.46 | 12.25 | 99.1 | 18.9 | 12.3
2007-05-18 | 13:00 | 6.55 | 7.47 | 12.17 | 99.1 | 18.9 | 12.3
2007-05-18 | 13:15 | 6.89 | 7.47 | 12.07 | 99.1 | 18.9 | 12.3
2007-05-18 | 13:30 | 7.2 | 7.47 | 11.97 | 99.1 | 19 | 12.3
2007-05-18 | 13:45 | 7.52 | 7.48 | 11.87 | 99.1 | 19 | 12.3
2007-05-18 | 14:00 | 7.72 | 7.49 | 11.81 | 99 | 19 | 12.3
2007-05-18 | 14:15 | 7.95 | 7.49 | 11.74 | 99 | 19 | 12.4
2007-05-18 | 14:30 | 8.18 | 7.49 | 11.68 | 99 | 19.1 | 12.4
2007-05-18 | 14:45 | 8.45 | 7.5 | 11.6 | 99 | 19.1 | 12.4
2007-05-18 | 15:00 | 8.54 | 7.5 | 11.56 | 98.9 | 19.1 | 12.4
2007-05-18 | 15:15 | 8.66 | 7.49 | 11.52 | 98.9 | 19.1 | 12.4
2007-05-18 | 15:30 | 8.89 | 7.49 | 11.47 | 98.9 | 19.1 | 12.4
2007-05-18 | 15:45 | 9.13 | 7.49 | 11.39 | 98.8 | 19.1 | 12.4
2007-05-18 | 16:00 | 9.46 | 7.5 | 11.32 | 99 | 19.1 | 12.4
2007-05-18 | 16:15 | 9.8 | 7.5 | 11.24 | 99.1 | 19.1 | 12.4
2007-05-18 | 16:30 | 10.02 | 7.5 | 11.17 | 99 | 19.1 | 12.4
2007-05-18 | 16:45 | 10.22 | 7.5 | 11.13 | 99.1 | 19.1 | 12.4
2007-05-18 | 17:00 | 10.42 | 7.51 | 11.08 | 99.1 | 19.1 | 12.4
2007-05-18 | 17:15 | 10.55 | 7.51 | 11.04 | 99.1 | 19.1 | 12.4
2007-05-18 | 17:30 | 10.65 | 7.51 | 11 | 99 | 19.2 | 12.5
2007-05-18 | 17:45 | 10.76 | 7.51 | 10.98 | 99 | 19.3 | 12.5
2007-05-18 | 18:00 | 10.81 | 7.51 | 10.96 | 99 | 19.3 | 12.5
2007-05-18 | 18:15 | 10.8 | 7.51 | 10.97 | 99 | 19.1 | 12.4
2007-05-18 | 18:30 | 10.72 | 7.51 | 10.99 | 99 | 19.1 | 12.4
2007-05-18 | 18:45 | 10.57 | 7.5 | 11.02 | 99 | 19.1 | 12.4
2007-05-18 | 19:00 | 10.37 | 7.49 | 11.07 | 98.9 | 19.1 | 12.4
2007-05-18 | 19:15 | 10.16 | 7.49 | 11.13 | 99 | 19.1 | 12.4
2007-05-18 | 19:30 | 9.94 | 7.48 | 11.19 | 99 | 19 | 12.4
2007-05-18 | 19:45 | 9.74 | 7.48 | 11.24 | 99 | 19 | 12.3
2007-05-18 | 20:00 | 9.54 | 7.47 | 11.3 | 99 | 18.9 | 12.3
2007-05-18 | 20:15 | 9.34 | 7.47 | 11.34 | 98.9 | 18.7 | 12.2
2007-05-18 | 20:30 | 9.14 | 7.46 | 11.41 | 99 | 18.6 | 12.1
2007-05-18 | 20:45 | 8.95 | 7.45 | 11.46 | 99 | 19.7 | 12.8
2007-05-18 | 21:00 | 8.77 | 7.45 | 11.51 | 99 | 20.3 | 13.2
2007-05-18 | 21:15 | 8.59 | 7.44 | 11.57 | 99.1 | 20.7 | 13.5
2007-05-18 | 21:30 | 8.41 | 7.44 | 11.62 | 99.1 | 21 | 13.7
2007-05-18 | 21:45 | 8.23 | 7.43 | 11.67 | 99.1 | 21.1 | 13.7
2007-05-18 | 22:00 | 8.03 | 7.43 | 11.71 | 98.9 | 21 | 13.7
2007-05-18 | 22:15 | 7.83 | 7.42 | 11.77 | 99 | 20.9 | 13.6
2007-05-18 | 22:30 | 7.62 | 7.42 | 11.84 | 99 | 20.8 | 13.5
2007-05-18 | 22:45 | 7.41 | 7.4 | 11.9 | 99.1 | 20.7 | 13.5
2007-05-18 | 23:00 | 7.22 | 7.4 | 11.95 | 99 | 20.7 | 13.4
2007-05-18 | 23:15 | 7.03 | 7.4 | 12.02 | 99.1 | 20.7 | 13.4
2007-05-18 | 23:30 | 6.85 | 7.38 | 12.05 | 98.9 | 20.6 | 13.4
2007-05-18 | 23:45 | 6.68 | 7.38 | 12.11 | 99 | 20.6 | 13.4
2007-05-19 | 0:00 | 6.53 | 7.38 | 12.16 | 99 | 20.6 | 13.4
2007-05-19 | 0:15 | 6.39 | 7.38 | 12.2 | 99 | 20.6 | 13.4
2007-05-19 | 0:30 | 6.26 | 7.38 | 12.24 | 99 | 20.6 | 13.4
2007-05-19 | 0:45 | 6.15 | 7.37 | 12.28 | 99 | 20.6 | 13.4
2007-05-19 | 1:01 | 6.05 | 7.38 | 12.31 | 99 | 20.4 | 13.2
2007-05-19 | 1:16 | 5.96 | 7.37 | 12.34 | 99 | 20.4 | 13.3
2007-05-19 | 1:31 | 5.87 | 7.38 | 12.36 | 99 | 20.3 | 13.2
2007-05-19 | 1:46 | 5.79 | 7.39 | 12.37 | 98.9 | 20.2 | 13.1
2007-05-19 | 2:01 | 5.71 | 7.38 | 12.41 | 98.9 | 20.1 | 13.1
2007-05-19 | 2:16 | 5.64 | 7.38 | 12.42 | 98.9 | 20 | 13
2007-05-19 | 2:31 | 5.58 | 7.38 | 12.46 | 99 | 20 | 13
2007-05-19 | 2:46 | 5.51 | 7.39 | 12.48 | 99 | 19.8 | 12.9
2007-05-19 | 3:01 | 5.45 | 7.37 | 12.5 | 99 | 19.8 | 12.9
2007-05-19 | 3:16 | 5.38 | 7.39 | 12.51 | 98.9 | 19.8 | 12.9
2007-05-19 | 3:31 | 5.32 | 7.37 | 12.53 | 98.9 | 19.7 | 12.8
2007-05-19 | 3:46 | 5.27 | 7.37 | 12.56 | 99 | 19.8 | 12.8
2007-05-19 | 4:01 | 5.21 | 7.38 | 12.57 | 99 | 19.7 | 12.8
2007-05-19 | 4:16 | 5.17 | 7.37 | 12.6 | 99.1 | 19.6 | 12.7
2007-05-19 | 4:31 | 5.14 | 7.38 | 12.58 | 98.9 | 19.6 | 12.7
2007-05-19 | 4:46 | 5.11 | 7.38 | 12.61 | 99 | 19.6 | 12.7
2007-05-19 | 5:01 | 5.09 | 7.39 | 12.62 | 99 | 19.6 | 12.7
2007-05-19 | 5:16 | 5.05 | 7.38 | 12.63 | 99 | 19.6 | 12.7
2007-05-19 | 5:31 | 5.02 | 7.38 | 12.63 | 99 | 19.5 | 12.7
2007-05-19 | 5:46 | 5.01 | 7.39 | 12.64 | 99 | 19.5 | 12.7
2007-05-19 | 6:01 | 4.99 | 7.38 | 12.65 | 99.1 | 19.5 | 12.7
2007-05-19 | 6:16 | 4.98 | 7.38 | 12.66 | 99.1 | 19.5 | 12.7
2007-05-19 | 6:31 | 4.97 | 7.38 | 12.67 | 99.1 | 19.5 | 12.7
2007-05-19 | 6:46 | 4.97 | 7.39 | 12.66 | 99.1 | 19.5 | 12.6
2007-05-19 | 7:01 | 4.97 | 7.39 | 12.66 | 99.1 | 19.3 | 12.6
2007-05-19 | 7:16 | 4.97 | 7.39 | 12.67 | 99.1 | 19.3 | 12.6
2007-05-19 | 7:31 | 4.96 | 7.4 | 12.68 | 99.2 | 19.3 | 12.6
2007-05-19 | 7:46 | 4.95 | 7.4 | 12.69 | 99.2 | 19.3 | 12.5
2007-05-19 | 8:01 | 4.95 | 7.4 | 12.68 | 99.2 | 19.3 | 12.5
2007-05-19 | 8:16 | 4.95 | 7.42 | 12.68 | 99.2 | 19.2 | 12.5
2007-05-19 | 8:31 | 4.97 | 7.42 | 12.69 | 99.3 | 19.1 | 12.4
2007-05-19 | 8:46 | 4.99 | 7.42 | 12.68 | 99.3 | 19.1 | 12.4
2007-05-19 | 9:01 | 5.01 | 7.42 | 12.67 | 99.2 | 19.1 | 12.4
2007-05-19 | 9:16 | 5.03 | 7.42 | 12.69 | 99.4 | 19.2 | 12.5
2007-05-19 | 9:31 | 5.06 | 7.43 | 12.65 | 99.2 | 19.2 | 12.5
2007-05-19 | 9:46 | 5.13 | 7.44 | 12.64 | 99.3 | 19.1 | 12.4
2007-05-19 | 10:01 | 5.21 | 7.43 | 12.61 | 99.3 | 19 | 12.4
2007-05-19 | 10:16 | 5.29 | 7.43 | 12.6 | 99.4 | 19.1 | 12.4
2007-05-19 | 10:31 | 5.36 | 7.43 | 12.56 | 99.3 | 19.1 | 12.4
2007-05-19 | 10:46 | 5.45 | 7.43 | 12.54 | 99.3 | 19 | 12.4
2007-05-19 | 11:01 | 5.55 | 7.44 | 12.51 | 99.3 | 19.1 | 12.4
2007-05-19 | 11:16 | 5.7 | 7.45 | 12.47 | 99.4 | 19 | 12.4
2007-05-19 | 11:31 | 5.87 | 7.44 | 12.46 | 99.7 | 19.1 | 12.4
2007-05-19 | 11:46 | 6.02 | 7.45 | 12.37 | 99.4 | 19 | 12.4
2007-05-19 | 12:01 | 6.16 | 7.45 | 12.31 | 99.3 | 19.1 | 12.4
2007-05-19 | 12:16 | 6.24 | 7.46 | 12.29 | 99.3 | 19 | 12.4
2007-05-19 | 12:31 | 6.33 | 7.46 | 12.29 | 99.5 | 19 | 12.4
2007-05-19 | 12:46 | 6.44 | 7.45 | 12.24 | 99.4 | 19 | 12.4
2007-05-19 | 13:00 | 6.68 | 7.45 | 12.16 | 99.4 | 19 | 12.4
2007-05-19 | 13:15 | 6.82 | 7.47 | 12.13 | 99.4 | 19.1 | 12.4
2007-05-19 | 13:30 | 7.02 | 7.47 | 12.09 | 99.6 | 19.1 | 12.4
2007-05-19 | 13:45 | 7.18 | 7.47 | 12.03 | 99.5 | 19.1 | 12.4
2007-05-19 | 14:00 | 7.34 | 7.47 | 11.97 | 99.5 | 19 | 12.4
2007-05-19 | 14:15 | 7.51 | 7.46 | 11.92 | 99.5 | 19 | 12.4
2007-05-19 | 14:30 | 7.75 | 7.48 | 11.84 | 99.4 | 19.1 | 12.4
2007-05-19 | 14:45 | 7.96 | 7.48 | 11.79 | 99.5 | 19.1 | 12.4
2007-05-19 | 15:00 | 8.2 | 7.48 | 11.72 | 99.4 | 19.1 | 12.4
2007-05-19 | 15:15 | 8.42 | 7.49 | 11.65 | 99.3 | 19.1 | 12.4
2007-05-19 | 15:30 | 8.6 | 7.48 | 11.6 | 99.4 | 19.2 | 12.5
2007-05-19 | 15:45 | 8.71 | 7.48 | 11.57 | 99.4 | 19.1 | 12.4
2007-05-19 | 16:00 | 8.66 | 7.48 | 11.6 | 99.5 | 19.1 | 12.4
2007-05-19 | 16:15 | 8.51 | 7.48 | 11.63 | 99.4 | 19.1 | 12.4
2007-05-19 | 16:30 | 8.36 | 7.48 | 11.69 | 99.6 | 19.1 | 12.4
2007-05-19 | 16:45 | 8.31 | 7.48 | 11.72 | 99.7 | 19.1 | 12.4
2007-05-19 | 17:00 | 8.29 | 7.48 | 11.72 | 99.7 | 19.1 | 12.4
2007-05-19 | 17:15 | 8.26 | 7.47 | 11.72 | 99.6 | 19.1 | 12.4
2007-05-19 | 17:30 | 8.23 | 7.48 | 11.72 | 99.5 | 19.1 | 12.4
2007-05-19 | 17:45 | 8.17 | 7.47 | 11.73 | 99.5 | 19.1 | 12.4
2007-05-19 | 18:00 | 8.12 | 7.47 | 11.75 | 99.5 | 19 | 12.4
2007-05-19 | 18:15 | 8.1 | 7.47 | 11.72 | 99.2 | 19 | 12.3
2007-05-19 | 18:30 | 8.13 | 7.48 | 11.72 | 99.3 | 19 | 12.3
2007-05-19 | 18:45 | 8.13 | 7.47 | 11.71 | 99.2 | 18.9 | 12.3
2007-05-19 | 19:00 | 8.14 | 7.48 | 11.7 | 99.1 | 18.9 | 12.3
2007-05-19 | 19:15 | 8.14 | 7.47 | 11.71 | 99.2 | 18.8 | 12.2
2007-05-19 | 19:30 | 8.17 | 7.46 | 11.69 | 99.1 | 18.8 | 12.2
2007-05-19 | 19:45 | 8.21 | 7.47 | 11.67 | 99 | 18.8 | 12.2
2007-05-19 | 20:00 | 8.22 | 7.46 | 11.67 | 99.1 | 18.8 | 12.2
2007-05-19 | 20:15 | 8.17 | 7.47 | 11.69 | 99.1 | 18.7 | 12.2
2007-05-19 | 20:30 | 8.07 | 7.46 | 11.72 | 99.1 | 18.7 | 12.1
2007-05-19 | 20:45 | 7.94 | 7.45 | 11.74 | 99 | 18.7 | 12.1
2007-05-19 | 21:00 | 7.79 | 7.44 | 11.8 | 99.1 | 18.6 | 12.1
2007-05-19 | 21:15 | 7.61 | 7.44 | 11.85 | 99.1 | 18.5 | 12
2007-05-19 | 21:30 | 7.42 | 7.44 | 11.9 | 99 | 18.5 | 12
2007-05-19 | 21:45 | 7.22 | 7.44 | 11.95 | 99 | 18.5 | 12
2007-05-19 | 22:00 | 7.04 | 7.43 | 12.01 | 99.1 | 18.4 | 12
2007-05-19 | 22:15 | 6.85 | 7.42 | 12.07 | 99 | 18.4 | 12
2007-05-19 | 22:30 | 6.68 | 7.42 | 12.12 | 99.1 | 18.4 | 11.9
2007-05-19 | 22:45 | 6.51 | 7.41 | 12.17 | 99 | 18.4 | 11.9
2007-05-19 | 23:00 | 6.35 | 7.41 | 12.22 | 99 | 18.4 | 11.9
2007-05-19 | 23:15 | 6.19 | 7.4 | 12.27 | 99 | 18.4 | 11.9
2007-05-19 | 23:30 | 6.03 | 7.41 | 12.31 | 99 | 18.4 | 11.9
2007-05-19 | 23:45 | 5.88 | 7.4 | 12.37 | 99 | 18.4 | 11.9
2007-05-20 | 0:00 | 5.74 | 7.39 | 12.4 | 98.9 | 18.4 | 11.9
2007-05-20 | 0:15 | 5.6 | 7.39 | 12.44 | 98.9 | 18.4 | 11.9
2007-05-20 | 0:30 | 5.48 | 7.38 | 12.49 | 99 | 18.3 | 11.9
2007-05-20 | 0:45 | 5.36 | 7.38 | 12.52 | 98.9 | 18.4 | 11.9
2007-05-20 | 1:01 | 5.25 | 7.38 | 12.56 | 98.9 | 18.4 | 11.9
2007-05-20 | 1:16 | 5.14 | 7.38 | 12.6 | 99 | 18.3 | 11.9
2007-05-20 | 1:31 | 5.04 | 7.38 | 12.62 | 98.9 | 18.4 | 11.9
2007-05-20 | 1:46 | 4.95 | 7.38 | 12.64 | 98.9 | 18.4 | 12
2007-05-20 | 2:01 | 4.87 | 7.38 | 12.67 | 98.9 | 18.4 | 12
2007-05-20 | 2:16 | 4.79 | 7.38 | 12.68 | 98.7 | 18.4 | 12
2007-05-20 | 2:31 | 4.71 | 7.38 | 12.73 | 98.9 | 18.4 | 12
2007-05-20 | 2:46 | 4.62 | 7.37 | 12.75 | 98.8 | 18.5 | 12
2007-05-20 | 3:01 | 4.55 | 7.38 | 12.75 | 98.7 | 18.5 | 12
2007-05-20 | 3:16 | 4.48 | 7.38 | 12.79 | 98.8 | 18.4 | 12
2007-05-20 | 3:31 | 4.42 | 7.38 | 12.81 | 98.8 | 18.4 | 12
2007-05-20 | 3:46 | 4.37 | 7.39 | 12.83 | 98.8 | 18.5 | 12
2007-05-20 | 4:01 | 4.32 | 7.38 | 12.84 | 98.8 | 18.5 | 12
2007-05-20 | 4:16 | 4.27 | 7.39 | 12.87 | 98.9 | 18.5 | 12
2007-05-20 | 4:31 | 4.23 | 7.39 | 12.9 | 99 | 18.6 | 12.1
2007-05-20 | 4:46 | 4.18 | 7.39 | 12.89 | 98.8 | 18.6 | 12.1
2007-05-20 | 5:01 | 4.14 | 7.38 | 12.89 | 98.7 | 18.6 | 12.1
2007-05-20 | 5:16 | 4.11 | 7.39 | 12.91 | 98.8 | 18.6 | 12.1
2007-05-20 | 5:31 | 4.07 | 7.39 | 12.92 | 98.7 | 18.6 | 12.1
2007-05-20 | 5:46 | 4.04 | 7.38 | 12.92 | 98.7 | 18.6 | 12.1
2007-05-20 | 6:01 | 4.01 | 7.39 | 12.93 | 98.6 | 18.6 | 12.1
2007-05-20 | 6:16 | 3.99 | 7.38 | 12.95 | 98.7 | 18.7 | 12.1
2007-05-20 | 6:31 | 3.97 | 7.38 | 12.94 | 98.7 | 18.7 | 12.2
2007-05-20 | 6:46 | 3.96 | 7.38 | 12.96 | 98.7 | 18.7 | 12.1
2007-05-20 | 7:01 | 3.95 | 7.39 | 12.97 | 98.8 | 18.7 | 12.1
2007-05-20 | 7:16 | 3.95 | 7.4 | 12.96 | 98.7 | 18.7 | 12.2
2007-05-20 | 7:31 | 3.96 | 7.4 | 12.98 | 98.9 | 18.7 | 12.1
2007-05-20 | 7:46 | 3.98 | 7.41 | 12.95 | 98.8 | 18.7 | 12.1
2007-05-20 | 8:01 | 4 | 7.42 | 12.95 | 98.8 | 18.7 | 12.1
2007-05-20 | 8:16 | 4.02 | 7.42 | 12.95 | 98.9 | 18.7 | 12.2
2007-05-20 | 8:31 | 4.03 | 7.42 | 12.94 | 98.8 | 18.7 | 12.2
2007-05-20 | 8:46 | 4.05 | 7.42 | 12.92 | 98.7 | 18.8 | 12.2
2007-05-20 | 9:01 | 4.08 | 7.43 | 12.94 | 98.9 | 18.8 | 12.2
2007-05-20 | 9:16 | 4.11 | 7.43 | 12.91 | 98.7 | 18.8 | 12.2
2007-05-20 | 9:31 | 4.13 | 7.43 | 12.89 | 98.7 | 18.8 | 12.2
2007-05-20 | 9:46 | 4.2 | 7.42 | 12.88 | 98.8 | 18.8 | 12.2
2007-05-20 | 10:01 | 4.27 | 7.43 | 12.86 | 98.8 | 18.8 | 12.2
2007-05-20 | 10:16 | 4.33 | 7.44 | 12.83 | 98.7 | 18.7 | 12.1
2007-05-20 | 10:31 | 4.37 | 7.43 | 12.81 | 98.6 | 18.7 | 12.1
2007-05-20 | 10:46 | 4.4 | 7.44 | 12.81 | 98.8 | 18.7 | 12.1
2007-05-20 | 11:01 | 4.42 | 7.43 | 12.8 | 98.7 | 18.8 | 12.2
2007-05-20 | 11:16 | 4.43 | 7.43 | 12.79 | 98.7 | 18.8 | 12.2
2007-05-20 | 11:31 | 4.43 | 7.43 | 12.79 | 98.7 | 18.7 | 12.2
2007-05-20 | 11:46 | 4.44 | 7.43 | 12.79 | 98.7 | 18.7 | 12.2
2007-05-20 | 12:01 | 4.49 | 7.44 | 12.81 | 98.9 | 18.7 | 12.2
2007-05-20 | 12:16 | 4.58 | 7.44 | 12.74 | 98.7 | 18.7 | 12.2
2007-05-20 | 12:31 | 4.67 | 7.44 | 12.71 | 98.7 | 18.7 | 12.1
2007-05-20 | 12:46 | 4.76 | 7.44 | 12.68 | 98.7 | 18.7 | 12.1
2007-05-20 | 13:00 | 4.8 | 7.44 | 12.67 | 98.7 | 20.1 | 13
2007-05-20 | 13:15 | 4.88 | 7.44 | 12.66 | 98.8 | 20.3 | 13.2
2007-05-20 | 13:30 | 5.07 | 7.45 | 12.61 | 98.9 | 20.4 | 13.2
2007-05-20 | 13:45 | 5.17 | 7.44 | 12.57 | 98.8 | 20.5 | 13.3
2007-05-20 | 14:00 | 5.24 | 7.45 | 12.53 | 98.7 | 20.6 | 13.4
2007-05-20 | 14:15 | 5.22 | 7.45 | 12.55 | 98.8 | 20.6 | 13.4
2007-05-20 | 14:30 | 5.22 | 7.45 | 12.56 | 98.9 | 20.6 | 13.4
2007-05-20 | 14:45 | 5.37 | 7.46 | 12.51 | 98.9 | 20.6 | 13.4
2007-05-20 | 15:00 | 5.38 | 7.45 | 12.5 | 98.9 | 20.6 | 13.4
2007-05-20 | 15:15 | 5.36 | 7.45 | 12.5 | 98.8 | 20.7 | 13.5
2007-05-20 | 15:30 | 5.35 | 7.45 | 12.49 | 98.7 | 20.7 | 13.5
2007-05-20 | 15:45 | 5.41 | 7.45 | 12.47 | 98.7 | 20.8 | 13.5
2007-05-20 | 16:00 | 5.43 | 7.46 | 12.45 | 98.6 | 20.8 | 13.5
2007-05-20 | 16:15 | 5.51 | 7.45 | 12.43 | 98.6 | 20.8 | 13.5
2007-05-20 | 16:30 | 5.63 | 7.46 | 12.38 | 98.6 | 20.8 | 13.5
2007-05-20 | 16:45 | 5.82 | 7.45 | 12.34 | 98.7 | 20.8 | 13.5
2007-05-20 | 17:00 | 6.08 | 7.45 | 12.26 | 98.7 | 20.9 | 13.6
2007-05-20 | 17:15 | 6.42 | 7.46 | 12.16 | 98.8 | 20.9 | 13.6
2007-05-20 | 17:30 | 6.74 | 7.46 | 12.06 | 98.7 | 20.9 | 13.6
2007-05-20 | 17:45 | 7.09 | 7.47 | 11.95 | 98.7 | 20.9 | 13.6
2007-05-20 | 18:00 | 7.41 | 7.47 | 11.85 | 98.6 | 21.2 | 13.8
2007-05-20 | 18:15 | 7.66 | 7.48 | 11.77 | 98.5 | 21.3 | 13.8
2007-05-20 | 18:30 | 7.86 | 7.48 | 11.7 | 98.5 | 21.2 | 13.8
2007-05-20 | 18:45 | 7.94 | 7.48 | 11.68 | 98.4 | 21.3 | 13.8
2007-05-20 | 19:00 | 7.86 | 7.47 | 11.69 | 98.4 | 21.3 | 13.8
2007-05-20 | 19:15 | 7.7 | 7.47 | 11.68 | 97.9 | 21.3 | 13.8
2007-05-20 | 19:30 | 7.52 | 7.46 | 11.77 | 98.2 | 21.4 | 13.9
2007-05-20 | 19:45 | 7.31 | 7.45 | 11.81 | 98 | 21.4 | 13.9
2007-05-20 | 20:00 | 7.11 | 7.45 | 11.89 | 98.2 | 21.4 | 13.9
2007-05-20 | 20:15 | 6.94 | 7.44 | 11.96 | 98.4 | 21.3 | 13.8
2007-05-20 | 20:30 | 6.79 | 7.44 | 12.01 | 98.4 | 21.2 | 13.8
2007-05-20 | 20:45 | 6.62 | 7.44 | 12.06 | 98.4 | 21.1 | 13.7
2007-05-20 | 21:00 | 6.47 | 7.44 | 12.11 | 98.5 | 21.2 | 13.8
2007-05-20 | 21:15 | 6.32 | 7.44 | 12.16 | 98.5 | 21.2 | 13.8
2007-05-20 | 21:30 | 6.18 | 7.43 | 12.2 | 98.5 | 21.2 | 13.8
2007-05-20 | 21:45 | 6.03 | 7.43 | 12.26 | 98.5 | 21.2 | 13.8
2007-05-20 | 22:00 | 5.89 | 7.43 | 12.3 | 98.5 | 21.2 | 13.8
2007-05-20 | 22:15 | 5.75 | 7.42 | 12.34 | 98.5 | 21.1 | 13.7
2007-05-20 | 22:30 | 5.62 | 7.42 | 12.39 | 98.6 | 21.1 | 13.7
2007-05-20 | 22:45 | 5.49 | 7.42 | 12.43 | 98.5 | 21.1 | 13.7
2007-05-20 | 23:00 | 5.35 | 7.42 | 12.46 | 98.4 | 21.1 | 13.7
2007-05-20 | 23:15 | 5.2 | 7.42 | 12.52 | 98.5 | 21.1 | 13.7
2007-05-20 | 23:30 | 5.06 | 7.42 | 12.56 | 98.5 | 21.1 | 13.7
2007-05-20 | 23:45 | 4.93 | 7.42 | 12.61 | 98.6 | 21.1 | 13.7
2007-05-21 | 0:00 | 4.79 | 7.42 | 12.67 | 98.6 | 21.1 | 13.7
2007-05-21 | 0:15 | 4.66 | 7.42 | 12.7 | 98.6 | 21 | 13.7
2007-05-21 | 0:30 | 4.53 | 7.41 | 12.73 | 98.4 | 21 | 13.7
2007-05-21 | 0:45 | 4.41 | 7.41 | 12.77 | 98.5 | 21 | 13.7
2007-05-21 | 1:01 | 4.3 | 7.41 | 12.83 | 98.6 | 21 | 13.6
2007-05-21 | 1:16 | 4.17 | 7.41 | 12.86 | 98.5 | 21 | 13.6
2007-05-21 | 1:31 | 4.05 | 7.41 | 12.9 | 98.5 | 21 | 13.6
2007-05-21 | 1:46 | 3.92 | 7.41 | 12.94 | 98.5 | 21 | 13.6
2007-05-21 | 2:01 | 3.8 | 7.41 | 12.99 | 98.6 | 21 | 13.6
2007-05-21 | 2:16 | 3.68 | 7.41 | 13.01 | 98.4 | 21 | 13.6
2007-05-21 | 2:31 | 3.57 | 7.41 | 13.08 | 98.6 | 21 | 13.6
2007-05-21 | 2:46 | 3.45 | 7.4 | 13.14 | 98.8 | 21 | 13.6
2007-05-21 | 3:01 | 3.34 | 7.4 | 13.15 | 98.6 | 21 | 13.6
2007-05-21 | 3:16 | 3.24 | 7.4 | 13.19 | 98.6 | 21 | 13.6
2007-05-21 | 3:31 | 3.14 | 7.4 | 13.22 | 98.6 | 21 | 13.7
2007-05-21 | 3:46 | 3.06 | 7.39 | 13.25 | 98.6 | 21 | 13.7
2007-05-21 | 4:01 | 3 | 7.39 | 13.29 | 98.7 | 21.1 | 13.7
2007-05-21 | 4:16 | 2.94 | 7.4 | 13.29 | 98.6 | 21.1 | 13.7
2007-05-21 | 4:31 | 2.89 | 7.4 | 13.31 | 98.6 | 21.1 | 13.7
2007-05-21 | 4:46 | 2.83 | 7.4 | 13.33 | 98.6 | 21.1 | 13.7
2007-05-21 | 5:01 | 2.77 | 7.4 | 13.35 | 98.5 | 21.1 | 13.7
2007-05-21 | 5:16 | 2.71 | 7.4 | 13.38 | 98.6 | 21.1 | 13.7
2007-05-21 | 5:31 | 2.65 | 7.4 | 13.41 | 98.7 | 21.2 | 13.8
2007-05-21 | 5:46 | 2.6 | 7.4 | 13.42 | 98.6 | 21.2 | 13.8
2007-05-21 | 6:01 | 2.57 | 7.39 | 13.42 | 98.5 | 21.2 | 13.8
2007-05-21 | 6:16 | 2.53 | 7.4 | 13.45 | 98.7 | 21.2 | 13.8
2007-05-21 | 6:31 | 2.49 | 7.4 | 13.46 | 98.6 | 21.3 | 13.8
2007-05-21 | 6:46 | 2.45 | 7.4 | 13.47 | 98.6 | 21.3 | 13.8
2007-05-21 | 7:01 | 2.4 | 7.4 | 13.5 | 98.7 | 21.3 | 13.9
2007-05-21 | 7:16 | 2.37 | 7.4 | 13.51 | 98.7 | 21.3 | 13.9
2007-05-21 | 7:31 | 2.35 | 7.41 | 13.51 | 98.6 | 21.4 | 13.9
2007-05-21 | 7:46 | 2.34 | 7.41 | 13.53 | 98.7 | 21.4 | 13.9
2007-05-21 | 8:01 | 2.33 | 7.4 | 13.53 | 98.7 | 21.4 | 13.9
2007-05-21 | 8:16 | 2.34 | 7.41 | 13.55 | 98.9 | 21.4 | 13.9
2007-05-21 | 8:31 | 2.35 | 7.4 | 13.52 | 98.7 | 21.4 | 13.9
2007-05-21 | 8:46 | 2.39 | 7.41 | 13.51 | 98.8 | 21.4 | 13.9
2007-05-21 | 9:01 | 2.44 | 7.41 | 13.51 | 98.8 | 21.4 | 13.9
2007-05-21 | 9:16 | 2.5 | 7.42 | 13.49 | 98.9 | 21.4 | 13.9
2007-05-21 | 9:31 | 2.58 | 7.42 | 13.46 | 98.9 | 21.4 | 13.9
2007-05-21 | 9:46 | 2.66 | 7.41 | 13.44 | 98.9 | 21.4 | 13.9
2007-05-21 | 10:01 | 2.77 | 7.42 | 13.4 | 98.9 | 21.5 | 14
2007-05-21 | 10:16 | 2.89 | 7.42 | 13.37 | 99 | 21.5 | 14
2007-05-21 | 10:31 | 3.02 | 7.43 | 13.33 | 99.1 | 21.5 | 13.9
2007-05-21 | 10:46 | 3.15 | 7.42 | 13.29 | 99.1 | 21.5 | 14
2007-05-21 | 11:01 | 3.31 | 7.44 | 13.24 | 99.2 | 21.5 | 14
2007-05-21 | 11:16 | 3.51 | 7.43 | 13.18 | 99.2 | 21.5 | 14
2007-05-21 | 11:31 | 3.71 | 7.44 | 13.11 | 99.2 | 21.5 | 14
2007-05-21 | 11:46 | 3.92 | 7.44 | 13.04 | 99.2 | 21.5 | 14
2007-05-21 | 12:01 | 4.15 | 7.44 | 12.97 | 99.3 | 21.5 | 14
2007-05-21 | 12:16 | 4.38 | 7.45 | 12.9 | 99.4 | 21.5 | 14
2007-05-21 | 12:31 | 4.63 | 7.45 | 12.82 | 99.4 | 21.5 | 14
2007-05-21 | 12:46 | 4.89 | 7.47 | 12.75 | 99.5 | 21.5 | 14
2007-05-21 | 13:00 | 5.15 | 7.46 | 12.65 | 99.5 | 21.6 | 14
2007-05-21 | 13:15 | 5.41 | 7.46 | 12.59 | 99.6 | 21.5 | 14
2007-05-21 | 13:30 | 5.69 | 7.46 | 12.5 | 99.6 | 21.5 | 14
2007-05-21 | 13:45 | 5.95 | 7.47 | 12.41 | 99.5 | 21.6 | 14
2007-05-21 | 14:00 | 6.2 | 7.47 | 12.36 | 99.8 | 21.6 | 14
2007-05-21 | 14:15 | 6.46 | 7.47 | 12.26 | 99.6 | 21.6 | 14
2007-05-21 | 14:30 | 6.71 | 7.48 | 12.18 | 99.6 | 21.6 | 14
2007-05-21 | 14:45 | 6.96 | 7.48 | 12.1 | 99.6 | 21.7 | 14.1
2007-05-21 | 15:00 | 7.2 | 7.48 | 12.02 | 99.5 | 21.7 | 14.1
2007-05-21 | 15:15 | 7.43 | 7.49 | 11.96 | 99.5 | 21.8 | 14.1
2007-05-21 | 15:30 | 7.61 | 7.49 | 11.91 | 99.6 | 21.7 | 14.1
2007-05-21 | 15:45 | 7.82 | 7.49 | 11.84 | 99.5 | 21.7 | 14.1
2007-05-21 | 16:00 | 7.9 | 7.49 | 11.82 | 99.5 | 21.6 | 14.1
2007-05-21 | 16:15 | 7.96 | 7.5 | 11.81 | 99.7 | 21.7 | 14.1
2007-05-21 | 16:30 | 7.98 | 7.49 | 11.8 | 99.6 | 21.7 | 14.1
2007-05-21 | 16:45 | 8.04 | 7.49 | 11.78 | 99.6 | 21.6 | 14
2007-05-21 | 17:00 | 8.09 | 7.49 | 11.76 | 99.6 | 21.6 | 14.1
2007-05-21 | 17:15 | 8.14 | 7.49 | 11.74 | 99.5 | 21.6 | 14.1
2007-05-21 | 17:30 | 8.16 | 7.5 | 11.75 | 99.6 | 21.7 | 14.1
2007-05-21 | 17:45 | 8.18 | 7.49 | 11.73 | 99.5 | 21.6 | 14.1
2007-05-21 | 18:00 | 8.17 | 7.49 | 11.72 | 99.4 | 21.8 | 14.1
2007-05-21 | 18:15 | 8.14 | 7.49 | 11.72 | 99.3 | 21.6 | 14.1
2007-05-21 | 18:30 | 8.07 | 7.49 | 11.73 | 99.2 | 21.7 | 14.1
2007-05-21 | 18:45 | 7.98 | 7.48 | 11.75 | 99.2 | 21.7 | 14.1
2007-05-21 | 19:00 | 7.87 | 7.49 | 11.79 | 99.2 | 21.8 | 14.1
2007-05-21 | 19:15 | 7.71 | 7.49 | 11.85 | 99.3 | 21.7 | 14.1
2007-05-21 | 19:30 | 7.53 | 7.48 | 11.88 | 99.1 | 21.7 | 14.1
2007-05-21 | 19:45 | 7.36 | 7.47 | 11.93 | 99.2 | 21.7 | 14.1
2007-05-21 | 20:00 | 7.2 | 7.47 | 11.96 | 99 | 21.7 | 14.1
2007-05-21 | 20:15 | 7.04 | 7.47 | 12.01 | 99.1 | 21.6 | 14.1
2007-05-21 | 20:30 | 6.89 | 7.47 | 12.06 | 99 | 21.6 | 14.1
2007-05-21 | 20:45 | 6.74 | 7.46 | 12.1 | 99 | 21.6 | 14.1
2007-05-21 | 21:00 | 6.6 | 7.46 | 12.15 | 99 | 21.6 | 14
2007-05-21 | 21:15 | 6.46 | 7.47 | 12.2 | 99.2 | 21.6 | 14
2007-05-21 | 21:30 | 6.33 | 7.45 | 12.23 | 99.1 | 21.5 | 14
2007-05-21 | 21:45 | 6.18 | 7.46 | 12.26 | 99 | 21.6 | 14
2007-05-21 | 22:00 | 6.03 | 7.46 | 12.32 | 99 | 21.5 | 14
2007-05-21 | 22:15 | 5.88 | 7.45 | 12.37 | 99 | 21.6 | 14
2007-05-21 | 22:30 | 5.73 | 7.45 | 12.42 | 99 | 21.6 | 14
2007-05-21 | 22:45 | 5.59 | 7.44 | 12.46 | 99 | 21.6 | 14
2007-05-21 | 23:00 | 5.45 | 7.44 | 12.5 | 99 | 21.5 | 14
2007-05-21 | 23:15 | 5.31 | 7.44 | 12.55 | 99.1 | 21.6 | 14
2007-05-21 | 23:30 | 5.19 | 7.44 | 12.58 | 99 | 21.6 | 14
2007-05-21 | 23:45 | 5.06 | 7.43 | 12.62 | 99 | 21.6 | 14
2007-05-22 | 0:00 | 4.94 | 7.44 | 12.67 | 99.1 | 21.6 | 14
2007-05-22 | 0:15 | 4.81 | 7.43 | 12.71 | 99 | 21.6 | 14
2007-05-22 | 0:30 | 4.69 | 7.43 | 12.75 | 99 | 21.5 | 14
2007-05-22 | 0:45 | 4.57 | 7.42 | 12.79 | 99 | 21.6 | 14
2007-05-22 | 1:01 | 4.46 | 7.43 | 12.82 | 99 | 21.5 | 14
2007-05-22 | 1:16 | 4.35 | 7.43 | 12.87 | 99 | 21.5 | 14
2007-05-22 | 1:31 | 4.22 | 7.42 | 12.91 | 99 | 21.5 | 14
2007-05-22 | 1:46 | 4.11 | 7.43 | 12.95 | 99.1 | 21.5 | 14
2007-05-22 | 2:01 | 3.98 | 7.42 | 12.99 | 99 | 21.5 | 14
2007-05-22 | 2:16 | 3.86 | 7.42 | 13.03 | 99.1 | 21.6 | 14
2007-05-22 | 2:31 | 3.74 | 7.41 | 13.06 | 99 | 21.6 | 14
2007-05-22 | 2:46 | 3.63 | 7.41 | 13.1 | 99 | 21.6 | 14
2007-05-22 | 3:01 | 3.51 | 7.42 | 13.15 | 99 | 21.5 | 14
2007-05-22 | 3:16 | 3.39 | 7.41 | 13.18 | 98.9 | 21.5 | 14
2007-05-22 | 3:31 | 3.29 | 7.41 | 13.22 | 99 | 21.5 | 14
2007-05-22 | 3:46 | 3.18 | 7.41 | 13.25 | 98.9 | 21.6 | 14
2007-05-22 | 4:01 | 3.09 | 7.41 | 13.3 | 99.1 | 21.6 | 14
2007-05-22 | 4:16 | 3.02 | 7.41 | 13.32 | 99 | 21.6 | 14
2007-05-22 | 4:31 | 2.94 | 7.41 | 13.34 | 99 | 21.6 | 14
2007-05-22 | 4:46 | 2.86 | 7.41 | 13.37 | 99 | 21.6 | 14
2007-05-22 | 5:01 | 2.79 | 7.4 | 13.37 | 98.8 | 21.6 | 14
2007-05-22 | 5:16 | 2.72 | 7.4 | 13.42 | 98.9 | 21.6 | 14
2007-05-22 | 5:31 | 2.66 | 7.4 | 13.43 | 98.8 | 21.6 | 14
2007-05-22 | 5:46 | 2.6 | 7.41 | 13.46 | 98.9 | 21.6 | 14
2007-05-22 | 6:01 | 2.54 | 7.41 | 13.48 | 98.9 | 21.7 | 14.1
2007-05-22 | 6:16 | 2.48 | 7.41 | 13.5 | 98.9 | 21.6 | 14.1
2007-05-22 | 6:31 | 2.42 | 7.41 | 13.51 | 98.8 | 21.6 | 14.1
2007-05-22 | 6:46 | 2.38 | 7.41 | 13.55 | 99 | 21.6 | 14.1
2007-05-22 | 7:01 | 2.33 | 7.42 | 13.57 | 99 | 21.6 | 14
2007-05-22 | 7:16 | 2.31 | 7.41 | 13.61 | 99.2 | 21.6 | 14
2007-05-22 | 7:31 | 2.29 | 7.42 | 13.6 | 99.1 | 21.6 | 14
2007-05-22 | 7:46 | 2.3 | 7.42 | 13.6 | 99.1 | 21.6 | 14
2007-05-22 | 8:01 | 2.31 | 7.43 | 13.6 | 99.1 | 21.6 | 14
2007-05-22 | 8:16 | 2.34 | 7.44 | 13.58 | 99.1 | 21.6 | 14
2007-05-22 | 8:31 | 2.38 | 7.43 | 13.57 | 99.2 | 21.6 | 14
2007-05-22 | 8:46 | 2.43 | 7.44 | 13.54 | 99 | 21.6 | 14
2007-05-22 | 9:01 | 2.5 | 7.44 | 13.53 | 99.2 | 21.6 | 14
2007-05-22 | 9:16 | 2.58 | 7.44 | 13.52 | 99.3 | 21.6 | 14
2007-05-22 | 9:31 | 2.67 | 7.45 | 13.49 | 99.3 | 21.6 | 14
2007-05-22 | 9:46 | 2.77 | 7.45 | 13.46 | 99.3 | 21.6 | 14
2007-05-22 | 10:01 | 2.88 | 7.44 | 13.42 | 99.3 | 21.5 | 14
2007-05-22 | 10:16 | 3.01 | 7.45 | 13.37 | 99.4 | 21.6 | 14
2007-05-22 | 10:31 | 3.15 | 7.45 | 13.32 | 99.4 | 21.5 | 14
2007-05-22 | 10:46 | 3.3 | 7.46 | 13.28 | 99.5 | 21.5 | 14
2007-05-22 | 11:01 | 3.48 | 7.46 | 13.18 | 99.2 | 21.5 | 14
2007-05-22 | 11:16 | 3.67 | 7.47 | 13.16 | 99.5 | 21.5 | 14
2007-05-22 | 11:31 | 3.87 | 7.48 | 13.1 | 99.6 | 21.5 | 14
2007-05-22 | 11:46 | 4.1 | 7.47 | 13.03 | 99.7 | 21.5 | 14
2007-05-22 | 12:01 | 4.34 | 7.48 | 12.96 | 99.7 | 21.5 | 14
2007-05-22 | 12:16 | 4.59 | 7.48 | 12.89 | 99.8 | 21.5 | 14
2007-05-22 | 12:31 | 4.85 | 7.48 | 12.81 | 99.9 | 21.5 | 14
2007-05-22 | 12:46 | 5.14 | 7.49 | 12.72 | 99.9 | 21.5 | 14
2007-05-22 | 13:00 | 5.43 | 7.49 | 12.61 | 99.9 | 21.5 | 14
2007-05-22 | 13:15 | 5.73 | 7.49 | 12.52 | 99.9 | 21.5 | 14
2007-05-22 | 13:30 | 6.03 | 7.5 | 12.41 | 99.7 | 21.4 | 13.9
2007-05-22 | 13:45 | 6.37 | 7.5 | 12.33 | 100 | 21.5 | 13.9
2007-05-22 | 14:00 | 6.69 | 7.5 | 12.25 | 100.1 | 21.5 | 14
2007-05-22 | 14:15 | 6.98 | 7.5 | 12.15 | 100 | 21.5 | 14
2007-05-22 | 14:30 | 7.19 | 7.51 | 12.08 | 100 | 21.5 | 14
2007-05-22 | 14:45 | 7.46 | 7.51 | 11.99 | 99.9 | 21.5 | 13.9
2007-05-22 | 15:00 | 7.72 | 7.52 | 11.94 | 100.1 | 21.5 | 14
2007-05-22 | 15:15 | 7.87 | 7.51 | 11.86 | 99.9 | 21.4 | 13.9
2007-05-22 | 15:30 | 8.04 | 7.52 | 11.8 | 99.7 | 21.4 | 13.9
2007-05-22 | 15:45 | 8.19 | 7.51 | 11.77 | 99.9 | 21.4 | 13.9
2007-05-22 | 16:00 | 8.39 | 7.52 | 11.72 | 99.9 | 21.4 | 13.9
2007-05-22 | 16:15 | 8.59 | 7.53 | 11.67 | 99.9 | 21.4 | 13.9
2007-05-22 | 16:30 | 8.72 | 7.53 | 11.63 | 99.9 | 21.4 | 13.9
2007-05-22 | 16:45 | 8.82 | 7.52 | 11.58 | 99.7 | 21.4 | 13.9
2007-05-22 | 17:00 | 8.93 | 7.52 | 11.55 | 99.7 | 21.4 | 13.9
2007-05-22 | 17:15 | 9.02 | 7.53 | 11.54 | 99.8 | 21.4 | 13.9
2007-05-22 | 17:30 | 9.05 | 7.53 | 11.5 | 99.6 | 21.4 | 13.9
2007-05-22 | 17:45 | 9.1 | 7.53 | 11.5 | 99.7 | 21.4 | 13.9
2007-05-22 | 18:00 | 9.05 | 7.52 | 11.5 | 99.6 | 21.4 | 13.9
2007-05-22 | 18:15 | 8.95 | 7.52 | 11.52 | 99.5 | 21.4 | 13.9
2007-05-22 | 18:30 | 8.84 | 7.51 | 11.55 | 99.5 | 21.4 | 13.9
2007-05-22 | 18:45 | 8.76 | 7.52 | 11.58 | 99.6 | 21.4 | 13.9
2007-05-22 | 19:00 | 8.68 | 7.52 | 11.6 | 99.6 | 21.4 | 13.9
2007-05-22 | 19:15 | 8.57 | 7.51 | 11.62 | 99.4 | 21.4 | 13.9
2007-05-22 | 19:30 | 8.41 | 7.5 | 11.66 | 99.4 | 21.4 | 13.9
2007-05-22 | 19:45 | 8.26 | 7.5 | 11.64 | 98.9 | 21.4 | 13.9
2007-05-22 | 20:00 | 8.1 | 7.5 | 11.74 | 99.3 | 21.5 | 14
2007-05-22 | 20:15 | 7.95 | 7.49 | 11.78 | 99.3 | 21.5 | 14
2007-05-22 | 20:30 | 7.79 | 7.48 | 11.84 | 99.4 | 21.5 | 14
2007-05-22 | 20:45 | 7.63 | 7.48 | 11.88 | 99.4 | 21.6 | 14
2007-05-22 | 21:00 | 7.47 | 7.47 | 11.93 | 99.4 | 21.5 | 14
2007-05-22 | 21:15 | 7.32 | 7.47 | 11.96 | 99.3 | 21.6 | 14
2007-05-22 | 21:30 | 7.18 | 7.47 | 12.01 | 99.4 | 21.5 | 14
2007-05-22 | 21:45 | 7.04 | 7.45 | 12.06 | 99.4 | 21.5 | 14
2007-05-22 | 22:00 | 6.89 | 7.46 | 12.11 | 99.5 | 21.5 | 14
2007-05-22 | 22:15 | 6.72 | 7.45 | 12.16 | 99.5 | 21.4 | 13.9
2007-05-22 | 22:30 | 6.56 | 7.44 | 12.21 | 99.5 | 21.4 | 13.9
2007-05-22 | 22:45 | 6.39 | 7.44 | 12.27 | 99.5 | 21.4 | 13.9
2007-05-22 | 23:00 | 6.24 | 7.44 | 12.31 | 99.5 | 21.4 | 13.9
2007-05-22 | 23:15 | 6.09 | 7.44 | 12.35 | 99.4 | 21.4 | 13.9
2007-05-22 | 23:30 | 5.94 | 7.43 | 12.39 | 99.3 | 21.4 | 13.9
2007-05-22 | 23:45 | 5.8 | 7.43 | 12.45 | 99.5 | 21.3 | 13.9
2007-05-23 | 0:00 | 5.66 | 7.43 | 12.49 | 99.5 | 21.3 | 13.9
2007-05-23 | 0:15 | 5.53 | 7.43 | 12.52 | 99.3 | 21.3 | 13.9
2007-05-23 | 0:30 | 5.4 | 7.43 | 12.57 | 99.5 | 21.3 | 13.8
2007-05-23 | 0:45 | 5.27 | 7.43 | 12.61 | 99.4 | 21.3 | 13.8
2007-05-23 | 1:01 | 5.14 | 7.43 | 12.65 | 99.4 | 21.2 | 13.8
2007-05-23 | 1:16 | 5.02 | 7.43 | 12.69 | 99.4 | 21.3 | 13.8
2007-05-23 | 1:31 | 4.89 | 7.43 | 12.73 | 99.4 | 21.2 | 13.8
2007-05-23 | 1:46 | 4.77 | 7.42 | 12.76 | 99.3 | 21.3 | 13.8
2007-05-23 | 2:01 | 4.65 | 7.43 | 12.81 | 99.4 | 21.3 | 13.8
2007-05-23 | 2:16 | 4.54 | 7.43 | 12.84 | 99.4 | 21.3 | 13.8
2007-05-23 | 2:31 | 4.42 | 7.42 | 12.88 | 99.3 | 21.3 | 13.8
2007-05-23 | 2:46 | 4.31 | 7.43 | 12.92 | 99.3 | 21.3 | 13.8
2007-05-23 | 3:01 | 4.2 | 7.42 | 12.95 | 99.3 | 21.3 | 13.8
2007-05-23 | 3:16 | 4.09 | 7.42 | 12.98 | 99.3 | 21.3 | 13.8
2007-05-23 | 3:31 | 3.98 | 7.42 | 13.01 | 99.2 | 21.2 | 13.8
2007-05-23 | 3:46 | 3.88 | 7.42 | 13.05 | 99.2 | 21.2 | 13.8
2007-05-23 | 4:01 | 3.78 | 7.42 | 13.09 | 99.3 | 21.2 | 13.8
2007-05-23 | 4:16 | 3.69 | 7.42 | 13.11 | 99.2 | 21.3 | 13.8
2007-05-23 | 4:31 | 3.61 | 7.41 | 13.15 | 99.3 | 21.3 | 13.8
2007-05-23 | 4:46 | 3.52 | 7.41 | 13.18 | 99.3 | 21.3 | 13.8
2007-05-23 | 5:01 | 3.44 | 7.41 | 13.2 | 99.2 | 21.2 | 13.8
2007-05-23 | 5:16 | 3.36 | 7.41 | 13.22 | 99.2 | 21.2 | 13.8
2007-05-23 | 5:31 | 3.29 | 7.42 | 13.21 | 98.9 | 21.2 | 13.8
2007-05-23 | 5:46 | 3.22 | 7.41 | 13.27 | 99.2 | 21.2 | 13.8
2007-05-23 | 6:01 | 3.16 | 7.41 | 13.3 | 99.2 | 21.2 | 13.8
2007-05-23 | 6:16 | 3.1 | 7.41 | 13.32 | 99.2 | 21.2 | 13.8
2007-05-23 | 6:31 | 3.05 | 7.41 | 13.34 | 99.2 | 21.2 | 13.8
2007-05-23 | 6:46 | 3.01 | 7.41 | 13.36 | 99.3 | 21.2 | 13.8
2007-05-23 | 7:01 | 2.98 | 7.42 | 13.31 | 98.9 | 21.1 | 13.7
2007-05-23 | 7:16 | 2.97 | 7.42 | 13.37 | 99.2 | 21.1 | 13.7
2007-05-23 | 7:31 | 2.96 | 7.42 | 13.36 | 99.2 | 21.1 | 13.7
2007-05-23 | 7:46 | 2.97 | 7.43 | 13.37 | 99.2 | 21.1 | 13.7
2007-05-23 | 8:01 | 2.98 | 7.43 | 13.39 | 99.4 | 21.1 | 13.7
2007-05-23 | 8:16 | 3.01 | 7.44 | 13.36 | 99.3 | 21.1 | 13.7
2007-05-23 | 8:31 | 3.04 | 7.45 | 13.37 | 99.4 | 21.1 | 13.7
2007-05-23 | 8:46 | 3.08 | 7.46 | 13.35 | 99.4 | 21.1 | 13.7
2007-05-23 | 9:01 | 3.15 | 7.47 | 13.35 | 99.6 | 21 | 13.7
2007-05-23 | 9:16 | 3.24 | 7.46 | 13.28 | 99.3 | 21 | 13.6
2007-05-23 | 9:31 | 3.34 | 7.47 | 13.26 | 99.4 | 21 | 13.6
2007-05-23 | 9:46 | 3.48 | 7.47 | 13.23 | 99.6 | 21 | 13.7
2007-05-23 | 10:01 | 3.61 | 7.47 | 13.17 | 99.5 | 21 | 13.7
2007-05-23 | 10:16 | 3.78 | 7.48 | 13.12 | 99.5 | 21.1 | 13.7
2007-05-23 | 10:31 | 3.95 | 7.48 | 13.05 | 99.4 | 21 | 13.7
2007-05-23 | 10:46 | 4.12 | 7.47 | 12.99 | 99.4 | 21.1 | 13.7
2007-05-23 | 11:01 | 4.29 | 7.49 | 12.95 | 99.5 | 21 | 13.7
2007-05-23 | 11:16 | 4.46 | 7.49 | 12.9 | 99.6 | 21 | 13.7
2007-05-23 | 11:31 | 4.67 | 7.49 | 12.83 | 99.6 | 21.1 | 13.7
2007-05-23 | 11:46 | 4.89 | 7.49 | 12.77 | 99.7 | 21.1 | 13.7
2007-05-23 | 12:01 | 5.06 | 7.49 | 12.72 | 99.8 | 21.1 | 13.7
2007-05-23 | 12:16 | 5.26 | 7.5 | 12.64 | 99.6 | 21.1 | 13.7
2007-05-23 | 12:31 | 5.51 | 7.5 | 12.58 | 99.8 | 21.1 | 13.7
2007-05-23 | 12:46 | 5.8 | 7.51 | 12.47 | 99.7 | 21.1 | 13.7
2007-05-23 | 13:00 | 6.14 | 7.51 | 12.37 | 99.7 | 21.1 | 13.7
2007-05-23 | 13:15 | 6.43 | 7.51 | 12.28 | 99.7 | 21.2 | 13.8
2007-05-23 | 13:30 | 6.71 | 7.52 | 12.2 | 99.7 | 21.2 | 13.8
2007-05-23 | 13:45 | 7.02 | 7.53 | 12.1 | 99.8 | 21.2 | 13.8
2007-05-23 | 14:00 | 7.35 | 7.52 | 12.01 | 99.8 | 21.1 | 13.7
2007-05-23 | 14:15 | 7.67 | 7.53 | 11.92 | 99.8 | 21.2 | 13.7
2007-05-23 | 14:30 | 7.95 | 7.53 | 11.82 | 99.7 | 21.2 | 13.7
2007-05-23 | 14:45 | 8.22 | 7.53 | 11.74 | 99.7 | 21.2 | 13.8
2007-05-23 | 15:00 | 8.47 | 7.54 | 11.66 | 99.6 | 21.2 | 13.8
2007-05-23 | 15:15 | 8.7 | 7.54 | 11.59 | 99.5 | 21.2 | 13.8
2007-05-23 | 15:30 | 8.95 | 7.54 | 11.53 | 99.6 | 21.2 | 13.8
2007-05-23 | 15:45 | 9.19 | 7.55 | 11.45 | 99.5 | 21.2 | 13.8
2007-05-23 | 16:00 | 9.39 | 7.54 | 11.39 | 99.5 | 21.2 | 13.8
2007-05-23 | 16:15 | 9.54 | 7.55 | 11.37 | 99.7 | 21.2 | 13.8
2007-05-23 | 16:30 | 9.61 | 7.55 | 11.34 | 99.5 | 21.3 | 13.8
2007-05-23 | 16:45 | 9.67 | 7.56 | 11.31 | 99.4 | 21.2 | 13.8
2007-05-23 | 17:00 | 9.8 | 7.55 | 11.28 | 99.5 | 21.2 | 13.8
2007-05-23 | 17:15 | 9.87 | 7.55 | 11.25 | 99.4 | 21.3 | 13.8
2007-05-23 | 17:30 | 9.87 | 7.54 | 11.24 | 99.3 | 21.2 | 13.8
2007-05-23 | 17:45 | 9.9 | 7.56 | 11.22 | 99.2 | 21.2 | 13.8
2007-05-23 | 18:00 | 9.93 | 7.55 | 11.22 | 99.2 | 21.2 | 13.8
2007-05-23 | 18:15 | 9.93 | 7.55 | 11.22 | 99.2 | 21.2 | 13.8
2007-05-23 | 18:30 | 9.87 | 7.55 | 11.23 | 99.2 | 21.2 | 13.8
2007-05-23 | 18:45 | 9.78 | 7.54 | 11.26 | 99.2 | 21.2 | 13.8
2007-05-23 | 19:00 | 9.67 | 7.54 | 11.28 | 99.2 | 21.2 | 13.8
2007-05-23 | 19:15 | 9.53 | 7.53 | 11.26 | 98.7 | 21.2 | 13.8
2007-05-23 | 19:30 | 9.39 | 7.53 | 11.35 | 99.1 | 21.1 | 13.7
2007-05-23 | 19:45 | 9.25 | 7.53 | 11.39 | 99.1 | 21.1 | 13.7
2007-05-23 | 20:00 | 9.12 | 7.52 | 11.42 | 99.1 | 21.1 | 13.7
2007-05-23 | 20:15 | 9.02 | 7.52 | 11.46 | 99.1 | 21.1 | 13.7
2007-05-23 | 20:30 | 8.92 | 7.52 | 11.48 | 99.1 | 21.1 | 13.7
2007-05-23 | 20:45 | 8.81 | 7.52 | 11.51 | 99.1 | 21 | 13.7
2007-05-23 | 21:00 | 8.71 | 7.52 | 11.55 | 99.2 | 21 | 13.7
2007-05-23 | 21:15 | 8.61 | 7.51 | 11.57 | 99.1 | 20.9 | 13.6
2007-05-23 | 21:30 | 8.5 | 7.51 | 11.6 | 99.1 | 20.9 | 13.6
2007-05-23 | 21:45 | 8.39 | 7.51 | 11.63 | 99.2 | 20.9 | 13.6
2007-05-23 | 22:00 | 8.27 | 7.5 | 11.66 | 99.1 | 20.8 | 13.5
2007-05-23 | 22:15 | 8.16 | 7.5 | 11.71 | 99.2 | 20.8 | 13.5
2007-05-23 | 22:30 | 8.06 | 7.5 | 11.76 | 99.4 | 20.8 | 13.5
2007-05-23 | 22:45 | 7.95 | 7.5 | 11.75 | 99.1 | 20.8 | 13.5
2007-05-23 | 23:00 | 7.83 | 7.5 | 11.79 | 99.1 | 20.8 | 13.5
2007-05-23 | 23:15 | 7.72 | 7.49 | 11.82 | 99.1 | 20.8 | 13.5
2007-05-23 | 23:30 | 7.6 | 7.48 | 11.85 | 99.1 | 20.7 | 13.5
2007-05-23 | 23:45 | 7.48 | 7.48 | 11.89 | 99.1 | 20.7 | 13.5
2007-05-24 | 0:00 | 7.36 | 7.48 | 11.91 | 99 | 20.7 | 13.5
2007-05-24 | 0:15 | 7.24 | 7.46 | 11.96 | 99.1 | 20.7 | 13.5
2007-05-24 | 0:30 | 7.12 | 7.47 | 11.99 | 99.1 | 20.6 | 13.4
2007-05-24 | 0:45 | 7.01 | 7.47 | 12.03 | 99.1 | 20.6 | 13.4
2007-05-24 | 1:01 | 6.89 | 7.46 | 12.07 | 99.1 | 20.6 | 13.4
2007-05-24 | 1:16 | 6.77 | 7.46 | 12.1 | 99.1 | 20.6 | 13.4
2007-05-24 | 1:31 | 6.66 | 7.45 | 12.12 | 99 | 20.6 | 13.4
2007-05-24 | 1:46 | 6.55 | 7.45 | 12.17 | 99.1 | 20.5 | 13.4
2007-05-24 | 2:01 | 6.45 | 7.45 | 12.19 | 99.1 | 20.6 | 13.4
2007-05-24 | 2:16 | 6.37 | 7.44 | 12.21 | 99 | 20.6 | 13.4
2007-05-24 | 2:31 | 6.29 | 7.44 | 12.23 | 99 | 20.5 | 13.3
2007-05-24 | 2:46 | 6.22 | 7.44 | 12.26 | 99 | 20.5 | 13.3
2007-05-24 | 3:01 | 6.15 | 7.44 | 12.28 | 99.1 | 20.5 | 13.3
2007-05-24 | 3:16 | 6.09 | 7.44 | 12.3 | 99 | 20.5 | 13.3
2007-05-24 | 3:31 | 6.02 | 7.44 | 12.32 | 99 | 20.4 | 13.3
2007-05-24 | 3:46 | 5.96 | 7.43 | 12.33 | 99 | 20.4 | 13.3
2007-05-24 | 4:01 | 5.9 | 7.43 | 12.36 | 99 | 20.4 | 13.2
2007-05-24 | 4:16 | 5.85 | 7.42 | 12.37 | 99 | 20.4 | 13.2
2007-05-24 | 4:31 | 5.79 | 7.42 | 12.34 | 98.6 | 20.3 | 13.2
2007-05-24 | 4:46 | 5.74 | 7.42 | 12.4 | 98.9 | 20.4 | 13.2
2007-05-24 | 5:01 | 5.7 | 7.42 | 12.42 | 99 | 20.3 | 13.2
2007-05-24 | 5:16 | 5.66 | 7.41 | 12.42 | 98.9 | 20.3 | 13.2
2007-05-24 | 5:31 | 5.62 | 7.41 | 12.46 | 99.1 | 20.2 | 13.1
2007-05-24 | 5:46 | 5.59 | 7.41 | 12.46 | 99 | 20.2 | 13.2
2007-05-24 | 6:01 | 5.56 | 7.41 | 12.47 | 99 | 20.2 | 13.1
2007-05-24 | 6:16 | 5.53 | 7.41 | 12.47 | 99 | 20.2 | 13.1
2007-05-24 | 6:31 | 5.5 | 7.42 | 12.49 | 99 | 20.2 | 13.1
2007-05-24 | 6:46 | 5.46 | 7.42 | 12.51 | 99.1 | 20.2 | 13.1
2007-05-24 | 7:01 | 5.42 | 7.42 | 12.51 | 99 | 20.2 | 13.1
2007-05-24 | 7:16 | 5.39 | 7.42 | 12.53 | 99.1 | 20.1 | 13.1
2007-05-24 | 7:31 | 5.39 | 7.43 | 12.55 | 99.2 | 20.1 | 13.1
2007-05-24 | 7:46 | 5.37 | 7.45 | 12.55 | 99.2 | 20.1 | 13.1
2007-05-24 | 8:01 | 5.36 | 7.45 | 12.54 | 99.1 | 20.1 | 13.1
2007-05-24 | 8:16 | 5.37 | 7.47 | 12.54 | 99.1 | 20.1 | 13.1
2007-05-24 | 8:31 | 5.36 | 7.47 | 12.54 | 99.1 | 20.1 | 13.1
2007-05-24 | 8:46 | 5.37 | 7.48 | 12.53 | 99.1 | 20.1 | 13.1
2007-05-24 | 9:01 | 5.38 | 7.48 | 12.54 | 99.1 | 20.1 | 13.1
2007-05-24 | 9:16 | 5.41 | 7.48 | 12.54 | 99.2 | 20.1 | 13
2007-05-24 | 9:31 | 5.46 | 7.48 | 12.5 | 99.1 | 20.1 | 13.1
2007-05-24 | 9:46 | 5.57 | 7.48 | 12.49 | 99.2 | 20 | 13
2007-05-24 | 10:01 | 5.65 | 7.48 | 12.47 | 99.3 | 20 | 13
2007-05-24 | 10:16 | 5.75 | 7.49 | 12.44 | 99.3 | 20.1 | 13.1
2007-05-24 | 10:31 | 5.79 | 7.49 | 12.4 | 99.1 | 20.1 | 13.1
2007-05-24 | 10:46 | 5.73 | 7.49 | 12.44 | 99.2 | 20.1 | 13
2007-05-24 | 11:01 | 5.8 | 7.49 | 12.42 | 99.2 | 20 | 13
2007-05-24 | 11:16 | 5.96 | 7.5 | 12.38 | 99.3 | 20 | 13
2007-05-24 | 11:31 | 6.21 | 7.5 | 12.29 | 99.3 | 20 | 13
2007-05-24 | 11:46 | 6.41 | 7.5 | 12.23 | 99.3 | 20 | 13
2007-05-24 | 12:01 | 6.59 | 7.5 | 12.19 | 99.4 | 20 | 13
2007-05-24 | 12:16 | 6.73 | 7.5 | 12.13 | 99.2 | 20 | 13
2007-05-24 | 12:31 | 6.89 | 7.51 | 12.09 | 99.3 | 20 | 13
2007-05-24 | 12:46 | 7.04 | 7.51 | 12.03 | 99.2 | 19.9 | 13
2007-05-24 | 13:00 | 7.48 | 7.52 | 11.94 | 99.6 | 19.9 | 13
2007-05-24 | 13:15 | 7.68 | 7.52 | 11.85 | 99.3 | 20 | 13
2007-05-24 | 13:30 | 7.78 | 7.52 | 11.81 | 99.2 | 20 | 13
2007-05-24 | 13:45 | 7.96 | 7.52 | 11.77 | 99.3 | 19.9 | 13
2007-05-24 | 14:00 | 8.31 | 7.54 | 11.68 | 99.4 | 20 | 13
2007-05-24 | 14:15 | 8.5 | 7.53 | 11.62 | 99.3 | 20 | 13
2007-05-24 | 14:30 | 9.01 | 7.54 | 11.51 | 99.6 | 20 | 13
2007-05-24 | 14:45 | 9.45 | 7.54 | 11.38 | 99.5 | 20 | 13
2007-05-24 | 15:00 | 9.84 | 7.53 | 11.27 | 99.5 | 20 | 13
2007-05-24 | 15:15 | 10.13 | 7.54 | 11.2 | 99.5 | 19.9 | 13
2007-05-24 | 15:30 | 10.36 | 7.55 | 11.13 | 99.5 | 20 | 13
2007-05-24 | 15:45 | 10.47 | 7.54 | 11.09 | 99.3 | 19.9 | 13
2007-05-24 | 16:00 | 10.46 | 7.56 | 11.09 | 99.3 | 19.9 | 13
2007-05-24 | 16:15 | 10.55 | 7.56 | 11.08 | 99.5 | 19.9 | 13
2007-05-24 | 16:30 | 10.76 | 7.55 | 11.02 | 99.4 | 19.9 | 12.9
2007-05-24 | 16:45 | 10.85 | 7.55 | 11 | 99.4 | 19.9 | 12.9
2007-05-24 | 17:00 | 10.75 | 7.55 | 11.02 | 99.4 | 19.9 | 12.9
2007-05-24 | 17:15 | 10.69 | 7.54 | 11.05 | 99.4 | 19.9 | 12.9
2007-05-24 | 17:30 | 10.73 | 7.55 | 11.02 | 99.3 | 19.9 | 12.9
2007-05-24 | 17:45 | 10.8 | 7.55 | 11 | 99.2 | 19.8 | 12.9
2007-05-24 | 18:00 | 11.01 | 7.54 | 10.94 | 99.2 | 19.8 | 12.9
2007-05-24 | 18:15 | 11.14 | 7.55 | 10.9 | 99.2 | 19.8 | 12.8
2007-05-24 | 18:30 | 11.13 | 7.55 | 10.9 | 99.1 | 19.7 | 12.8
2007-05-24 | 18:45 | 11.09 | 7.53 | 10.9 | 99 | 19.7 | 12.8
2007-05-24 | 19:00 | 11.04 | 7.53 | 10.94 | 99.3 | 19.6 | 12.8
2007-05-24 | 19:15 | 10.97 | 7.53 | 10.93 | 99 | 19.5 | 12.7
2007-05-24 | 19:30 | 10.83 | 7.53 | 10.98 | 99.1 | 19.5 | 12.7
2007-05-24 | 19:45 | 10.68 | 7.52 | 11 | 99 | 19.4 | 12.6
2007-05-24 | 20:00 | 10.53 | 7.51 | 11.05 | 99.1 | 19.3 | 12.5
2007-05-24 | 20:15 | 10.37 | 7.5 | 11.09 | 99.1 | 19.3 | 12.5
2007-05-24 | 20:30 | 10.22 | 7.5 | 11.13 | 99.1 | 19.2 | 12.5
2007-05-24 | 20:45 | 10.05 | 7.49 | 11.18 | 99.2 | 19.2 | 12.4
2007-05-24 | 21:00 | 9.88 | 7.47 | 11.21 | 99 | 19.1 | 12.4
2007-05-24 | 21:15 | 9.71 | 7.47 | 11.27 | 99.1 | 19.1 | 12.4
2007-05-24 | 21:30 | 9.53 | 7.46 | 11.31 | 99.1 | 19 | 12.4
2007-05-24 | 21:45 | 9.35 | 7.46 | 11.36 | 99.1 | 19 | 12.3
2007-05-24 | 22:00 | 9.16 | 7.45 | 11.41 | 99.1 | 19.2 | 12.5
2007-05-24 | 22:15 | 8.98 | 7.45 | 11.47 | 99.2 | 19.1 | 12.4
2007-05-24 | 22:30 | 8.79 | 7.45 | 11.51 | 99.1 | 19.1 | 12.4
2007-05-24 | 22:45 | 8.61 | 7.43 | 11.56 | 99 | 19.2 | 12.5
2007-05-24 | 23:00 | 8.43 | 7.43 | 11.62 | 99.1 | 19.1 | 12.4
2007-05-24 | 23:15 | 8.24 | 7.42 | 11.66 | 99 | 19.2 | 12.5
2007-05-24 | 23:30 | 8.06 | 7.42 | 11.73 | 99.2 | 19.1 | 12.4
2007-05-24 | 23:45 | 7.87 | 7.41 | 11.79 | 99.2 | 19.1 | 12.4
2007-05-25 | 0:00 | 7.69 | 7.41 | 11.83 | 99.1 | 19.2 | 12.5
2007-05-25 | 0:15 | 7.52 | 7.4 | 11.88 | 99.1 | 19.2 | 12.5
2007-05-25 | 0:30 | 7.35 | 7.4 | 11.94 | 99.2 | 18.8 | 12.2
2007-05-25 | 0:45 | 7.2 | 7.41 | 11.97 | 99.1 | 19 | 12.4
2007-05-25 | 1:01 | 7.04 | 7.39 | 12.03 | 99.2 | 19 | 12.3
2007-05-25 | 1:16 | 6.9 | 7.39 | 12.06 | 99.1 | 18.9 | 12.3
2007-05-25 | 1:31 | 6.77 | 7.4 | 12.1 | 99.1 | 18.9 | 12.3
2007-05-25 | 1:46 | 6.64 | 7.39 | 12.14 | 99.1 | 18.8 | 12.3
2007-05-25 | 2:01 | 6.51 | 7.4 | 12.19 | 99.2 | 18.8 | 12.2
2007-05-25 | 2:16 | 6.39 | 7.39 | 12.22 | 99.2 | 18.9 | 12.3
2007-05-25 | 2:31 | 6.27 | 7.39 | 12.25 | 99.1 | 18.7 | 12.1
2007-05-25 | 2:46 | 6.16 | 7.39 | 12.29 | 99.1 | 18.7 | 12.1
2007-05-25 | 3:01 | 6.06 | 7.39 | 12.32 | 99.1 | 18.7 | 12.1
2007-05-25 | 3:16 | 5.97 | 7.4 | 12.34 | 99.1 | 18.6 | 12.1
2007-05-25 | 3:31 | 5.88 | 7.39 | 12.36 | 99 | 18.6 | 12.1
2007-05-25 | 3:46 | 5.79 | 7.4 | 12.4 | 99.1 | 19 | 12.3
2007-05-25 | 4:01 | 5.7 | 7.39 | 12.43 | 99.1 | 18.9 | 12.3
2007-05-25 | 4:16 | 5.61 | 7.39 | 12.45 | 99 | 18.9 | 12.3
2007-05-25 | 4:31 | 5.52 | 7.39 | 12.49 | 99.1 | 18.2 | 11.8
2007-05-25 | 4:46 | 5.42 | 7.39 | 12.5 | 98.9 | 18.2 | 11.8
2007-05-25 | 5:01 | 5.33 | 7.39 | 12.54 | 99 | 18.2 | 11.8
2007-05-25 | 5:16 | 5.25 | 7.38 | 12.57 | 99.1 | 18.2 | 11.8
2007-05-25 | 5:31 | 5.17 | 7.38 | 12.59 | 99 | 18.2 | 11.8
2007-05-25 | 5:46 | 5.09 | 7.38 | 12.62 | 99 | 18.2 | 11.8
2007-05-25 | 6:01 | 5.02 | 7.39 | 12.64 | 99 | 18.2 | 11.8
2007-05-25 | 6:16 | 4.95 | 7.39 | 12.67 | 99.1 | 18.2 | 11.8
2007-05-25 | 6:31 | 4.89 | 7.39 | 12.7 | 99.1 | 18.2 | 11.8
2007-05-25 | 6:46 | 4.84 | 7.39 | 12.71 | 99.1 | 18.2 | 11.8
2007-05-25 | 7:01 | 4.8 | 7.39 | 12.72 | 99.1 | 18.2 | 11.8
2007-05-25 | 7:16 | 4.77 | 7.4 | 12.73 | 99.1 | 18.2 | 11.8
2007-05-25 | 7:31 | 4.75 | 7.41 | 12.74 | 99.1 | 18.2 | 11.8
2007-05-25 | 7:46 | 4.75 | 7.42 | 12.74 | 99.1 | 18.2 | 11.8
2007-05-25 | 8:01 | 4.74 | 7.43 | 12.74 | 99.1 | 18.3 | 11.9
2007-05-25 | 8:16 | 4.78 | 7.44 | 12.74 | 99.1 | 18.2 | 11.8
2007-05-25 | 8:31 | 4.79 | 7.44 | 12.74 | 99.2 | 18.2 | 11.9
2007-05-25 | 8:46 | 4.84 | 7.46 | 12.74 | 99.3 | 18.2 | 11.8
2007-05-25 | 9:01 | 4.9 | 7.45 | 12.71 | 99.3 | 18.2 | 11.8
2007-05-25 | 9:16 | 5 | 7.45 | 12.68 | 99.3 | 18.3 | 11.9
2007-05-25 | 9:31 | 5.1 | 7.46 | 12.65 | 99.3 | 18.2 | 11.9
2007-05-25 | 9:46 | 5.22 | 7.46 | 12.61 | 99.3 | 18.3 | 11.9
2007-05-25 | 10:01 | 5.36 | 7.46 | 12.54 | 99.1 | 18.3 | 11.9
2007-05-25 | 10:16 | 5.49 | 7.46 | 12.53 | 99.3 | 18.2 | 11.9
2007-05-25 | 10:31 | 5.66 | 7.47 | 12.46 | 99.2 | 18.3 | 11.9
2007-05-25 | 10:46 | 5.85 | 7.48 | 12.44 | 99.5 | 18.3 | 11.9
2007-05-25 | 11:01 | 6.07 | 7.48 | 12.36 | 99.5 | 18.3 | 11.9
2007-05-25 | 11:16 | 6.2 | 7.48 | 12.31 | 99.4 | 18.3 | 11.9
2007-05-25 | 11:31 | 6.43 | 7.48 | 12.25 | 99.5 | 18.3 | 11.9
2007-05-25 | 11:46 | 6.66 | 7.49 | 12.19 | 99.6 | 18.4 | 11.9
2007-05-25 | 12:01 | 6.92 | 7.48 | 12.11 | 99.5 | 18.3 | 11.9
2007-05-25 | 12:16 | 7.14 | 7.49 | 12.03 | 99.5 | 18.3 | 11.9
2007-05-25 | 12:31 | 7.26 | 7.5 | 11.99 | 99.4 | 18.3 | 11.9
2007-05-25 | 12:46 | 7.56 | 7.5 | 11.91 | 99.5 | 18.3 | 11.9
2007-05-25 | 13:00 | 7.82 | 7.5 | 11.85 | 99.6 | 18.3 | 11.9
2007-05-25 | 13:15 | 8.02 | 7.52 | 11.78 | 99.5 | 18.3 | 11.9
2007-05-25 | 13:30 | 8.27 | 7.52 | 11.72 | 99.6 | 18.3 | 11.9
2007-05-25 | 13:45 | 8.48 | 7.52 | 11.64 | 99.4 | 18.3 | 11.9
2007-05-25 | 14:00 | 8.64 | 7.51 | 11.58 | 99.3 | 18.4 | 11.9
2007-05-25 | 14:15 | 8.76 | 7.52 | 11.55 | 99.3 | 18.3 | 11.9
2007-05-25 | 14:30 | 8.86 | 7.52 | 11.51 | 99.2 | 18.4 | 11.9
2007-05-25 | 14:45 | 8.93 | 7.52 | 11.49 | 99.2 | 18.4 | 11.9
2007-05-25 | 15:00 | 9 | 7.51 | 11.47 | 99.3 | 18.3 | 11.9
2007-05-25 | 15:15 | 9.07 | 7.51 | 11.45 | 99.2 | 18.2 | 11.8
2007-05-25 | 15:30 | 9.15 | 7.52 | 11.43 | 99.2 | 18.2 | 11.9
2007-05-25 | 15:45 | 9.31 | 7.51 | 11.41 | 99.5 | 18.2 | 11.8
2007-05-25 | 16:00 | 9.49 | 7.52 | 11.35 | 99.3 | 18.2 | 11.8
2007-05-25 | 16:15 | 9.58 | 7.52 | 11.33 | 99.3 | 18.2 | 11.8
2007-05-25 | 16:30 | 9.62 | 7.53 | 11.31 | 99.3 | 18.1 | 11.8
2007-05-25 | 16:45 | 9.67 | 7.52 | 11.32 | 99.5 | 18.1 | 11.8
2007-05-25 | 17:00 | 9.68 | 7.52 | 11.3 | 99.3 | 18.1 | 11.8
2007-05-25 | 17:15 | 9.63 | 7.52 | 11.32 | 99.4 | 18.1 | 11.7
2007-05-25 | 17:30 | 9.61 | 7.52 | 11.34 | 99.5 | 18.1 | 11.7
2007-05-25 | 17:45 | 9.63 | 7.52 | 11.33 | 99.5 | 18.1 | 11.8
2007-05-25 | 18:00 | 9.63 | 7.52 | 11.3 | 99.2 | 18 | 11.7
2007-05-25 | 18:15 | 9.67 | 7.51 | 11.28 | 99.2 | 18.1 | 11.8
2007-05-25 | 18:30 | 9.75 | 7.51 | 11.27 | 99.2 | 18 | 11.7
2007-05-25 | 18:45 | 9.78 | 7.51 | 11.24 | 99.1 | 18.2 | 11.8
2007-05-25 | 19:00 | 9.77 | 7.51 | 11.26 | 99.2 | 18.2 | 11.8
2007-05-25 | 19:15 | 9.67 | 7.51 | 11.28 | 99.2 | 18.1 | 11.8
2007-05-25 | 19:30 | 9.58 | 7.5 | 11.31 | 99.2 | 18.1 | 11.8
2007-05-25 | 19:45 | 9.47 | 7.5 | 11.35 | 99.3 | 18.1 | 11.7
2007-05-25 | 20:00 | 9.31 | 7.48 | 11.4 | 99.3 | 18.1 | 11.7
2007-05-25 | 20:15 | 9.17 | 7.48 | 11.42 | 99.2 | 18 | 11.7
2007-05-25 | 20:30 | 9.06 | 7.47 | 11.44 | 99.1 | 18 | 11.7
2007-05-25 | 20:45 | 8.96 | 7.47 | 11.44 | 98.9 | 18 | 11.7
2007-05-25 | 21:00 | 8.86 | 7.46 | 11.49 | 99.1 | 17.9 | 11.7
2007-05-25 | 21:15 | 8.74 | 7.46 | 11.55 | 99.3 | 17.9 | 11.6
2007-05-25 | 21:30 | 8.63 | 7.45 | 11.58 | 99.3 | 17.9 | 11.6
2007-05-25 | 21:45 | 8.52 | 7.45 | 11.6 | 99.2 | 17.9 | 11.6
2007-05-25 | 22:00 | 8.4 | 7.45 | 11.64 | 99.2 | 17.9 | 11.6
2007-05-25 | 22:15 | 8.28 | 7.44 | 11.68 | 99.3 | 17.8 | 11.6
2007-05-25 | 22:30 | 8.16 | 7.43 | 11.7 | 99.2 | 17.8 | 11.6
2007-05-25 | 22:45 | 8.05 | 7.42 | 11.75 | 99.3 | 17.8 | 11.6
2007-05-25 | 23:00 | 7.92 | 7.42 | 11.78 | 99.2 | 17.8 | 11.6
2007-05-25 | 23:15 | 7.8 | 7.42 | 11.81 | 99.2 | 17.8 | 11.5
2007-05-25 | 23:30 | 7.67 | 7.41 | 11.86 | 99.3 | 17.8 | 11.6
2007-05-25 | 23:45 | 7.54 | 7.42 | 11.9 | 99.3 | 17.7 | 11.5
2007-05-26 | 0:00 | 7.42 | 7.41 | 11.92 | 99.2 | 17.7 | 11.5
2007-05-26 | 0:15 | 7.3 | 7.41 | 11.96 | 99.3 | 17.8 | 11.5
2007-05-26 | 0:30 | 7.17 | 7.41 | 12 | 99.2 | 17.7 | 11.5
2007-05-26 | 0:45 | 7.05 | 7.41 | 12.01 | 99.1 | 17.6 | 11.5
2007-05-26 | 1:01 | 6.92 | 7.41 | 12.07 | 99.2 | 17.7 | 11.5
2007-05-26 | 1:16 | 6.8 | 7.41 | 12.1 | 99.2 | 17.7 | 11.5
2007-05-26 | 1:31 | 6.68 | 7.41 | 12.14 | 99.2 | 17.6 | 11.5
2007-05-26 | 1:46 | 6.56 | 7.41 | 12.17 | 99.2 | 17.7 | 11.5
2007-05-26 | 2:01 | 6.43 | 7.41 | 12.21 | 99.1 | 17.7 | 11.5
2007-05-26 | 2:16 | 6.32 | 7.41 | 12.24 | 99.1 | 17.7 | 11.5
2007-05-26 | 2:31 | 6.2 | 7.41 | 12.28 | 99.1 | 17.6 | 11.5
2007-05-26 | 2:46 | 6.09 | 7.41 | 12.31 | 99.1 | 17.6 | 11.4
2007-05-26 | 3:01 | 5.98 | 7.4 | 12.35 | 99.1 | 17.6 | 11.4
2007-05-26 | 3:16 | 5.88 | 7.4 | 12.37 | 99.1 | 17.6 | 11.4
2007-05-26 | 3:31 | 5.79 | 7.4 | 12.4 | 99.1 | 17.6 | 11.5
2007-05-26 | 3:46 | 5.69 | 7.4 | 12.44 | 99.1 | 17.6 | 11.4
2007-05-26 | 4:01 | 5.61 | 7.4 | 12.47 | 99.2 | 17.6 | 11.5
2007-05-26 | 4:16 | 5.52 | 7.4 | 12.48 | 99.1 | 17.6 | 11.5
2007-05-26 | 4:31 | 5.44 | 7.4 | 12.52 | 99.2 | 17.6 | 11.5
2007-05-26 | 4:46 | 5.37 | 7.4 | 12.54 | 99.1 | 17.6 | 11.5
2007-05-26 | 5:01 | 5.31 | 7.4 | 12.56 | 99.2 | 17.6 | 11.5
2007-05-26 | 5:16 | 5.25 | 7.4 | 12.58 | 99.2 | 17.7 | 11.5
2007-05-26 | 5:31 | 5.19 | 7.4 | 12.59 | 99.1 | 17.6 | 11.5
2007-05-26 | 5:46 | 5.13 | 7.39 | 12.61 | 99.1 | 17.6 | 11.5
2007-05-26 | 6:01 | 5.08 | 7.39 | 12.63 | 99.1 | 17.6 | 11.5
2007-05-26 | 6:16 | 5.04 | 7.39 | 12.65 | 99.1 | 17.6 | 11.5
2007-05-26 | 6:31 | 5 | 7.4 | 12.66 | 99.1 | 17.7 | 11.5
2007-05-26 | 6:46 | 4.97 | 7.4 | 12.67 | 99.1 | 17.7 | 11.5
2007-05-26 | 7:01 | 4.95 | 7.41 | 12.68 | 99.1 | 17.7 | 11.5
2007-05-26 | 7:16 | 4.93 | 7.41 | 12.68 | 99.1 | 17.7 | 11.5
2007-05-26 | 7:31 | 4.93 | 7.42 | 12.68 | 99.1 | 17.7 | 11.5
2007-05-26 | 7:46 | 4.93 | 7.42 | 12.69 | 99.1 | 17.7 | 11.5
2007-05-26 | 8:01 | 4.94 | 7.42 | 12.69 | 99.2 | 17.6 | 11.5
2007-05-26 | 8:16 | 4.95 | 7.43 | 12.68 | 99.1 | 17.7 | 11.5
2007-05-26 | 8:31 | 4.98 | 7.45 | 12.67 | 99.2 | 17.7 | 11.5
2007-05-26 | 8:46 | 5.03 | 7.46 | 12.66 | 99.2 | 17.6 | 11.4
2007-05-26 | 9:01 | 5.09 | 7.46 | 12.65 | 99.2 | 17.6 | 11.4
2007-05-26 | 9:16 | 5.16 | 7.46 | 12.6 | 99.1 | 17.6 | 11.5
2007-05-26 | 9:31 | 5.26 | 7.47 | 12.59 | 99.2 | 17.6 | 11.5
2007-05-26 | 9:46 | 5.37 | 7.47 | 12.56 | 99.3 | 17.6 | 11.4
2007-05-26 | 10:01 | 5.5 | 7.47 | 12.53 | 99.3 | 17.6 | 11.4
2007-05-26 | 10:16 | 5.65 | 7.48 | 12.49 | 99.5 | 17.6 | 11.4
2007-05-26 | 10:31 | 5.81 | 7.48 | 12.43 | 99.4 | 17.6 | 11.4
2007-05-26 | 10:46 | 5.98 | 7.48 | 12.37 | 99.3 | 17.5 | 11.4
2007-05-26 | 11:01 | 6.13 | 7.49 | 12.34 | 99.4 | 17.6 | 11.4
2007-05-26 | 11:16 | 6.32 | 7.49 | 12.27 | 99.4 | 17.6 | 11.4
2007-05-26 | 11:31 | 6.48 | 7.49 | 12.22 | 99.4 | 17.6 | 11.4
2007-05-26 | 11:46 | 6.74 | 7.49 | 12.15 | 99.5 | 17.6 | 11.5
2007-05-26 | 12:01 | 6.97 | 7.5 | 12.08 | 99.4 | 17.6 | 11.5
2007-05-26 | 12:16 | 7.22 | 7.51 | 12.01 | 99.5 | 17.6 | 11.4
2007-05-26 | 12:31 | 7.47 | 7.51 | 11.95 | 99.6 | 17.6 | 11.4
2007-05-26 | 12:46 | 7.77 | 7.51 | 11.86 | 99.6 | 17.6 | 11.4
2007-05-26 | 13:00 | 8.1 | 7.51 | 11.76 | 99.5 | 17.6 | 11.4
2007-05-26 | 13:15 | 8.4 | 7.52 | 11.68 | 99.6 | 17.6 | 11.4
2007-05-26 | 13:30 | 8.69 | 7.52 | 11.61 | 99.6 | 17.6 | 11.5
2007-05-26 | 13:45 | 8.98 | 7.52 | 11.51 | 99.5 | 17.6 | 11.5
2007-05-26 | 14:00 | 9.28 | 7.53 | 11.43 | 99.5 | 17.7 | 11.5
2007-05-26 | 14:15 | 9.55 | 7.53 | 11.36 | 99.6 | 17.7 | 11.5
2007-05-26 | 14:30 | 9.91 | 7.53 | 11.27 | 99.6 | 17.6 | 11.4
2007-05-26 | 14:45 | 10.08 | 7.53 | 11.18 | 99.2 | 17.6 | 11.5
2007-05-26 | 15:00 | 10.18 | 7.53 | 11.17 | 99.4 | 17.6 | 11.5
2007-05-26 | 15:15 | 10.25 | 7.54 | 11.15 | 99.4 | 17.6 | 11.5
2007-05-26 | 15:30 | 10.43 | 7.54 | 11.08 | 99.1 | 17.6 | 11.5
2007-05-26 | 15:45 | 10.41 | 7.53 | 11.09 | 99.2 | 17.6 | 11.5
2007-05-26 | 16:00 | 10.76 | 7.54 | 11.01 | 99.3 | 17.6 | 11.4
2007-05-26 | 16:15 | 11.23 | 7.54 | 10.89 | 99.2 | 17.6 | 11.4
2007-05-26 | 16:30 | 11.57 | 7.55 | 10.79 | 99.2 | 17.6 | 11.4
2007-05-26 | 16:45 | 11.82 | 7.55 | 10.71 | 99 | 17.6 | 11.4
2007-05-26 | 17:00 | 11.93 | 7.55 | 10.7 | 99.2 | 17.6 | 11.4
2007-05-26 | 17:15 | 11.99 | 7.55 | 10.7 | 99.3 | 17.6 | 11.4
2007-05-26 | 17:30 | 12.02 | 7.55 | 10.68 | 99.2 | 17.7 | 11.5
2007-05-26 | 17:45 | 12.06 | 7.54 | 10.66 | 99.1 | 17.7 | 11.5
2007-05-26 | 18:00 | 12.05 | 7.54 | 10.67 | 99.1 | 17.7 | 11.5
2007-05-26 | 18:15 | 12 | 7.54 | 10.67 | 99 | 17.7 | 11.5
2007-05-26 | 18:30 | 11.9 | 7.54 | 10.71 | 99.2 | 17.7 | 11.5
2007-05-26 | 18:45 | 11.75 | 7.54 | 10.75 | 99.2 | 17.6 | 11.5
2007-05-26 | 19:00 | 11.56 | 7.53 | 10.77 | 99 | 17.6 | 11.4
2007-05-26 | 19:15 | 11.41 | 7.53 | 10.82 | 99 | 17.5 | 11.4
2007-05-26 | 19:30 | 11.29 | 7.53 | 10.8 | 98.6 | 17.5 | 11.4
2007-05-26 | 19:45 | 11.19 | 7.53 | 10.86 | 98.9 | 17.5 | 11.3
2007-05-26 | 20:00 | 11.1 | 7.52 | 10.89 | 99 | 17.4 | 11.3
2007-05-26 | 20:15 | 11.02 | 7.51 | 10.93 | 99.1 | 17.4 | 11.3
2007-05-26 | 20:30 | 10.96 | 7.5 | 10.92 | 99 | 17.3 | 11.3
2007-05-26 | 20:45 | 10.88 | 7.49 | 10.94 | 99 | 17.3 | 11.2
2007-05-26 | 21:00 | 10.77 | 7.49 | 10.96 | 98.9 | 17.2 | 11.2
2007-05-26 | 21:15 | 10.64 | 7.47 | 11 | 98.9 | 17.2 | 11.2
2007-05-26 | 21:30 | 10.5 | 7.46 | 11.06 | 99.2 | 17.1 | 11.1
2007-05-26 | 21:45 | 10.36 | 7.45 | 11.09 | 99 | 17.1 | 11.1
2007-05-26 | 22:00 | 10.2 | 7.44 | 11.14 | 99.2 | 17.1 | 11.1
2007-05-26 | 22:15 | 10.04 | 7.43 | 11.18 | 99.2 | 17 | 11.1
2007-05-26 | 22:30 | 9.88 | 7.43 | 11.22 | 99.1 | 17 | 11.1
2007-05-26 | 22:45 | 9.73 | 7.43 | 11.27 | 99.1 | 17 | 11
2007-05-26 | 23:00 | 9.58 | 7.42 | 11.31 | 99.2 | 16.9 | 11
2007-05-26 | 23:15 | 9.44 | 7.42 | 11.35 | 99.2 | 16.8 | 10.9
2007-05-26 | 23:30 | 9.3 | 7.41 | 11.38 | 99.1 | 16.7 | 10.9
2007-05-26 | 23:45 | 9.17 | 7.4 | 11.42 | 99.2 | 16.6 | 10.8
2007-05-27 | 0:00 | 9.02 | 7.4 | 11.45 | 99.1 | 16.6 | 10.8
2007-05-27 | 0:15 | 8.87 | 7.4 | 11.5 | 99.2 | 16.5 | 10.7
2007-05-27 | 0:30 | 8.71 | 7.4 | 11.54 | 99.2 | 16.4 | 10.7
2007-05-27 | 0:45 | 8.55 | 7.39 | 11.6 | 99.3 | 16.4 | 10.6
2007-05-27 | 1:01 | 8.4 | 7.39 | 11.62 | 99.1 | 16.3 | 10.6
2007-05-27 | 1:16 | 8.25 | 7.39 | 11.68 | 99.2 | 16.5 | 10.7
2007-05-27 | 1:31 | 8.11 | 7.39 | 11.72 | 99.2 | 16.4 | 10.7
2007-05-27 | 1:46 | 7.99 | 7.38 | 11.75 | 99.1 | 16.4 | 10.6
2007-05-27 | 2:01 | 7.87 | 7.38 | 11.79 | 99.2 | 16.6 | 10.8
2007-05-27 | 2:16 | 7.75 | 7.38 | 11.8 | 99 | 16.6 | 10.8
2007-05-27 | 2:31 | 7.64 | 7.38 | 11.84 | 99.1 | 16.7 | 10.8
2007-05-27 | 2:46 | 7.52 | 7.37 | 11.89 | 99.2 | 17 | 11.1
2007-05-27 | 3:01 | 7.42 | 7.37 | 11.93 | 99.3 | 16.9 | 11
2007-05-27 | 3:16 | 7.32 | 7.38 | 11.95 | 99.2 | 17 | 11.1
2007-05-27 | 3:31 | 7.22 | 7.4 | 11.99 | 99.3 | 17.1 | 11.1
2007-05-27 | 3:46 | 7.12 | 7.41 | 12.01 | 99.2 | 17.2 | 11.2
2007-05-27 | 4:01 | 7.03 | 7.4 | 12.03 | 99.1 | 17.2 | 11.2
2007-05-27 | 4:16 | 6.95 | 7.43 | 12.04 | 99.1 | 17.2 | 11.2
2007-05-27 | 4:31 | 6.88 | 7.43 | 12.06 | 99.1 | 17.4 | 11.3
2007-05-27 | 4:46 | 6.82 | 7.42 | 12.09 | 99.1 | 17.4 | 11.3
2007-05-27 | 5:01 | 6.77 | 7.43 | 12.11 | 99.1 | 17.4 | 11.3
2007-05-27 | 5:16 | 6.71 | 7.42 | 12.13 | 99.2 | 17.4 | 11.3
2007-05-27 | 5:31 | 6.65 | 7.43 | 12.14 | 99.1 | 17.3 | 11.2
2007-05-27 | 5:46 | 6.6 | 7.43 | 12.15 | 99.1 | 17.4 | 11.3
2007-05-27 | 6:01 | 6.55 | 7.43 | 12.19 | 99.2 | 17.4 | 11.3
2007-05-27 | 6:16 | 6.5 | 7.43 | 12.2 | 99.3 | 17.4 | 11.3
2007-05-27 | 6:31 | 6.46 | 7.43 | 12.2 | 99.2 | 17.4 | 11.3
2007-05-27 | 6:46 | 6.42 | 7.43 | 12.22 | 99.2 | 17.3 | 11.3
2007-05-27 | 7:01 | 6.39 | 7.43 | 12.23 | 99.2 | 17.3 | 11.3
2007-05-27 | 7:16 | 6.36 | 7.43 | 12.24 | 99.2 | 17.5 | 11.4
2007-05-27 | 7:31 | 6.35 | 7.43 | 12.26 | 99.4 | 17.5 | 11.3
2007-05-27 | 7:46 | 6.33 | 7.43 | 12.23 | 99.1 | 17.5 | 11.4
2007-05-27 | 8:01 | 6.31 | 7.44 | 12.26 | 99.3 | 17.5 | 11.3
2007-05-27 | 8:16 | 6.3 | 7.44 | 12.26 | 99.2 | 17.5 | 11.4
2007-05-27 | 8:31 | 6.28 | 7.43 | 12.27 | 99.3 | 17.5 | 11.4
2007-05-27 | 8:46 | 6.28 | 7.44 | 12.28 | 99.3 | 17.6 | 11.4
2007-05-27 | 9:01 | 6.29 | 7.44 | 12.27 | 99.3 | 17.6 | 11.5
2007-05-27 | 9:16 | 6.32 | 7.44 | 12.26 | 99.3 | 17.6 | 11.4
2007-05-27 | 9:31 | 6.36 | 7.44 | 12.27 | 99.5 | 17.7 | 11.5
2007-05-27 | 9:46 | 6.42 | 7.45 | 12.24 | 99.4 | 17.7 | 11.5
2007-05-27 | 10:01 | 6.5 | 7.45 | 12.22 | 99.4 | 17.7 | 11.5
2007-05-27 | 10:16 | 6.58 | 7.45 | 12.2 | 99.5 | 17.6 | 11.5
2007-05-27 | 10:31 | 6.68 | 7.45 | 12.16 | 99.4 | 17.7 | 11.5
2007-05-27 | 10:46 | 6.78 | 7.45 | 12.14 | 99.5 | 17.7 | 11.5
2007-05-27 | 11:01 | 6.9 | 7.45 | 12.1 | 99.4 | 17.6 | 11.5
2007-05-27 | 11:16 | 7.03 | 7.46 | 12.07 | 99.5 | 17.9 | 11.7
2007-05-27 | 11:31 | 7.17 | 7.46 | 12.03 | 99.5 | 17.9 | 11.6
2007-05-27 | 11:46 | 7.33 | 7.46 | 12 | 99.7 | 17.9 | 11.6
2007-05-27 | 12:01 | 7.5 | 7.47 | 11.96 | 99.7 | 17.9 | 11.6
2007-05-27 | 12:16 | 7.7 | 7.47 | 11.89 | 99.7 | 18 | 11.7
2007-05-27 | 12:31 | 7.9 | 7.47 | 11.83 | 99.7 | 18 | 11.7
2007-05-27 | 12:46 | 8.11 | 7.48 | 11.78 | 99.7 | 18.1 | 11.7
2007-05-27 | 13:00 | 8.33 | 7.47 | 11.72 | 99.8 | 18 | 11.7
2007-05-27 | 13:15 | 8.56 | 7.48 | 11.66 | 99.8 | 18 | 11.7
2007-05-27 | 13:30 | 8.79 | 7.48 | 11.61 | 99.9 | 18 | 11.7
2007-05-27 | 13:45 | 9.01 | 7.49 | 11.55 | 99.9 | 17.9 | 11.6
2007-05-27 | 14:00 | 9.22 | 7.49 | 11.48 | 99.9 | 18 | 11.7
2007-05-27 | 14:15 | 9.4 | 7.49 | 11.44 | 99.9 | 18.2 | 11.8
2007-05-27 | 14:30 | 9.59 | 7.49 | 11.38 | 99.8 | 18.1 | 11.7
2007-05-27 | 14:45 | 9.75 | 7.5 | 11.33 | 99.8 | 18 | 11.7
2007-05-27 | 15:00 | 9.95 | 7.49 | 11.29 | 99.9 | 18.1 | 11.7
2007-05-27 | 15:15 | 10.19 | 7.49 | 11.22 | 99.8 | 18 | 11.7
2007-05-27 | 15:30 | 10.38 | 7.5 | 11.17 | 99.8 | 18 | 11.7
2007-05-27 | 15:45 | 10.53 | 7.5 | 11.13 | 99.8 | 17.9 | 11.7
2007-05-27 | 16:00 | 10.59 | 7.51 | 11.11 | 99.8 | 18 | 11.7
2007-05-27 | 16:15 | 10.61 | 7.5 | 11.11 | 99.8 | 18 | 11.7
2007-05-27 | 16:30 | 10.64 | 7.5 | 11.05 | 99.4 | 17.9 | 11.6
2007-05-27 | 16:45 | 10.62 | 7.5 | 11.09 | 99.7 | 17.9 | 11.6
2007-05-27 | 17:00 | 10.61 | 7.49 | 11.1 | 99.7 | 17.9 | 11.6
2007-05-27 | 17:15 | 10.56 | 7.49 | 11.12 | 99.8 | 17.8 | 11.6
2007-05-27 | 17:30 | 10.39 | 7.49 | 11.14 | 99.6 | 17.8 | 11.6
2007-05-27 | 17:45 | 10.19 | 7.49 | 11.2 | 99.7 | 17.7 | 11.5
2007-05-27 | 18:00 | 10.02 | 7.48 | 11.25 | 99.7 | 17.5 | 11.4
2007-05-27 | 18:15 | 9.86 | 7.48 | 11.29 | 99.7 | 17.7 | 11.5
2007-05-27 | 18:30 | 9.67 | 7.48 | 11.32 | 99.5 | 17.7 | 11.5
2007-05-27 | 18:45 | 9.48 | 7.47 | 11.38 | 99.6 | 17.5 | 11.4
2007-05-27 | 19:00 | 9.3 | 7.47 | 11.43 | 99.6 | 17.5 | 11.4
2007-05-27 | 19:15 | 9.12 | 7.46 | 11.46 | 99.4 | 17.4 | 11.3
2007-05-27 | 19:30 | 8.89 | 7.46 | 11.55 | 99.6 | 17.4 | 11.3
2007-05-27 | 19:45 | 8.66 | 7.46 | 11.61 | 99.6 | 17.3 | 11.2
2007-05-27 | 20:00 | 8.42 | 7.45 | 11.67 | 99.5 | 17.2 | 11.2
2007-05-27 | 20:15 | 8.19 | 7.44 | 11.74 | 99.6 | 17.3 | 11.2
2007-05-27 | 20:30 | 7.97 | 7.44 | 11.81 | 99.6 | 17.2 | 11.2
2007-05-27 | 20:45 | 7.75 | 7.43 | 11.86 | 99.6 | 17.2 | 11.2
2007-05-27 | 21:00 | 7.53 | 7.44 | 11.93 | 99.6 | 17.2 | 11.1
2007-05-27 | 21:15 | 7.31 | 7.43 | 12 | 99.6 | 17.3 | 11.2
2007-05-27 | 21:30 | 7.08 | 7.42 | 12.06 | 99.6 | 17.3 | 11.2
2007-05-27 | 21:45 | 6.86 | 7.42 | 12.13 | 99.6 | 17.4 | 11.3
2007-05-27 | 22:00 | 6.64 | 7.43 | 12.21 | 99.7 | 17.5 | 11.4
2007-05-27 | 22:15 | 6.43 | 7.41 | 12.26 | 99.6 | 17.4 | 11.3
2007-05-27 | 22:30 | 6.23 | 7.42 | 12.32 | 99.6 | 17.4 | 11.3
2007-05-27 | 22:45 | 6.04 | 7.41 | 12.38 | 99.6 | 17.4 | 11.3
2007-05-27 | 23:00 | 5.86 | 7.41 | 12.44 | 99.6 | 17.4 | 11.3
2007-05-27 | 23:15 | 5.68 | 7.4 | 12.5 | 99.6 | 17.4 | 11.3
2007-05-27 | 23:30 | 5.51 | 7.4 | 12.54 | 99.5 | 17.3 | 11.3
2007-05-27 | 23:45 | 5.34 | 7.4 | 12.58 | 99.3 | 17.4 | 11.3
2007-05-28 | 0:00 | 5.18 | 7.39 | 12.63 | 99.3 | 17.4 | 11.3
2007-05-28 | 0:15 | 5.04 | 7.39 | 12.7 | 99.5 | 17.2 | 11.2
2007-05-28 | 0:30 | 4.9 | 7.39 | 12.74 | 99.5 | 17.3 | 11.2
2007-05-28 | 0:45 | 4.76 | 7.39 | 12.78 | 99.4 | 17.4 | 11.3
2007-05-28 | 1:01 | 4.63 | 7.38 | 12.82 | 99.4 | 17.3 | 11.2
2007-05-28 | 1:16 | 4.49 | 7.38 | 12.87 | 99.4 | 17.3 | 11.3
2007-05-28 | 1:31 | 4.34 | 7.37 | 12.91 | 99.4 | 17.4 | 11.3
2007-05-28 | 1:46 | 4.21 | 7.37 | 12.96 | 99.4 | 17.4 | 11.3
2007-05-28 | 2:01 | 4.11 | 7.38 | 13 | 99.5 | 17.3 | 11.2
2007-05-28 | 2:16 | 3.97 | 7.37 | 13.03 | 99.3 | 17.5 | 11.3
2007-05-28 | 2:31 | 3.81 | 7.37 | 13.09 | 99.4 | 17.4 | 11.3
2007-05-28 | 2:46 | 3.66 | 7.37 | 13.15 | 99.4 | 17.4 | 11.3
2007-05-28 | 3:01 | 3.53 | 7.37 | 13.17 | 99.2 | 17.4 | 11.3
2007-05-28 | 3:16 | 3.4 | 7.37 | 13.23 | 99.3 | 17.5 | 11.4
2007-05-28 | 3:31 | 3.27 | 7.36 | 13.28 | 99.3 | 17.4 | 11.3
2007-05-28 | 3:46 | 3.16 | 7.36 | 13.32 | 99.4 | 17.6 | 11.4
2007-05-28 | 4:01 | 3.05 | 7.36 | 13.36 | 99.4 | 17.6 | 11.4
2007-05-28 | 4:16 | 2.96 | 7.35 | 13.38 | 99.3 | 17.5 | 11.4
2007-05-28 | 4:31 | 2.88 | 7.36 | 13.41 | 99.3 | 17.6 | 11.4
2007-05-28 | 4:46 | 2.79 | 7.35 | 13.43 | 99.2 | 17.6 | 11.5
2007-05-28 | 5:01 | 2.7 | 7.35 | 13.47 | 99.3 | 17.6 | 11.5
2007-05-28 | 5:16 | 2.62 | 7.35 | 13.5 | 99.2 | 17.7 | 11.5
2007-05-28 | 5:31 | 2.53 | 7.35 | 13.53 | 99.2 | 17.8 | 11.5
2007-05-28 | 5:46 | 2.45 | 7.35 | 13.56 | 99.2 | 17.7 | 11.5
2007-05-28 | 6:01 | 2.37 | 7.34 | 13.59 | 99.2 | 17.6 | 11.4
2007-05-28 | 6:16 | 2.3 | 7.34 | 13.59 | 99.1 | 17.8 | 11.6
2007-05-28 | 6:31 | 2.24 | 7.35 | 13.64 | 99.3 | 17.8 | 11.6
2007-05-28 | 6:46 | 2.19 | 7.35 | 13.66 | 99.3 | 17.9 | 11.6
2007-05-28 | 7:01 | 2.15 | 7.35 | 13.67 | 99.3 | 17.8 | 11.6
2007-05-28 | 7:16 | 2.12 | 7.35 | 13.69 | 99.3 | 17.8 | 11.6
2007-05-28 | 7:31 | 2.1 | 7.35 | 13.69 | 99.3 | 17.9 | 11.6
2007-05-28 | 7:46 | 2.09 | 7.36 | 13.71 | 99.3 | 17.9 | 11.6
2007-05-28 | 8:01 | 2.1 | 7.38 | 13.7 | 99.3 | 17.9 | 11.6
2007-05-28 | 8:16 | 2.11 | 7.38 | 13.7 | 99.3 | 17.9 | 11.6
2007-05-28 | 8:31 | 2.14 | 7.39 | 13.7 | 99.4 | 17.9 | 11.7
2007-05-28 | 8:46 | 2.18 | 7.39 | 13.7 | 99.5 | 17.9 | 11.7
2007-05-28 | 9:01 | 2.23 | 7.39 | 13.66 | 99.4 | 17.9 | 11.6
2007-05-28 | 9:16 | 2.3 | 7.4 | 13.64 | 99.4 | 17.7 | 11.5
2007-05-28 | 9:31 | 2.39 | 7.39 | 13.61 | 99.5 | 17.8 | 11.5
2007-05-28 | 9:46 | 2.49 | 7.4 | 13.58 | 99.5 | 17.8 | 11.6
2007-05-28 | 10:01 | 2.61 | 7.4 | 13.53 | 99.5 | 17.9 | 11.6
2007-05-28 | 10:16 | 2.75 | 7.41 | 13.47 | 99.4 | 17.8 | 11.6
2007-05-28 | 10:31 | 2.91 | 7.41 | 13.44 | 99.6 | 17.8 | 11.6
2007-05-28 | 10:46 | 3.07 | 7.42 | 13.38 | 99.6 | 17.8 | 11.6
2007-05-28 | 11:01 | 3.27 | 7.42 | 13.31 | 99.6 | 17.8 | 11.6
2007-05-28 | 11:16 | 3.49 | 7.43 | 13.25 | 99.7 | 17.8 | 11.6
2007-05-28 | 11:31 | 3.72 | 7.43 | 13.17 | 99.7 | 17.8 | 11.6
2007-05-28 | 11:46 | 3.96 | 7.43 | 13.09 | 99.8 | 17.8 | 11.6
2007-05-28 | 12:01 | 4.23 | 7.43 | 13.01 | 99.8 | 17.8 | 11.6
2007-05-28 | 12:16 | 4.52 | 7.44 | 12.93 | 100 | 17.8 | 11.6
2007-05-28 | 12:31 | 4.83 | 7.44 | 12.82 | 99.9 | 17.8 | 11.6
2007-05-28 | 12:46 | 5.16 | 7.45 | 12.72 | 100 | 17.9 | 11.6
2007-05-28 | 13:00 | 5.5 | 7.45 | 12.61 | 100 | 18.1 | 11.7
2007-05-28 | 13:15 | 5.85 | 7.46 | 12.5 | 100 | 18 | 11.7
2007-05-28 | 13:30 | 6.21 | 7.47 | 12.39 | 100 | 18 | 11.7
2007-05-28 | 13:45 | 6.55 | 7.47 | 12.3 | 100.2 | 18.1 | 11.8
2007-05-28 | 14:00 | 6.88 | 7.47 | 12.2 | 100.2 | 18.1 | 11.7
2007-05-28 | 14:15 | 7.22 | 7.48 | 12.09 | 100.1 | 18 | 11.7
2007-05-28 | 14:30 | 7.55 | 7.49 | 11.99 | 100.1 | 18.1 | 11.8
2007-05-28 | 14:45 | 7.86 | 7.48 | 11.89 | 100.1 | 18.1 | 11.8
2007-05-28 | 15:00 | 8.16 | 7.48 | 11.81 | 100.1 | 18.1 | 11.8
2007-05-28 | 15:15 | 8.45 | 7.49 | 11.71 | 99.9 | 18.1 | 11.8
2007-05-28 | 15:30 | 8.71 | 7.49 | 11.64 | 100 | 18.2 | 11.8
2007-05-28 | 15:45 | 8.96 | 7.49 | 11.57 | 100 | 18.3 | 11.9
2007-05-28 | 16:00 | 9.19 | 7.5 | 11.51 | 100.1 | 18.4 | 12
2007-05-28 | 16:15 | 9.4 | 7.51 | 11.46 | 100.1 | 18.4 | 12
2007-05-28 | 16:30 | 9.58 | 7.5 | 11.41 | 100 | 18.5 | 12
2007-05-28 | 16:45 | 9.73 | 7.5 | 11.37 | 100.1 | 18.5 | 12
2007-05-28 | 17:00 | 9.83 | 7.5 | 11.33 | 100 | 18.5 | 12
2007-05-28 | 17:15 | 9.9 | 7.5 | 11.33 | 100.1 | 18.4 | 12
2007-05-28 | 17:30 | 9.95 | 7.51 | 11.3 | 100 | 18.5 | 12
2007-05-28 | 17:45 | 9.95 | 7.5 | 11.29 | 99.9 | 18.6 | 12.1
2007-05-28 | 18:00 | 9.95 | 7.49 | 11.27 | 99.8 | 18.5 | 12
2007-05-28 | 18:15 | 9.92 | 7.49 | 11.28 | 99.7 | 18.5 | 12.1
2007-05-28 | 18:30 | 9.85 | 7.5 | 11.28 | 99.5 | 18.6 | 12.1
2007-05-28 | 18:45 | 9.75 | 7.5 | 11.3 | 99.5 | 18.5 | 12
2007-05-28 | 19:00 | 9.63 | 7.49 | 11.34 | 99.6 | 18.5 | 12
2007-05-28 | 19:15 | 9.49 | 7.48 | 11.39 | 99.7 | 18.5 | 12.1
2007-05-28 | 19:30 | 9.32 | 7.48 | 11.42 | 99.6 | 18.9 | 12.3
2007-05-28 | 19:45 | 9.15 | 7.47 | 11.46 | 99.5 | 18.9 | 12.3
2007-05-28 | 20:00 | 8.99 | 7.47 | 11.53 | 99.7 | 18.9 | 12.3
2007-05-28 | 20:15 | 8.84 | 7.46 | 11.55 | 99.5 | 18.8 | 12.2
2007-05-28 | 20:30 | 8.68 | 7.45 | 11.61 | 99.6 | 18.8 | 12.2
2007-05-28 | 20:45 | 8.51 | 7.45 | 11.65 | 99.6 | 18.7 | 12.2
2007-05-28 | 21:00 | 8.34 | 7.44 | 11.69 | 99.5 | 18.7 | 12.1
2007-05-28 | 21:15 | 8.16 | 7.44 | 11.76 | 99.7 | 18.6 | 12.1
2007-05-28 | 21:30 | 7.99 | 7.43 | 11.79 | 99.5 | 18.5 | 12
2007-05-28 | 21:45 | 7.81 | 7.43 | 11.84 | 99.5 | 18.8 | 12.2
2007-05-28 | 22:00 | 7.63 | 7.42 | 11.9 | 99.5 | 18.6 | 12.1
2007-05-28 | 22:15 | 7.44 | 7.42 | 11.96 | 99.6 | 18.6 | 12.1
2007-05-28 | 22:30 | 7.25 | 7.41 | 12.01 | 99.5 | 18.6 | 12.1
2007-05-28 | 22:45 | 7.07 | 7.41 | 12.06 | 99.5 | 18.6 | 12.1
2007-05-28 | 23:00 | 6.89 | 7.41 | 12.13 | 99.6 | 18.6 | 12.1
2007-05-28 | 23:15 | 6.71 | 7.41 | 12.16 | 99.5 | 18.7 | 12.1
2007-05-28 | 23:30 | 6.55 | 7.4 | 12.22 | 99.6 | 18.6 | 12.1
2007-05-28 | 23:45 | 6.39 | 7.4 | 12.27 | 99.5 | 18.5 | 12
2007-05-29 | 0:00 | 6.23 | 7.4 | 12.33 | 99.7 | 18.5 | 12
2007-05-29 | 0:15 | 6.08 | 7.4 | 12.37 | 99.6 | 18.4 | 12
2007-05-29 | 0:30 | 5.93 | 7.4 | 12.41 | 99.5 | 18.4 | 11.9
2007-05-29 | 0:45 | 5.79 | 7.39 | 12.48 | 99.7 | 18.6 | 12.1
2007-05-29 | 1:01 | 5.65 | 7.39 | 12.49 | 99.4 | 18.4 | 12
2007-05-29 | 1:16 | 5.51 | 7.4 | 12.54 | 99.5 | 18.5 | 12
2007-05-29 | 1:31 | 5.37 | 7.39 | 12.59 | 99.5 | 18.4 | 12
2007-05-29 | 1:46 | 5.24 | 7.39 | 12.64 | 99.6 | 18.4 | 12
2007-05-29 | 2:01 | 5.11 | 7.39 | 12.68 | 99.5 | 18.3 | 11.9
2007-05-29 | 2:16 | 4.98 | 7.4 | 12.72 | 99.6 | 18.3 | 11.9
2007-05-29 | 2:31 | 4.85 | 7.39 | 12.76 | 99.6 | 18.3 | 11.9
2007-05-29 | 2:46 | 4.73 | 7.38 | 12.77 | 99.3 | 18.3 | 11.9
2007-05-29 | 3:01 | 4.61 | 7.39 | 12.83 | 99.4 | 18.3 | 11.9
2007-05-29 | 3:16 | 4.49 | 7.39 | 12.87 | 99.5 | 18.4 | 11.9
2007-05-29 | 3:31 | 4.38 | 7.38 | 12.91 | 99.5 | 18.3 | 11.9
2007-05-29 | 3:46 | 4.28 | 7.38 | 12.94 | 99.4 | 18.4 | 11.9
2007-05-29 | 4:01 | 4.18 | 7.39 | 12.97 | 99.4 | 18.4 | 11.9
2007-05-29 | 4:16 | 4.08 | 7.38 | 13.02 | 99.5 | 18.4 | 11.9
2007-05-29 | 4:31 | 3.99 | 7.38 | 13.05 | 99.5 | 18.3 | 11.9
2007-05-29 | 4:46 | 3.9 | 7.38 | 13.07 | 99.4 | 18.3 | 11.9
2007-05-29 | 5:01 | 3.82 | 7.38 | 13.09 | 99.4 | 18.3 | 11.9
2007-05-29 | 5:16 | 3.74 | 7.38 | 13.12 | 99.4 | 18.3 | 11.9
2007-05-29 | 5:31 | 3.66 | 7.39 | 13.16 | 99.5 | 18.3 | 11.9
2007-05-29 | 5:46 | 3.59 | 7.37 | 13.19 | 99.6 | 18.2 | 11.8
2007-05-29 | 6:01 | 3.52 | 7.38 | 13.22 | 99.6 | 18.2 | 11.8
2007-05-29 | 6:16 | 3.47 | 7.38 | 13.23 | 99.5 | 18.2 | 11.8
2007-05-29 | 6:31 | 3.42 | 7.38 | 13.25 | 99.5 | 18.6 | 12.1
2007-05-29 | 6:46 | 3.38 | 7.38 | 13.27 | 99.6 | 18.5 | 12
2007-05-29 | 7:01 | 3.35 | 7.38 | 13.27 | 99.5 | 18.5 | 12
2007-05-29 | 7:16 | 3.33 | 7.39 | 13.28 | 99.6 | 18.5 | 12
2007-05-29 | 7:31 | 3.33 | 7.39 | 13.29 | 99.6 | 18.5 | 12
2007-05-29 | 7:46 | 3.33 | 7.4 | 13.29 | 99.6 | 18.4 | 12
2007-05-29 | 8:01 | 3.35 | 7.41 | 13.29 | 99.6 | 18.5 | 12
2007-05-29 | 8:16 | 3.38 | 7.41 | 13.29 | 99.7 | 18.6 | 12.1
2007-05-29 | 8:31 | 3.42 | 7.42 | 13.28 | 99.8 | 18.6 | 12.1
2007-05-29 | 8:46 | 3.47 | 7.42 | 13.26 | 99.8 | 18.6 | 12.1
2007-05-29 | 9:01 | 3.53 | 7.43 | 13.25 | 99.8 | 18.6 | 12.1
2007-05-29 | 9:16 | 3.63 | 7.43 | 13.21 | 99.7 | 18.6 | 12.1
2007-05-29 | 9:31 | 3.73 | 7.43 | 13.18 | 99.8 | 18.5 | 12
2007-05-29 | 9:46 | 3.85 | 7.44 | 13.16 | 100 | 18.5 | 12
2007-05-29 | 10:01 | 3.99 | 7.44 | 13.09 | 99.8 | 18.5 | 12
2007-05-29 | 10:16 | 4.14 | 7.45 | 13.05 | 100 | 18.5 | 12
2007-05-29 | 10:31 | 4.32 | 7.46 | 12.99 | 99.9 | 18.6 | 12.1
2007-05-29 | 10:46 | 4.51 | 7.46 | 12.94 | 100.1 | 18.5 | 12
2007-05-29 | 11:01 | 4.73 | 7.46 | 12.87 | 100 | 18.6 | 12.1
2007-05-29 | 11:16 | 4.96 | 7.47 | 12.79 | 100 | 18.5 | 12
2007-05-29 | 11:31 | 5.21 | 7.47 | 12.71 | 100.1 | 18.5 | 12
2007-05-29 | 11:46 | 5.47 | 7.47 | 12.63 | 100.1 | 18.5 | 12
2007-05-29 | 12:01 | 5.74 | 7.48 | 12.55 | 100.1 | 18.6 | 12.1
2007-05-29 | 12:16 | 6.04 | 7.48 | 12.46 | 100.2 | 18.8 | 12.2
2007-05-29 | 12:31 | 6.36 | 7.48 | 12.37 | 100.3 | 18.8 | 12.2
2007-05-29 | 12:46 | 6.69 | 7.48 | 12.26 | 100.2 | 18.9 | 12.3
2007-05-29 | 13:00 | 7.04 | 7.49 | 12.15 | 100.2 | 18.9 | 12.3
2007-05-29 | 13:15 | 7.4 | 7.5 | 12.05 | 100.2 | 18.9 | 12.3
2007-05-29 | 13:30 | 7.75 | 7.49 | 11.94 | 100.2 | 19 | 12.4
2007-05-29 | 13:45 | 8.12 | 7.51 | 11.85 | 100.4 | 19 | 12.4
2007-05-29 | 14:00 | 8.47 | 7.51 | 11.76 | 100.4 | 19 | 12.4
2007-05-29 | 14:15 | 8.82 | 7.51 | 11.65 | 100.3 | 19.1 | 12.4
2007-05-29 | 14:30 | 9.15 | 7.52 | 11.55 | 100.2 | 19.2 | 12.5
2007-05-29 | 14:45 | 9.47 | 7.51 | 11.46 | 100.3 | 19.1 | 12.4
2007-05-29 | 15:00 | 9.77 | 7.52 | 11.37 | 100.2 | 19.1 | 12.4
2007-05-29 | 15:15 | 10.06 | 7.53 | 11.3 | 100.3 | 19.1 | 12.4
2007-05-29 | 15:30 | 10.32 | 7.53 | 11.22 | 100.2 | 19.2 | 12.5
2007-05-29 | 15:45 | 10.57 | 7.53 | 11.15 | 100.1 | 19.2 | 12.5
2007-05-29 | 16:00 | 10.8 | 7.53 | 11.1 | 100.2 | 19.3 | 12.5
2007-05-29 | 16:15 | 11.01 | 7.53 | 11.04 | 100.1 | 19.3 | 12.5
2007-05-29 | 16:30 | 11.2 | 7.54 | 10.99 | 100.1 | 19.3 | 12.6
2007-05-29 | 16:45 | 11.35 | 7.54 | 10.96 | 100.2 | 19.6 | 12.8
2007-05-29 | 17:00 | 11.47 | 7.53 | 10.91 | 100 | 19.7 | 12.8
2007-05-29 | 17:15 | 11.55 | 7.54 | 10.89 | 100 | 19.7 | 12.8
2007-05-29 | 17:30 | 11.61 | 7.53 | 10.88 | 100 | 19.8 | 12.8
2007-05-29 | 17:45 | 11.64 | 7.53 | 10.86 | 100 | 19.8 | 12.9
2007-05-29 | 18:00 | 11.63 | 7.54 | 10.86 | 99.9 | 19.8 | 12.9
2007-05-29 | 18:15 | 11.6 | 7.53 | 10.84 | 99.7 | 19.9 | 12.9
2007-05-29 | 18:30 | 11.53 | 7.53 | 10.87 | 99.7 | 19.9 | 12.9
2007-05-29 | 18:45 | 11.43 | 7.53 | 10.88 | 99.7 | 19.9 | 12.9
2007-05-29 | 19:00 | 11.31 | 7.52 | 10.92 | 99.7 | 19.7 | 12.8
2007-05-29 | 19:15 | 11.16 | 7.52 | 10.95 | 99.6 | 19.7 | 12.8
2007-05-29 | 19:30 | 10.99 | 7.51 | 10.99 | 99.6 | 19.8 | 12.8
2007-05-29 | 19:45 | 10.82 | 7.5 | 11.04 | 99.7 | 19.6 | 12.8
2007-05-29 | 20:00 | 10.64 | 7.48 | 11.06 | 99.5 | 19.6 | 12.8
2007-05-29 | 20:15 | 10.45 | 7.47 | 11.13 | 99.7 | 19.7 | 12.8
2007-05-29 | 20:30 | 10.27 | 7.47 | 11.17 | 99.5 | 19.6 | 12.8
2007-05-29 | 20:45 | 10.08 | 7.46 | 11.22 | 99.6 | 19.6 | 12.8
2007-05-29 | 21:00 | 9.89 | 7.46 | 11.28 | 99.7 | 19.6 | 12.7
2007-05-29 | 21:15 | 9.7 | 7.45 | 11.34 | 99.7 | 19.4 | 12.6
2007-05-29 | 21:30 | 9.51 | 7.44 | 11.39 | 99.7 | 19.4 | 12.6
2007-05-29 | 21:45 | 9.32 | 7.43 | 11.45 | 99.8 | 19.3 | 12.6
2007-05-29 | 22:00 | 9.12 | 7.43 | 11.45 | 99.4 | 19.2 | 12.5
2007-05-29 | 22:15 | 8.92 | 7.43 | 11.54 | 99.6 | 19.2 | 12.5
2007-05-29 | 22:30 | 8.72 | 7.42 | 11.61 | 99.8 | 19.2 | 12.5
2007-05-29 | 22:45 | 8.53 | 7.39 | 11.66 | 99.7 | 19.1 | 12.4
2007-05-29 | 23:00 | 8.33 | 7.42 | 11.71 | 99.7 | 19 | 12.4
2007-05-29 | 23:15 | 8.14 | 7.4 | 11.77 | 99.8 | 19 | 12.4
2007-05-29 | 23:30 | 7.96 | 7.41 | 11.81 | 99.7 | 19 | 12.4
2007-05-29 | 23:45 | 7.78 | 7.4 | 11.87 | 99.7 | 19 | 12.4
2007-05-30 | 0:00 | 7.6 | 7.4 | 11.94 | 99.9 | 19 | 12.3
2007-05-30 | 0:15 | 7.43 | 7.4 | 11.98 | 99.8 | 18.8 | 12.2
2007-05-30 | 0:30 | 7.27 | 7.4 | 12.04 | 99.9 | 18.8 | 12.2
2007-05-30 | 0:45 | 7.11 | 7.39 | 12.07 | 99.7 | 18.7 | 12.2
2007-05-30 | 1:01 | 6.96 | 7.39 | 12.13 | 99.8 | 18.9 | 12.3
2007-05-30 | 1:16 | 6.8 | 7.39 | 12.18 | 99.8 | 18.9 | 12.3
2007-05-30 | 1:31 | 6.66 | 7.39 | 12.21 | 99.7 | 19 | 12.3
2007-05-30 | 1:46 | 6.52 | 7.39 | 12.25 | 99.7 | 18.9 | 12.3
2007-05-30 | 2:01 | 6.38 | 7.39 | 12.3 | 99.8 | 18.8 | 12.2
2007-05-30 | 2:16 | 6.25 | 7.38 | 12.31 | 99.5 | 18.9 | 12.3
2007-05-30 | 2:31 | 6.11 | 7.38 | 12.37 | 99.7 | 18.8 | 12.2
2007-05-30 | 2:46 | 5.98 | 7.38 | 12.42 | 99.7 | 18.9 | 12.3
2007-05-30 | 3:01 | 5.86 | 7.38 | 12.44 | 99.6 | 18.9 | 12.3
2007-05-30 | 3:16 | 5.75 | 7.38 | 12.49 | 99.7 | 19 | 12.4
2007-05-30 | 3:31 | 5.63 | 7.38 | 12.5 | 99.5 | 19 | 12.3
2007-05-30 | 3:46 | 5.53 | 7.38 | 12.55 | 99.6 | 18.9 | 12.3
2007-05-30 | 4:01 | 5.43 | 7.38 | 12.58 | 99.6 | 18.8 | 12.3
2007-05-30 | 4:16 | 5.33 | 7.38 | 12.62 | 99.6 | 18.9 | 12.3
2007-05-30 | 4:31 | 5.24 | 7.38 | 12.64 | 99.6 | 18.9 | 12.3
2007-05-30 | 4:46 | 5.16 | 7.38 | 12.66 | 99.5 | 18.9 | 12.3
2007-05-30 | 5:01 | 5.08 | 7.38 | 12.7 | 99.6 | 18.8 | 12.2
2007-05-30 | 5:16 | 5 | 7.38 | 12.72 | 99.6 | 18.9 | 12.3
2007-05-30 | 5:31 | 4.93 | 7.39 | 12.78 | 99.9 | 18.8 | 12.2
2007-05-30 | 5:46 | 4.86 | 7.38 | 12.77 | 99.6 | 18.8 | 12.2
2007-05-30 | 6:01 | 4.8 | 7.38 | 12.81 | 99.8 | 18.8 | 12.2
2007-05-30 | 6:16 | 4.74 | 7.38 | 12.82 | 99.7 | 18.7 | 12.2
2007-05-30 | 6:31 | 4.69 | 7.39 | 12.83 | 99.7 | 18.7 | 12.2
2007-05-30 | 6:46 | 4.66 | 7.39 | 12.85 | 99.7 | 18.8 | 12.2
2007-05-30 | 7:01 | 4.63 | 7.39 | 12.86 | 99.8 | 18.8 | 12.2
2007-05-30 | 7:16 | 4.61 | 7.39 | 12.88 | 99.8 | 18.8 | 12.2
2007-05-30 | 7:31 | 4.6 | 7.39 | 12.88 | 99.8 | 18.7 | 12.2
2007-05-30 | 7:46 | 4.6 | 7.4 | 12.88 | 99.8 | 18.7 | 12.2
2007-05-30 | 8:01 | 4.61 | 7.4 | 12.9 | 100 | 19.1 | 12.4
2007-05-30 | 8:16 | 4.62 | 7.41 | 12.89 | 99.9 | 19 | 12.3
2007-05-30 | 8:31 | 4.65 | 7.42 | 12.89 | 100 | 18.9 | 12.3
2007-05-30 | 8:46 | 4.68 | 7.43 | 12.87 | 100 | 18.8 | 12.2
2007-05-30 | 9:01 | 4.74 | 7.43 | 12.85 | 99.9 | 18.7 | 12.2
2007-05-30 | 9:16 | 4.81 | 7.45 | 12.82 | 99.9 | 18.7 | 12.1
2007-05-30 | 9:31 | 4.91 | 7.44 | 12.81 | 100.1 | 18.7 | 12.2
2007-05-30 | 9:46 | 5.02 | 7.46 | 12.76 | 99.9 | 18.8 | 12.2
2007-05-30 | 10:01 | 5.15 | 7.45 | 12.72 | 100 | 18.7 | 12.2
2007-05-30 | 10:16 | 5.3 | 7.46 | 12.67 | 100 | 18.6 | 12.1
2007-05-30 | 10:31 | 5.46 | 7.47 | 12.62 | 100 | 18.6 | 12.1
2007-05-30 | 10:46 | 5.63 | 7.47 | 12.58 | 100.1 | 18.6 | 12.1
2007-05-30 | 11:01 | 5.82 | 7.47 | 12.51 | 100 | 18.7 | 12.2
2007-05-30 | 11:16 | 6.04 | 7.47 | 12.44 | 100 | 18.6 | 12.1
2007-05-30 | 11:31 | 6.28 | 7.47 | 12.37 | 100.1 | 18.6 | 12.1
2007-05-30 | 11:46 | 6.54 | 7.48 | 12.29 | 100.1 | 18.6 | 12.1
2007-05-30 | 12:01 | 6.81 | 7.48 | 12.22 | 100.2 | 18.6 | 12.1
2007-05-30 | 12:16 | 7.11 | 7.48 | 12.14 | 100.3 | 18.6 | 12.1
2007-05-30 | 12:31 | 7.42 | 7.48 | 12.05 | 100.3 | 18.7 | 12.1
2007-05-30 | 12:46 | 7.75 | 7.49 | 11.96 | 100.4 | 18.7 | 12.2
2007-05-30 | 13:00 | 8.1 | 7.5 | 11.85 | 100.3 | 18.7 | 12.1
2007-05-30 | 13:15 | 8.46 | 7.5 | 11.74 | 100.2 | 18.7 | 12.1
2007-05-30 | 13:30 | 8.83 | 7.5 | 11.63 | 100.2 | 18.7 | 12.2
2007-05-30 | 13:45 | 9.19 | 7.5 | 11.55 | 100.4 | 18.7 | 12.2
2007-05-30 | 14:00 | 9.55 | 7.51 | 11.45 | 100.4 | 18.7 | 12.2
2007-05-30 | 14:15 | 9.9 | 7.51 | 11.35 | 100.3 | 18.7 | 12.2
2007-05-30 | 14:30 | 10.23 | 7.52 | 11.26 | 100.3 | 18.7 | 12.1
2007-05-30 | 14:45 | 10.56 | 7.53 | 11.16 | 100.2 | 18.7 | 12.2
2007-05-30 | 15:00 | 10.88 | 7.53 | 11.09 | 100.3 | 18.7 | 12.2
2007-05-30 | 15:15 | 11.18 | 7.52 | 10.99 | 100.1 | 18.7 | 12.2
2007-05-30 | 15:30 | 11.46 | 7.54 | 10.92 | 100.1 | 18.7 | 12.2
2007-05-30 | 15:45 | 11.73 | 7.53 | 10.84 | 100 | 18.8 | 12.2
2007-05-30 | 16:00 | 11.97 | 7.53 | 10.79 | 100.1 | 18.8 | 12.2
2007-05-30 | 16:15 | 12.2 | 7.53 | 10.74 | 100.1 | 18.7 | 12.2
2007-05-30 | 16:30 | 12.39 | 7.55 | 10.69 | 100 | 19.4 | 12.6
2007-05-30 | 16:45 | 12.54 | 7.54 | 10.65 | 100.1 | 19.3 | 12.6
2007-05-30 | 17:00 | 12.67 | 7.55 | 10.61 | 99.9 | 19.2 | 12.5
2007-05-30 | 17:15 | 12.76 | 7.54 | 10.58 | 99.9 | 19.2 | 12.5
2007-05-30 | 17:30 | 12.83 | 7.54 | 10.57 | 99.9 | 19.2 | 12.5
2007-05-30 | 17:45 | 12.86 | 7.54 | 10.57 | 100 | 19.2 | 12.5
2007-05-30 | 18:00 | 12.86 | 7.54 | 10.55 | 99.8 | 19.2 | 12.5
2007-05-30 | 18:15 | 12.82 | 7.53 | 10.54 | 99.6 | 19.1 | 12.4
2007-05-30 | 18:30 | 12.75 | 7.53 | 10.56 | 99.7 | 19 | 12.4
2007-05-30 | 18:45 | 12.63 | 7.53 | 10.58 | 99.6 | 19.1 | 12.4
2007-05-30 | 19:00 | 12.48 | 7.52 | 10.63 | 99.7 | 19 | 12.3
2007-05-30 | 19:15 | 12.33 | 7.52 | 10.66 | 99.7 | 19.1 | 12.4
2007-05-30 | 19:30 | 12.16 | 7.51 | 10.7 | 99.7 | 19 | 12.4
2007-05-30 | 19:45 | 11.99 | 7.49 | 10.75 | 99.7 | 19 | 12.3
2007-05-30 | 20:00 | 11.82 | 7.49 | 10.75 | 99.3 | 18.8 | 12.2
2007-05-30 | 20:15 | 11.63 | 7.48 | 10.83 | 99.7 | 18.7 | 12.2
2007-05-30 | 20:30 | 11.44 | 7.47 | 10.88 | 99.7 | 18.7 | 12.2
2007-05-30 | 20:45 | 11.24 | 7.46 | 10.92 | 99.6 | 18.6 | 12.1
2007-05-30 | 21:00 | 11.02 | 7.45 | 10.98 | 99.6 | 18.5 | 12
2007-05-30 | 21:15 | 10.78 | 7.45 | 11.06 | 99.7 | 18.4 | 12
2007-05-30 | 21:30 | 10.54 | 7.44 | 11.12 | 99.8 | 18.3 | 11.9
2007-05-30 | 21:45 | 10.31 | 7.43 | 11.18 | 99.8 | 18.2 | 11.8
2007-05-30 | 22:00 | 10.08 | 7.43 | 11.24 | 99.7 | 18.2 | 11.8
2007-05-30 | 22:15 | 9.87 | 7.42 | 11.29 | 99.7 | 18.1 | 11.8
2007-05-30 | 22:30 | 9.65 | 7.42 | 11.35 | 99.7 | 18.1 | 11.8
2007-05-30 | 22:45 | 9.45 | 7.41 | 11.41 | 99.8 | 18 | 11.7
2007-05-30 | 23:00 | 9.24 | 7.41 | 11.47 | 99.8 | 18 | 11.7
2007-05-30 | 23:15 | 9.04 | 7.4 | 11.51 | 99.7 | 17.9 | 11.6
2007-05-30 | 23:30 | 8.85 | 7.4 | 11.57 | 99.7 | 17.7 | 11.5
2007-05-30 | 23:45 | 8.66 | 7.39 | 11.62 | 99.7 | 17.7 | 11.5
2007-05-31 | 0:00 | 8.48 | 7.39 | 11.67 | 99.7 | 17.6 | 11.4
2007-05-31 | 0:15 | 8.3 | 7.39 | 11.73 | 99.7 | 17.5 | 11.4
2007-05-31 | 0:30 | 8.12 | 7.39 | 11.77 | 99.7 | 17.4 | 11.3
2007-05-31 | 0:45 | 7.96 | 7.38 | 11.81 | 99.6 | 17.3 | 11.2
2007-05-31 | 1:01 | 7.8 | 7.38 | 11.87 | 99.7 | 17.4 | 11.3
2007-05-31 | 1:16 | 7.63 | 7.38 | 11.9 | 99.6 | 17.4 | 11.3
2007-05-31 | 1:31 | 7.48 | 7.38 | 11.95 | 99.6 | 17.3 | 11.2
2007-05-31 | 1:46 | 7.33 | 7.37 | 11.99 | 99.6 | 17.7 | 11.5
2007-05-31 | 2:01 | 7.19 | 7.37 | 12.06 | 99.8 | 17.5 | 11.4
2007-05-31 | 2:16 | 7.04 | 7.38 | 12.08 | 99.6 | 17.3 | 11.3
2007-05-31 | 2:31 | 6.9 | 7.37 | 12.13 | 99.6 | 17.2 | 11.2
2007-05-31 | 2:46 | 6.76 | 7.37 | 12.16 | 99.6 | 17 | 11.1
2007-05-31 | 3:01 | 6.63 | 7.36 | 12.2 | 99.6 | 17 | 11
2007-05-31 | 3:16 | 6.51 | 7.37 | 12.24 | 99.6 | 16.9 | 11
2007-05-31 | 3:31 | 6.4 | 7.36 | 12.28 | 99.7 | 17 | 11
2007-05-31 | 3:46 | 6.28 | 7.37 | 12.3 | 99.6 | 16.8 | 10.9
2007-05-31 | 4:01 | 6.18 | 7.36 | 12.34 | 99.6 | 17 | 11
2007-05-31 | 4:16 | 6.08 | 7.36 | 12.33 | 99.2 | 17 | 11.1
2007-05-31 | 4:31 | 5.98 | 7.36 | 12.39 | 99.5 | 17.1 | 11.1
2007-05-31 | 4:46 | 5.89 | 7.36 | 12.42 | 99.5 | 17.1 | 11.1
2007-05-31 | 5:01 | 5.81 | 7.36 | 12.44 | 99.4 | 17 | 11.1
2007-05-31 | 5:16 | 5.73 | 7.36 | 12.48 | 99.6 | 16.9 | 11
2007-05-31 | 5:31 | 5.65 | 7.36 | 12.5 | 99.5 | 16.9 | 11
2007-05-31 | 5:46 | 5.58 | 7.36 | 12.53 | 99.6 | 16.9 | 11
2007-05-31 | 6:01 | 5.51 | 7.36 | 12.55 | 99.5 | 16.9 | 11
2007-05-31 | 6:16 | 5.46 | 7.36 | 12.56 | 99.5 | 16.9 | 11
2007-05-31 | 6:31 | 5.41 | 7.36 | 12.58 | 99.5 | 16.9 | 11
2007-05-31 | 6:46 | 5.37 | 7.36 | 12.57 | 99.4 | 16.9 | 11
2007-05-31 | 7:01 | 5.34 | 7.37 | 12.62 | 99.6 | 16.9 | 11
2007-05-31 | 7:16 | 5.32 | 7.37 | 12.62 | 99.6 | 17.3 | 11.3
2007-05-31 | 7:31 | 5.3 | 7.38 | 12.63 | 99.6 | 17.4 | 11.3
2007-05-31 | 7:46 | 5.3 | 7.38 | 12.63 | 99.7 | 17.3 | 11.3
2007-05-31 | 8:01 | 5.31 | 7.39 | 12.63 | 99.6 | 17.3 | 11.3
2007-05-31 | 8:16 | 5.31 | 7.4 | 12.63 | 99.7 | 17.3 | 11.3
2007-05-31 | 8:31 | 5.33 | 7.4 | 12.62 | 99.7 | 17.2 | 11.2
2007-05-31 | 8:46 | 5.37 | 7.42 | 12.62 | 99.7 | 17.2 | 11.2
2007-05-31 | 9:01 | 5.42 | 7.42 | 12.61 | 99.8 | 17.3 | 11.2
2007-05-31 | 9:16 | 5.5 | 7.43 | 12.58 | 99.8 | 17.2 | 11.2
2007-05-31 | 9:31 | 5.59 | 7.44 | 12.56 | 99.8 | 17.3 | 11.3
2007-05-31 | 9:46 | 5.7 | 7.44 | 12.53 | 99.9 | 17.3 | 11.3
2007-05-31 | 10:01 | 5.82 | 7.44 | 12.49 | 99.9 | 17.3 | 11.2
2007-05-31 | 10:16 | 5.96 | 7.45 | 12.43 | 99.8 | 17.3 | 11.2
2007-05-31 | 10:31 | 6.12 | 7.44 | 12.41 | 100 | 17.3 | 11.3
2007-05-31 | 10:46 | 6.3 | 7.45 | 12.34 | 99.9 | 17.3 | 11.2
2007-05-31 | 11:01 | 6.48 | 7.45 | 12.29 | 99.9 | 17.3 | 11.2
2007-05-31 | 11:16 | 6.7 | 7.46 | 12.22 | 99.9 | 17.3 | 11.3
2007-05-31 | 11:31 | 6.94 | 7.46 | 12.14 | 99.9 | 17.3 | 11.3
2007-05-31 | 11:46 | 7.18 | 7.47 | 12.08 | 99.9 | 17.3 | 11.3
2007-05-31 | 12:01 | 7.46 | 7.47 | 12 | 100 | 17.2 | 11.2
2007-05-31 | 12:16 | 7.76 | 7.47 | 11.92 | 100.1 | 17.3 | 11.3
2007-05-31 | 12:31 | 8.07 | 7.47 | 11.86 | 100.3 | 17.4 | 11.3
2007-05-31 | 12:46 | 8.4 | 7.48 | 11.75 | 100.1 | 17.4 | 11.3
2007-05-31 | 13:00 | 8.75 | 7.48 | 11.64 | 100.1 | 17.4 | 11.3
2007-05-31 | 13:15 | 9.11 | 7.49 | 11.52 | 99.9 | 17.7 | 11.5
2007-05-31 | 13:30 | 9.47 | 7.49 | 11.44 | 100.1 | 17.7 | 11.5
2007-05-31 | 13:45 | 9.82 | 7.5 | 11.35 | 100.1 | 17.8 | 11.6
2007-05-31 | 14:00 | 10.16 | 7.5 | 11.25 | 100 | 17.8 | 11.6
2007-05-31 | 14:15 | 10.5 | 7.51 | 11.17 | 100.1 | 17.7 | 11.5
2007-05-31 | 14:30 | 10.83 | 7.51 | 11.07 | 100 | 17.9 | 11.6
2007-05-31 | 14:45 | 11.16 | 7.51 | 10.99 | 100 | 17.9 | 11.6
2007-05-31 | 15:00 | 11.43 | 7.51 | 10.91 | 99.9 | 18 | 11.7
2007-05-31 | 15:15 | 11.63 | 7.52 | 10.86 | 99.9 | 18.1 | 11.7
2007-05-31 | 15:30 | 11.72 | 7.52 | 10.81 | 99.6 | 17.9 | 11.7
2007-05-31 | 15:45 | 11.79 | 7.52 | 10.81 | 99.8 | 18 | 11.7
2007-05-31 | 16:00 | 11.95 | 7.52 | 10.77 | 99.8 | 18 | 11.7
2007-05-31 | 16:15 | 12.25 | 7.52 | 10.7 | 99.8 | 18.1 | 11.7
2007-05-31 | 16:30 | 12.47 | 7.52 | 10.65 | 99.8 | 18.1 | 11.7
2007-05-31 | 16:45 | 12.56 | 7.53 | 10.63 | 99.9 | 18.1 | 11.7
2007-05-31 | 17:00 | 12.54 | 7.53 | 10.63 | 99.8 | 18.1 | 11.8
2007-05-31 | 17:15 | 12.64 | 7.52 | 10.6 | 99.8 | 18.2 | 11.8
2007-05-31 | 17:30 | 12.58 | 7.53 | 10.59 | 99.5 | 18.1 | 11.8
2007-05-31 | 17:45 | 12.62 | 7.52 | 10.6 | 99.7 | 18.2 | 11.8
2007-05-31 | 18:00 | 12.63 | 7.52 | 10.57 | 99.5 | 18.2 | 11.8
2007-05-31 | 18:15 | 12.64 | 7.52 | 10.58 | 99.6 | 18.1 | 11.8
2007-05-31 | 18:30 | 12.71 | 7.52 | 10.56 | 99.6 | 18.2 | 11.9
2007-05-31 | 18:45 | 12.75 | 7.51 | 10.55 | 99.5 | 18.1 | 11.8
2007-05-31 | 19:00 | 12.74 | 7.51 | 10.56 | 99.7 | 18 | 11.7
2007-05-31 | 19:15 | 12.67 | 7.51 | 10.56 | 99.5 | 18 | 11.7
2007-05-31 | 19:30 | 12.55 | 7.5 | 10.59 | 99.5 | 17.9 | 11.6
2007-05-31 | 19:45 | 12.39 | 7.48 | 10.63 | 99.5 | 17.9 | 11.6
2007-05-31 | 20:00 | 12.21 | 7.49 | 10.68 | 99.6 | 18.1 | 11.8
2007-05-31 | 20:15 | 12.03 | 7.47 | 10.73 | 99.6 | 18 | 11.7
2007-05-31 | 20:30 | 11.85 | 7.45 | 10.76 | 99.5 | 18 | 11.7
2007-05-31 | 20:45 | 11.65 | 7.45 | 10.81 | 99.5 | 17.9 | 11.6
2007-05-31 | 21:00 | 11.44 | 7.43 | 10.88 | 99.6 | 17.9 | 11.6
2007-05-31 | 21:15 | 11.23 | 7.42 | 10.92 | 99.6 | 17.7 | 11.5
2007-05-31 | 21:30 | 11.01 | 7.41 | 11.01 | 99.8 | 17.4 | 11.3
2007-05-31 | 21:45 | 10.8 | 7.41 | 11.02 | 99.5 | 17.4 | 11.3
2007-05-31 | 22:00 | 10.57 | 7.4 | 11.09 | 99.6 | 17.2 | 11.2
2007-05-31 | 22:15 | 10.35 | 7.39 | 11.13 | 99.4 | 17 | 11.1
2007-05-31 | 22:30 | 10.14 | 7.39 | 11.21 | 99.6 | 16.9 | 11
2007-05-31 | 22:45 | 9.93 | 7.39 | 11.26 | 99.6 | 16.9 | 11
2007-05-31 | 23:00 | 9.73 | 7.38 | 11.32 | 99.7 | 16.7 | 10.9
2007-05-31 | 23:15 | 9.53 | 7.37 | 11.4 | 99.8 | 16.7 | 10.9
2007-05-31 | 23:30 | 9.34 | 7.38 | 11.43 | 99.7 | 17.4 | 11.3
2007-05-31 | 23:45 | 9.16 | 7.36 | 11.49 | 99.7 | 17.1 | 11.1
2007-06-01 | 0:00 | 8.97 | 7.36 | 11.52 | 99.5 | 17.2 | 11.2
2007-06-01 | 0:15 | 8.79 | 7.36 | 11.57 | 99.6 | 17 | 11.1
2007-06-01 | 0:30 | 8.61 | 7.36 | 11.64 | 99.7 | 16.9 | 11
2007-06-01 | 0:45 | 8.44 | 7.36 | 11.68 | 99.6 | 17.2 | 11.2
2007-06-01 | 1:01 | 8.27 | 7.35 | 11.72 | 99.6 | 17.2 | 11.2
2007-06-01 | 1:16 | 8.11 | 7.35 | 11.77 | 99.6 | 16.7 | 10.8
2007-06-01 | 1:31 | 7.95 | 7.35 | 11.81 | 99.6 | 16.8 | 10.9
2007-06-01 | 1:46 | 7.79 | 7.34 | 11.86 | 99.7 | 16.6 | 10.8
2007-06-01 | 2:01 | 7.65 | 7.34 | 11.88 | 99.5 | 16.7 | 10.9
2007-06-01 | 2:16 | 7.5 | 7.35 | 11.95 | 99.7 | 16.5 | 10.8
2007-06-01 | 2:31 | 7.37 | 7.34 | 11.97 | 99.5 | 16.3 | 10.6
2007-06-01 | 2:46 | 7.24 | 7.34 | 12 | 99.5 | 16 | 10.4
2007-06-01 | 3:01 | 7.12 | 7.33 | 12.05 | 99.6 | 16.7 | 10.9
2007-06-01 | 3:16 | 7 | 7.33 | 12.07 | 99.4 | 16.5 | 10.7
2007-06-01 | 3:31 | 6.89 | 7.33 | 12.1 | 99.4 | 16.2 | 10.5
2007-06-01 | 3:46 | 6.79 | 7.33 | 12.12 | 99.3 | 16.2 | 10.5
2007-06-01 | 4:01 | 6.69 | 7.34 | 12.14 | 99.2 | 16.2 | 10.5
2007-06-01 | 4:16 | 6.6 | 7.33 | 12.18 | 99.3 | 16.2 | 10.6
2007-06-01 | 4:31 | 6.52 | 7.33 | 12.2 | 99.3 | 17.4 | 11.3
2007-06-01 | 4:46 | 6.44 | 7.33 | 12.23 | 99.4 | 17.4 | 11.3
2007-06-01 | 5:01 | 6.36 | 7.33 | 12.26 | 99.3 | 17.5 | 11.3
2007-06-01 | 5:16 | 6.29 | 7.33 | 12.27 | 99.3 | 17.6 | 11.4
2007-06-01 | 5:31 | 6.22 | 7.33 | 12.3 | 99.3 | 17.9 | 11.7
2007-06-01 | 5:46 | 6.15 | 7.33 | 12.32 | 99.3 | 18 | 11.7
2007-06-01 | 6:01 | 6.09 | 7.34 | 12.35 | 99.4 | 18.1 | 11.8
2007-06-01 | 6:16 | 6.04 | 7.33 | 12.35 | 99.3 | 18.1 | 11.7
2007-06-01 | 6:31 | 5.99 | 7.33 | 12.37 | 99.4 | 18 | 11.7
2007-06-01 | 6:46 | 5.95 | 7.35 | 12.39 | 99.4 | 18 | 11.7
2007-06-01 | 7:01 | 5.92 | 7.35 | 12.41 | 99.5 | 18 | 11.7
2007-06-01 | 7:16 | 5.89 | 7.35 | 12.41 | 99.4 | 18 | 11.7
2007-06-01 | 7:31 | 5.87 | 7.35 | 12.42 | 99.4 | 17.9 | 11.6
2007-06-01 | 7:46 | 5.85 | 7.36 | 12.43 | 99.5 | 17.8 | 11.6
2007-06-01 | 8:01 | 5.85 | 7.37 | 12.43 | 99.4 | 17.9 | 11.6
2007-06-01 | 8:16 | 5.86 | 7.37 | 12.43 | 99.5 | 17.9 | 11.6
2007-06-01 | 8:31 | 5.87 | 7.39 | 12.41 | 99.4 | 17.9 | 11.6
2007-06-01 | 8:46 | 5.89 | 7.4 | 12.42 | 99.5 | 17.9 | 11.6
2007-06-01 | 9:01 | 5.95 | 7.4 | 12.39 | 99.4 | 17.9 | 11.6
2007-06-01 | 9:16 | 6.02 | 7.41 | 12.36 | 99.4 | 17.9 | 11.6
2007-06-01 | 9:31 | 6.09 | 7.41 | 12.34 | 99.4 | 17.8 | 11.6
2007-06-01 | 9:46 | 6.18 | 7.41 | 12.3 | 99.3 | 17.8 | 11.6
2007-06-01 | 10:01 | 6.29 | 7.42 | 12.29 | 99.5 | 17.8 | 11.5
2007-06-01 | 10:16 | 6.41 | 7.43 | 12.26 | 99.5 | 17.8 | 11.6
2007-06-01 | 10:31 | 6.57 | 7.43 | 12.21 | 99.5 | 17.8 | 11.6
2007-06-01 | 10:46 | 6.74 | 7.43 | 12.15 | 99.4 | 17.8 | 11.5
2007-06-01 | 11:01 | 6.93 | 7.45 | 12.1 | 99.5 | 17.8 | 11.6
2007-06-01 | 11:16 | 7.1 | 7.45 | 12.05 | 99.5 | 17.8 | 11.6
2007-06-01 | 11:31 | 7.34 | 7.45 | 11.98 | 99.5 | 17.8 | 11.6
2007-06-01 | 11:46 | 7.59 | 7.45 | 11.91 | 99.5 | 17.8 | 11.6
2007-06-01 | 12:01 | 7.84 | 7.45 | 11.83 | 99.5 | 17.8 | 11.6
2007-06-01 | 12:16 | 8.13 | 7.45 | 11.76 | 99.6 | 17.8 | 11.6
2007-06-01 | 12:31 | 8.42 | 7.46 | 11.68 | 99.6 | 17.8 | 11.6
2007-06-01 | 12:46 | 8.77 | 7.47 | 11.59 | 99.7 | 17.8 | 11.6
2007-06-01 | 13:00 | 9.1 | 7.47 | 11.49 | 99.6 | 17.8 | 11.6
2007-06-01 | 13:15 | 9.42 | 7.48 | 11.4 | 99.6 | 17.7 | 11.5
2007-06-01 | 13:30 | 9.77 | 7.48 | 11.31 | 99.6 | 17.7 | 11.5
2007-06-01 | 13:45 | 10.12 | 7.48 | 11.22 | 99.7 | 17.8 | 11.6
2007-06-01 | 14:00 | 10.46 | 7.49 | 11.13 | 99.7 | 17.8 | 11.6
2007-06-01 | 14:15 | 10.82 | 7.49 | 11.05 | 99.7 | 17.9 | 11.6
2007-06-01 | 14:30 | 11.16 | 7.49 | 10.95 | 99.7 | 17.9 | 11.6
2007-06-01 | 14:45 | 11.36 | 7.5 | 10.87 | 99.4 | 17.9 | 11.6
2007-06-01 | 15:00 | 11.48 | 7.49 | 10.83 | 99.3 | 17.9 | 11.6
2007-06-01 | 15:15 | 11.58 | 7.48 | 10.78 | 99.1 | 17.9 | 11.6
2007-06-01 | 15:30 | 11.69 | 7.48 | 10.76 | 99.1 | 17.9 | 11.7
2007-06-01 | 15:45 | 11.82 | 7.49 | 10.73 | 99.2 | 17.9 | 11.7
2007-06-01 | 16:00 | 11.91 | 7.48 | 10.72 | 99.2 | 18 | 11.7
2007-06-01 | 16:15 | 11.9 | 7.49 | 10.72 | 99.2 | 18 | 11.7
2007-06-01 | 16:30 | 12 | 7.48 | 10.71 | 99.4 | 18 | 11.7
2007-06-01 | 16:45 | 12.09 | 7.48 | 10.67 | 99.2 | 17.9 | 11.7
2007-06-01 | 17:00 | 12.06 | 7.48 | 10.67 | 99.1 | 17.9 | 11.6
2007-06-01 | 17:15 | 11.94 | 7.47 | 10.71 | 99.2 | 18.5 | 12
2007-06-01 | 17:30 | 11.84 | 7.46 | 10.72 | 99.1 | 18.4 | 12
2007-06-01 | 17:45 | 11.79 | 7.47 | 10.72 | 99 | 18.4 | 12
2007-06-01 | 18:00 | 11.75 | 7.46 | 10.75 | 99.2 | 18.5 | 12
2007-06-01 | 18:15 | 11.71 | 7.46 | 10.76 | 99.2 | 18.8 | 12.2
2007-06-01 | 18:30 | 11.62 | 7.46 | 10.78 | 99.2 | 19 | 12.4
2007-06-01 | 18:45 | 11.53 | 7.46 | 10.8 | 99.2 | 19 | 12.3
2007-06-01 | 19:00 | 11.43 | 7.46 | 10.84 | 99.3 | 18.8 | 12.2
2007-06-01 | 19:15 | 11.34 | 7.45 | 10.86 | 99.3 | 18.8 | 12.2
2007-06-01 | 19:30 | 11.24 | 7.44 | 10.89 | 99.3 | 18.6 | 12.1
2007-06-01 | 19:45 | 11.09 | 7.43 | 10.92 | 99.2 | 18.5 | 12
2007-06-01 | 20:00 | 10.9 | 7.43 | 10.97 | 99.2 | 18.3 | 11.9
2007-06-01 | 20:15 | 10.71 | 7.42 | 11 | 99.1 | 18.4 | 11.9
2007-06-01 | 20:30 | 10.54 | 7.41 | 11.06 | 99.2 | 18 | 11.7
2007-06-01 | 20:45 | 10.36 | 7.4 | 11.12 | 99.4 | 18 | 11.7
2007-06-01 | 21:00 | 10.18 | 7.38 | 11.16 | 99.3 | 18.2 | 11.8
2007-06-01 | 21:15 | 9.97 | 7.37 | 11.2 | 99.2 | 18.2 | 11.8
2007-06-01 | 21:30 | 9.76 | 7.37 | 11.26 | 99.2 | 18 | 11.7
2007-06-01 | 21:45 | 9.55 | 7.36 | 11.33 | 99.3 | 17.9 | 11.6
2007-06-01 | 22:00 | 9.36 | 7.35 | 11.39 | 99.3 | 17.8 | 11.6
2007-06-01 | 22:15 | 9.19 | 7.34 | 11.44 | 99.4 | 17.8 | 11.5
2007-06-01 | 22:30 | 9.04 | 7.34 | 11.46 | 99.2 | 17.6 | 11.5
2007-06-01 | 22:45 | 8.91 | 7.33 | 11.51 | 99.3 | 17.6 | 11.4
2007-06-01 | 23:00 | 8.8 | 7.32 | 11.53 | 99.3 | 17.4 | 11.3
2007-06-01 | 23:15 | 8.69 | 7.33 | 11.57 | 99.3 | 18.2 | 11.8
2007-06-01 | 23:30 | 8.58 | 7.32 | 11.59 | 99.3 | 18.1 | 11.8
2007-06-01 | 23:45 | 8.48 | 7.31 | 11.63 | 99.3 | 18 | 11.7
2007-06-02 | 0:00 | 8.37 | 7.31 | 11.64 | 99.2 | 18 | 11.7
2007-06-02 | 0:15 | 8.26 | 7.32 | 11.7 | 99.4 | 17.8 | 11.6
2007-06-02 | 0:30 | 8.15 | 7.32 | 11.74 | 99.5 | 17.8 | 11.5
2007-06-02 | 0:45 | 8.05 | 7.31 | 11.72 | 99.1 | 17.7 | 11.5
2007-06-02 | 1:01 | 7.95 | 7.31 | 11.76 | 99.2 | 18 | 11.7
2007-06-02 | 1:16 | 7.84 | 7.3 | 11.79 | 99.2 | 18.5 | 12
2007-06-02 | 1:31 | 7.74 | 7.29 | 11.82 | 99.1 | 18.6 | 12.1
2007-06-02 | 1:46 | 7.63 | 7.29 | 11.83 | 99 | 18.6 | 12.1
2007-06-02 | 2:01 | 7.53 | 7.29 | 11.87 | 99.1 | 18.4 | 12
2007-06-02 | 2:16 | 7.44 | 7.3 | 11.92 | 99.3 | 18.1 | 11.8
2007-06-02 | 2:31 | 7.34 | 7.3 | 11.93 | 99.1 | 17.8 | 11.6
2007-06-02 | 2:46 | 7.25 | 7.28 | 11.96 | 99.1 | 17.6 | 11.5
2007-06-02 | 3:01 | 7.17 | 7.3 | 11.98 | 99.1 | 17.5 | 11.3
2007-06-02 | 3:16 | 7.09 | 7.3 | 12 | 99.1 | 17.4 | 11.3
2007-06-02 | 3:31 | 7.01 | 7.29 | 12.03 | 99.1 | 17.4 | 11.3
2007-06-02 | 3:46 | 6.94 | 7.3 | 12.04 | 99 | 17.4 | 11.3
2007-06-02 | 4:01 | 6.87 | 7.29 | 12.07 | 99.1 | 17.4 | 11.3
2007-06-02 | 4:16 | 6.81 | 7.29 | 12.09 | 99.2 | 17.4 | 11.3
2007-06-02 | 4:31 | 6.76 | 7.29 | 12.1 | 99.1 | 17.9 | 11.7
2007-06-02 | 4:46 | 6.71 | 7.29 | 12.13 | 99.1 | 18 | 11.7
2007-06-02 | 5:01 | 6.66 | 7.29 | 12.14 | 99.1 | 18 | 11.7
2007-06-02 | 5:16 | 6.61 | 7.28 | 12.15 | 99.1 | 17.8 | 11.6
2007-06-02 | 5:31 | 6.56 | 7.28 | 12.17 | 99.1 | 17.7 | 11.5
2007-06-02 | 5:46 | 6.52 | 7.29 | 12.19 | 99.2 | 17.9 | 11.6
2007-06-02 | 6:01 | 6.48 | 7.29 | 12.2 | 99.2 | 17.9 | 11.6
2007-06-02 | 6:16 | 6.43 | 7.29 | 12.22 | 99.2 | 17.9 | 11.6
2007-06-02 | 6:31 | 6.39 | 7.28 | 12.22 | 99.2 | 17.9 | 11.6
2007-06-02 | 6:46 | 6.35 | 7.3 | 12.25 | 99.3 | 17.9 | 11.6
2007-06-02 | 7:01 | 6.31 | 7.29 | 12.28 | 99.4 | 18 | 11.7
2007-06-02 | 7:16 | 6.28 | 7.3 | 12.28 | 99.4 | 18.1 | 11.7
2007-06-02 | 7:31 | 6.26 | 7.3 | 12.29 | 99.4 | 18.1 | 11.8
2007-06-02 | 7:46 | 6.25 | 7.31 | 12.29 | 99.4 | 18 | 11.7
2007-06-02 | 8:01 | 6.24 | 7.33 | 12.31 | 99.5 | 17.9 | 11.6
2007-06-02 | 8:16 | 6.25 | 7.34 | 12.31 | 99.5 | 17.7 | 11.5
2007-06-02 | 8:31 | 6.27 | 7.35 | 12.31 | 99.5 | 17.6 | 11.5
2007-06-02 | 8:46 | 6.3 | 7.37 | 12.29 | 99.5 | 17.5 | 11.4
2007-06-02 | 9:01 | 6.34 | 7.38 | 12.28 | 99.5 | 17.8 | 11.6
2007-06-02 | 9:16 | 6.38 | 7.37 | 12.27 | 99.5 | 17.8 | 11.6
2007-06-02 | 9:31 | 6.43 | 7.38 | 12.25 | 99.5 | 17.6 | 11.4
2007-06-02 | 9:46 | 6.52 | 7.38 | 12.22 | 99.5 | 17.6 | 11.4
2007-06-02 | 10:01 | 6.63 | 7.39 | 12.19 | 99.5 | 17.7 | 11.5
2007-06-02 | 10:16 | 6.76 | 7.4 | 12.14 | 99.4 | 17.6 | 11.4
2007-06-02 | 10:31 | 6.89 | 7.4 | 12.1 | 99.4 | 17.5 | 11.4
2007-06-02 | 10:46 | 7.03 | 7.41 | 12.07 | 99.5 | 17.6 | 11.4
2007-06-02 | 11:01 | 7.21 | 7.41 | 12.02 | 99.5 | 17.7 | 11.5
2007-06-02 | 11:16 | 7.41 | 7.41 | 11.96 | 99.5 | 18.3 | 11.9
2007-06-02 | 11:31 | 7.63 | 7.41 | 11.9 | 99.6 | 18.2 | 11.8
2007-06-02 | 11:46 | 7.85 | 7.43 | 11.84 | 99.6 | 18.2 | 11.8
2007-06-02 | 12:01 | 8.11 | 7.42 | 11.76 | 99.6 | 18.1 | 11.8
2007-06-02 | 12:16 | 8.41 | 7.42 | 11.69 | 99.7 | 18.3 | 11.9
2007-06-02 | 12:31 | 8.71 | 7.44 | 11.62 | 99.8 | 18.2 | 11.8
2007-06-02 | 12:46 | 9.02 | 7.43 | 11.52 | 99.7 | 18.1 | 11.8
2007-06-02 | 13:00 | 9.35 | 7.44 | 11.43 | 99.7 | 17.9 | 11.7
2007-06-02 | 13:15 | 9.69 | 7.44 | 11.34 | 99.7 | 17.9 | 11.6
2007-06-02 | 13:30 | 10.03 | 7.44 | 11.25 | 99.7 | 17.9 | 11.7
2007-06-02 | 13:45 | 10.35 | 7.45 | 11.17 | 99.7 | 17.9 | 11.6
2007-06-02 | 14:00 | 10.69 | 7.45 | 11.07 | 99.7 | 17.8 | 11.6
2007-06-02 | 14:15 | 10.97 | 7.46 | 11 | 99.7 | 18 | 11.7
2007-06-02 | 14:30 | 11.26 | 7.47 | 10.92 | 99.6 | 17.9 | 11.6
2007-06-02 | 14:45 | 11.47 | 7.46 | 10.84 | 99.4 | 17.8 | 11.6
2007-06-02 | 15:00 | 11.89 | 7.47 | 10.76 | 99.6 | 17.9 | 11.6
2007-06-02 | 15:15 | 12.17 | 7.47 | 10.69 | 99.6 | 17.8 | 11.6
2007-06-02 | 15:30 | 12.19 | 7.47 | 10.65 | 99.2 | 17.8 | 11.6
2007-06-02 | 15:45 | 12.14 | 7.45 | 10.66 | 99.2 | 17.8 | 11.5
2007-06-02 | 16:00 | 12 | 7.46 | 10.71 | 99.4 | 17.7 | 11.5
2007-06-02 | 16:15 | 11.85 | 7.44 | 10.74 | 99.3 | 17.7 | 11.5
2007-06-02 | 16:30 | 11.71 | 7.44 | 10.77 | 99.3 | 17.7 | 11.5
2007-06-02 | 16:45 | 11.71 | 7.44 | 10.78 | 99.4 | 17.6 | 11.5
2007-06-02 | 17:00 | 11.72 | 7.44 | 10.78 | 99.4 | 17.6 | 11.5
2007-06-02 | 17:15 | 11.62 | 7.45 | 10.8 | 99.4 | 18.3 | 11.9
2007-06-02 | 17:30 | 11.61 | 7.44 | 10.8 | 99.3 | 18 | 11.7
2007-06-02 | 17:45 | 11.64 | 7.46 | 10.79 | 99.3 | 17.7 | 11.5
2007-06-02 | 18:00 | 11.68 | 7.44 | 10.78 | 99.3 | 17.5 | 11.4
2007-06-02 | 18:15 | 11.66 | 7.43 | 10.78 | 99.3 | 17.1 | 11.1
2007-06-02 | 18:30 | 11.57 | 7.43 | 10.8 | 99.2 | 17.4 | 11.3
2007-06-02 | 18:45 | 11.47 | 7.43 | 10.84 | 99.4 | 17.4 | 11.3
2007-06-02 | 19:00 | 11.34 | 7.42 | 10.87 | 99.3 | 17.6 | 11.4
2007-06-02 | 19:15 | 11.18 | 7.41 | 10.9 | 99.3 | 17.4 | 11.3
2007-06-02 | 19:30 | 11 | 7.41 | 10.94 | 99.2 | 18 | 11.7
2007-06-02 | 19:45 | 10.83 | 7.4 | 10.99 | 99.3 | 17.8 | 11.6
2007-06-02 | 20:00 | 10.66 | 7.4 | 11.03 | 99.3 | 17.4 | 11.3
2007-06-02 | 20:15 | 10.47 | 7.39 | 11.08 | 99.3 | 17.3 | 11.3
2007-06-02 | 20:30 | 10.26 | 7.36 | 11.12 | 99.1 | 17.4 | 11.3
2007-06-02 | 20:45 | 10.03 | 7.37 | 11.19 | 99.2 | 17.2 | 11.2
2007-06-02 | 21:00 | 9.8 | 7.34 | 11.25 | 99.2 | 16.7 | 10.9
2007-06-02 | 21:15 | 9.56 | 7.36 | 11.34 | 99.4 | 16.9 | 11
2007-06-02 | 21:30 | 9.33 | 7.34 | 11.34 | 98.9 | 16.9 | 11
2007-06-02 | 21:45 | 9.13 | 7.33 | 11.43 | 99.2 | 17.5 | 11.4
2007-06-02 | 22:00 | 8.94 | 7.32 | 11.48 | 99.2 | 16.9 | 11
2007-06-02 | 22:15 | 8.78 | 7.33 | 11.52 | 99.1 | 16.9 | 11
2007-06-02 | 22:30 | 8.62 | 7.31 | 11.58 | 99.3 | 16.9 | 11
2007-06-02 | 22:45 | 8.47 | 7.31 | 11.62 | 99.2 | 16.8 | 10.9
2007-06-02 | 23:00 | 8.33 | 7.31 | 11.65 | 99.2 | 16.7 | 10.9
2007-06-02 | 23:15 | 8.19 | 7.3 | 11.69 | 99.2 | 16.7 | 10.9
2007-06-02 | 23:30 | 8.07 | 7.32 | 11.73 | 99.2 | 16.6 | 10.8
2007-06-02 | 23:45 | 7.95 | 7.31 | 11.76 | 99.2 | 16.6 | 10.8
2007-06-03 | 0:00 | 7.85 | 7.3 | 11.8 | 99.2 | 16.4 | 10.7
2007-06-03 | 0:15 | 7.76 | 7.3 | 11.82 | 99.2 | 16.4 | 10.6
2007-06-03 | 0:30 | 7.67 | 7.3 | 11.83 | 99.1 | 16.3 | 10.6
2007-06-03 | 0:45 | 7.58 | 7.3 | 11.88 | 99.2 | 15.9 | 10.4
2007-06-03 | 1:01 | 7.49 | 7.31 | 11.88 | 99.1 | 16 | 10.4
2007-06-03 | 1:16 | 7.4 | 7.3 | 11.92 | 99.2 | 15.8 | 10.3
2007-06-03 | 1:31 | 7.31 | 7.3 | 11.93 | 99.1 | 15.8 | 10.3
2007-06-03 | 1:46 | 7.21 | 7.3 | 11.98 | 99.2 | 15.9 | 10.3
2007-06-03 | 2:01 | 7.12 | 7.29 | 12 | 99.2 | 15.8 | 10.3
2007-06-03 | 2:16 | 7.02 | 7.3 | 12.04 | 99.2 | 16.4 | 10.7
2007-06-03 | 2:31 | 6.92 | 7.29 | 12.06 | 99.1 | 16.6 | 10.8
2007-06-03 | 2:46 | 6.83 | 7.29 | 12.1 | 99.2 | 16.6 | 10.8
2007-06-03 | 3:01 | 6.75 | 7.29 | 12.12 | 99.2 | 16.5 | 10.7
2007-06-03 | 3:16 | 6.67 | 7.28 | 12.14 | 99.2 | 16.6 | 10.8
2007-06-03 | 3:31 | 6.59 | 7.29 | 12.18 | 99.3 | 16.5 | 10.7
2007-06-03 | 3:46 | 6.52 | 7.29 | 12.19 | 99.2 | 16.4 | 10.7
2007-06-03 | 4:01 | 6.45 | 7.29 | 12.2 | 99.1 | 16.6 | 10.8
2007-06-03 | 4:16 | 6.39 | 7.29 | 12.23 | 99.2 | 16.9 | 11
2007-06-03 | 4:31 | 6.32 | 7.29 | 12.25 | 99.2 | 16.7 | 10.9
2007-06-03 | 4:46 | 6.26 | 7.28 | 12.28 | 99.3 | 17.5 | 11.4
2007-06-03 | 5:01 | 6.2 | 7.27 | 12.3 | 99.3 | 17.7 | 11.5
2007-06-03 | 5:16 | 6.14 | 7.28 | 12.31 | 99.2 | 17.7 | 11.5
2007-06-03 | 5:31 | 6.09 | 7.28 | 12.34 | 99.3 | 17.4 | 11.3
2007-06-03 | 5:46 | 6.04 | 7.29 | 12.35 | 99.3 | 16.9 | 11
2007-06-03 | 6:01 | 5.99 | 7.29 | 12.37 | 99.3 | 16.8 | 10.9
2007-06-03 | 6:16 | 5.95 | 7.29 | 12.38 | 99.3 | 16.6 | 10.8
2007-06-03 | 6:31 | 5.91 | 7.29 | 12.4 | 99.3 | 16.4 | 10.6
2007-06-03 | 6:46 | 5.87 | 7.29 | 12.4 | 99.3 | 16.7 | 10.8
2007-06-03 | 7:01 | 5.85 | 7.3 | 12.42 | 99.4 | 16.6 | 10.8
2007-06-03 | 7:16 | 5.83 | 7.3 | 12.43 | 99.4 | 16.6 | 10.8
2007-06-03 | 7:31 | 5.81 | 7.3 | 12.44 | 99.5 | 17 | 11
2007-06-03 | 7:46 | 5.81 | 7.31 | 12.43 | 99.4 | 17 | 11
2007-06-03 | 8:01 | 5.8 | 7.32 | 12.45 | 99.5 | 16.9 | 11
2007-06-03 | 8:16 | 5.82 | 7.33 | 12.45 | 99.5 | 16.8 | 11
2007-06-03 | 8:31 | 5.85 | 7.34 | 12.42 | 99.4 | 17 | 11.1
2007-06-03 | 8:46 | 5.9 | 7.36 | 12.42 | 99.6 | 16.9 | 11
2007-06-03 | 9:01 | 5.96 | 7.37 | 12.4 | 99.5 | 16.9 | 11
2007-06-03 | 9:16 | 6.03 | 7.37 | 12.38 | 99.5 | 16.8 | 10.9
2007-06-03 | 9:31 | 6.11 | 7.38 | 12.36 | 99.6 | 16.8 | 10.9
2007-06-03 | 9:46 | 6.22 | 7.39 | 12.34 | 99.6 | 16.9 | 11
2007-06-03 | 10:01 | 6.34 | 7.39 | 12.3 | 99.6 | 16.8 | 10.9
2007-06-03 | 10:16 | 6.48 | 7.4 | 12.26 | 99.7 | 16.8 | 10.9
2007-06-03 | 10:31 | 6.63 | 7.4 | 12.2 | 99.6 | 16.8 | 10.9
2007-06-03 | 10:46 | 6.79 | 7.4 | 12.15 | 99.6 | 16.8 | 10.9
2007-06-03 | 11:01 | 6.98 | 7.41 | 12.1 | 99.7 | 16.9 | 11
2007-06-03 | 11:16 | 7.19 | 7.41 | 12.04 | 99.7 | 17 | 11
2007-06-03 | 11:31 | 7.43 | 7.42 | 11.97 | 99.7 | 16.9 | 11
2007-06-03 | 11:46 | 7.69 | 7.42 | 11.9 | 99.7 | 17 | 11
2007-06-03 | 12:01 | 7.97 | 7.42 | 11.83 | 99.8 | 17 | 11
2007-06-03 | 12:16 | 8.26 | 7.43 | 11.75 | 99.8 | 17 | 11
2007-06-03 | 12:31 | 8.49 | 7.43 | 11.68 | 99.8 | 17 | 11.1
2007-06-03 | 12:46 | 8.74 | 7.43 | 11.61 | 99.8 | 17.1 | 11.1
2007-06-03 | 13:00 | 9.11 | 7.44 | 11.51 | 99.8 | 17.1 | 11.1
2007-06-03 | 13:15 | 9.44 | 7.44 | 11.41 | 99.7 | 17.1 | 11.1
2007-06-03 | 13:30 | 9.63 | 7.44 | 11.36 | 99.7 | 17.1 | 11.1
2007-06-03 | 13:45 | 9.85 | 7.44 | 11.3 | 99.8 | 17.2 | 11.2
2007-06-03 | 14:00 | 10.05 | 7.45 | 11.25 | 99.7 | 17.2 | 11.2
2007-06-03 | 14:15 | 10.4 | 7.45 | 11.16 | 99.8 | 17.2 | 11.2
2007-06-03 | 14:30 | 10.66 | 7.45 | 11.08 | 99.7 | 17.2 | 11.2
2007-06-03 | 14:45 | 10.77 | 7.46 | 11.04 | 99.6 | 17.2 | 11.2
2007-06-03 | 15:00 | 11.01 | 7.46 | 11 | 99.7 | 17.2 | 11.2
2007-06-03 | 15:15 | 11.23 | 7.46 | 10.94 | 99.7 | 17.2 | 11.2
2007-06-03 | 15:30 | 11.51 | 7.46 | 10.86 | 99.6 | 17.3 | 11.2
2007-06-03 | 15:45 | 11.68 | 7.46 | 10.8 | 99.5 | 17.3 | 11.2
2007-06-03 | 16:00 | 11.82 | 7.46 | 10.76 | 99.5 | 17.3 | 11.2
2007-06-03 | 16:15 | 11.8 | 7.48 | 10.76 | 99.4 | 17.3 | 11.3
2007-06-03 | 16:30 | 11.7 | 7.47 | 10.81 | 99.6 | 17.3 | 11.2
2007-06-03 | 16:45 | 11.56 | 7.47 | 10.83 | 99.4 | 17.4 | 11.3
2007-06-03 | 17:00 | 11.51 | 7.46 | 10.84 | 99.5 | 17.5 | 11.3
2007-06-03 | 17:15 | 11.49 | 7.47 | 10.85 | 99.5 | 17.4 | 11.3
2007-06-03 | 17:30 | 11.46 | 7.47 | 10.85 | 99.4 | 17.4 | 11.3
2007-06-03 | 17:45 | 11.38 | 7.46 | 10.87 | 99.4 | 17.3 | 11.3
2007-06-03 | 18:00 | 11.32 | 7.46 | 10.89 | 99.5 | 17.3 | 11.2
2007-06-03 | 18:15 | 11.23 | 7.45 | 10.9 | 99.4 | 17.1 | 11.1
2007-06-03 | 18:30 | 11.11 | 7.44 | 10.93 | 99.4 | 16.9 | 11
2007-06-03 | 18:45 | 10.97 | 7.45 | 10.97 | 99.4 | 16.9 | 11
2007-06-03 | 19:00 | 10.83 | 7.42 | 11.01 | 99.4 | 17 | 11.1
2007-06-03 | 19:15 | 10.71 | 7.43 | 11.05 | 99.5 | 16.6 | 10.8
2007-06-03 | 19:30 | 10.57 | 7.42 | 11.09 | 99.5 | 16.4 | 10.6
2007-06-03 | 19:45 | 10.42 | 7.42 | 11.12 | 99.5 | 16.2 | 10.5
2007-06-03 | 20:00 | 10.29 | 7.42 | 11.15 | 99.4 | 16 | 10.4
2007-06-03 | 20:15 | 10.14 | 7.41 | 11.18 | 99.4 | 16 | 10.4
2007-06-03 | 20:30 | 9.99 | 7.41 | 11.23 | 99.5 | 15.9 | 10.3
2007-06-03 | 20:45 | 9.82 | 7.4 | 11.27 | 99.4 | 16 | 10.4
2007-06-03 | 21:00 | 9.65 | 7.38 | 11.33 | 99.5 | 15.8 | 10.3
2007-06-03 | 21:15 | 9.48 | 7.38 | 11.37 | 99.5 | 15.7 | 10.2
2007-06-03 | 21:30 | 9.33 | 7.37 | 11.41 | 99.5 | 15.6 | 10.1
2007-06-03 | 21:45 | 9.19 | 7.39 | 11.45 | 99.5 | 16.1 | 10.4
2007-06-03 | 22:00 | 9.06 | 7.37 | 11.49 | 99.5 | 16.3 | 10.6
2007-06-03 | 22:15 | 8.96 | 7.36 | 11.51 | 99.5 | 16.3 | 10.6
2007-06-03 | 22:30 | 8.85 | 7.33 | 11.55 | 99.6 | 16.3 | 10.6
2007-06-03 | 22:45 | 8.76 | 7.34 | 11.57 | 99.5 | 16.2 | 10.5
2007-06-03 | 23:00 | 8.66 | 7.34 | 11.6 | 99.5 | 15.9 | 10.3
2007-06-03 | 23:15 | 8.55 | 7.33 | 11.65 | 99.6 | 15.6 | 10.1
2007-06-03 | 23:30 | 8.43 | 7.33 | 11.63 | 99.2 | 15.7 | 10.2
2007-06-03 | 23:45 | 8.32 | 7.32 | 11.69 | 99.5 | 15.9 | 10.3
2007-06-04 | 0:00 | 8.2 | 7.31 | 11.73 | 99.5 | 15.9 | 10.3
2007-06-04 | 0:15 | 8.09 | 7.31 | 11.77 | 99.6 | 16 | 10.4
2007-06-04 | 0:30 | 7.96 | 7.31 | 11.8 | 99.6 | 15.9 | 10.4
2007-06-04 | 0:45 | 7.83 | 7.29 | 11.83 | 99.5 | 15.8 | 10.3
2007-06-04 | 1:01 | 7.71 | 7.29 | 11.86 | 99.4 | 15.9 | 10.3
2007-06-04 | 1:16 | 7.59 | 7.29 | 11.9 | 99.4 | 15.9 | 10.3
2007-06-04 | 1:31 | 7.48 | 7.28 | 11.93 | 99.4 | 15.8 | 10.3
2007-06-04 | 1:46 | 7.38 | 7.28 | 11.97 | 99.5 | 15.6 | 10.1
2007-06-04 | 2:01 | 7.29 | 7.27 | 12 | 99.6 | 15.7 | 10.2
2007-06-04 | 2:16 | 7.21 | 7.28 | 12.02 | 99.5 | 15.4 | 10
2007-06-04 | 2:31 | 7.13 | 7.27 | 12.03 | 99.4 | 15.3 | 10
2007-06-04 | 2:46 | 7.05 | 7.28 | 12.07 | 99.5 | 15.3 | 10
2007-06-04 | 3:01 | 6.98 | 7.27 | 12.08 | 99.4 | 15.5 | 10.1
2007-06-04 | 3:16 | 6.91 | 7.26 | 12.11 | 99.6 | 15.6 | 10.1
2007-06-04 | 3:31 | 6.86 | 7.26 | 12.1 | 99.3 | 15.5 | 10.1
2007-06-04 | 3:46 | 6.8 | 7.25 | 12.14 | 99.5 | 15.4 | 10
2007-06-04 | 4:01 | 6.75 | 7.25 | 12.15 | 99.5 | 15.4 | 10
2007-06-04 | 4:16 | 6.7 | 7.25 | 12.18 | 99.6 | 15.5 | 10.1
2007-06-04 | 4:31 | 6.66 | 7.26 | 12.18 | 99.5 | 15.4 | 10
2007-06-04 | 4:46 | 6.61 | 7.26 | 12.19 | 99.4 | 15.4 | 10
2007-06-04 | 5:01 | 6.56 | 7.27 | 12.21 | 99.5 | 15.3 | 10
2007-06-04 | 5:16 | 6.52 | 7.27 | 12.22 | 99.5 | 15.2 | 9.9
2007-06-04 | 5:31 | 6.48 | 7.26 | 12.23 | 99.4 | 15.2 | 9.9
2007-06-04 | 5:46 | 6.44 | 7.26 | 12.25 | 99.5 | 15.2 | 9.9
2007-06-04 | 6:01 | 6.4 | 7.27 | 12.26 | 99.5 | 15.3 | 9.9
2007-06-04 | 6:16 | 6.37 | 7.27 | 12.29 | 99.6 | 15.3 | 9.9
2007-06-04 | 6:31 | 6.34 | 7.26 | 12.29 | 99.5 | 15.2 | 9.9
2007-06-04 | 6:46 | 6.31 | 7.27 | 12.29 | 99.5 | 15.2 | 9.9
2007-06-04 | 7:01 | 6.28 | 7.26 | 12.3 | 99.5 | 15.2 | 9.9
2007-06-04 | 7:16 | 6.25 | 7.28 | 12.3 | 99.5 | 15.1 | 9.8
2007-06-04 | 7:31 | 6.24 | 7.28 | 12.32 | 99.5 | 15.1 | 9.8
2007-06-04 | 7:46 | 6.25 | 7.3 | 12.3 | 99.4 | 15.1 | 9.8
2007-06-04 | 8:01 | 6.24 | 7.32 | 12.3 | 99.4 | 15.1 | 9.8
2007-06-04 | 8:16 | 6.24 | 7.32 | 12.33 | 99.6 | 15.1 | 9.8
2007-06-04 | 8:31 | 6.26 | 7.33 | 12.31 | 99.5 | 15.1 | 9.8
2007-06-04 | 8:46 | 6.29 | 7.33 | 12.3 | 99.5 | 15.1 | 9.8
2007-06-04 | 9:01 | 6.32 | 7.35 | 12.29 | 99.5 | 15.1 | 9.8
2007-06-04 | 9:16 | 6.36 | 7.34 | 12.29 | 99.6 | 15.1 | 9.8
2007-06-04 | 9:31 | 6.42 | 7.36 | 12.26 | 99.6 | 15.1 | 9.8
2007-06-04 | 9:46 | 6.48 | 7.38 | 12.25 | 99.6 | 15 | 9.7
2007-06-04 | 10:01 | 6.53 | 7.38 | 12.22 | 99.5 | 14.9 | 9.7
2007-06-04 | 10:16 | 6.64 | 7.37 | 12.19 | 99.5 | 14.9 | 9.7
2007-06-04 | 10:31 | 6.74 | 7.38 | 12.15 | 99.4 | 15 | 9.7
2007-06-04 | 10:46 | 6.86 | 7.4 | 12.12 | 99.5 | 14.9 | 9.7
2007-06-04 | 11:01 | 6.99 | 7.39 | 12.07 | 99.4 | 14.9 | 9.7
2007-06-04 | 11:16 | 7.16 | 7.4 | 12.04 | 99.6 | 14.9 | 9.7
2007-06-04 | 11:31 | 7.35 | 7.4 | 11.98 | 99.5 | 14.8 | 9.6
2007-06-04 | 11:46 | 7.57 | 7.4 | 11.92 | 99.6 | 15.4 | 10
2007-06-04 | 12:01 | 7.75 | 7.4 | 11.85 | 99.5 | 15.4 | 10
2007-06-04 | 12:16 | 7.92 | 7.41 | 11.81 | 99.5 | 15.5 | 10.1
2007-06-04 | 12:31 | 8.1 | 7.42 | 11.77 | 99.6 | 15.5 | 10.1
2007-06-04 | 12:46 | 8.33 | 7.42 | 11.7 | 99.6 | 15.4 | 10
2007-06-04 | 13:00 | 8.35 | 7.41 | 11.68 | 99.4 | 15.4 | 10
2007-06-04 | 13:15 | 8.49 | 7.41 | 11.65 | 99.6 | 15.5 | 10
2007-06-04 | 13:30 | 8.68 | 7.42 | 11.6 | 99.6 | 15.5 | 10.1
2007-06-04 | 13:45 | 8.81 | 7.42 | 11.55 | 99.5 | 15.5 | 10.1
2007-06-04 | 14:00 | 8.84 | 7.43 | 11.54 | 99.4 | 15.5 | 10.1
2007-06-04 | 14:15 | 8.96 | 7.42 | 11.5 | 99.3 | 15.5 | 10.1
2007-06-04 | 14:30 | 9.25 | 7.44 | 11.44 | 99.6 | 15.4 | 10
2007-06-04 | 14:45 | 9.52 | 7.44 | 11.39 | 99.7 | 15.4 | 10
2007-06-04 | 15:00 | 9.74 | 7.44 | 11.31 | 99.6 | 15.4 | 10
2007-06-04 | 15:15 | 9.8 | 7.44 | 11.3 | 99.6 | 15.3 | 10
2007-06-04 | 15:30 | 9.76 | 7.45 | 11.3 | 99.5 | 15.4 | 10
2007-06-04 | 15:45 | 9.74 | 7.45 | 11.29 | 99.4 | 15.4 | 10
2007-06-04 | 16:00 | 9.81 | 7.43 | 11.28 | 99.5 | 15.4 | 10
2007-06-04 | 16:15 | 9.89 | 7.43 | 11.22 | 99.1 | 15.3 | 10
2007-06-04 | 16:30 | 10.11 | 7.43 | 11.2 | 99.5 | 15.3 | 9.9
2007-06-04 | 16:45 | 10.27 | 7.44 | 11.15 | 99.4 | 15.3 | 9.9
2007-06-04 | 17:00 | 10.39 | 7.42 | 11.1 | 99.3 | 15.3 | 9.9
2007-06-04 | 17:15 | 10.53 | 7.44 | 11.08 | 99.3 | 15.4 | 10
2007-06-04 | 17:30 | 10.69 | 7.42 | 11.03 | 99.3 | 15.2 | 9.9
2007-06-04 | 17:45 | 10.83 | 7.43 | 10.95 | 98.9 | 15.3 | 10
2007-06-04 | 18:00 | 10.92 | 7.43 | 10.97 | 99.3 | 15.3 | 10
2007-06-04 | 18:15 | 10.99 | 7.42 | 10.94 | 99.2 | 16.5 | 10.7
2007-06-04 | 18:30 | 10.97 | 7.42 | 10.94 | 99.2 | 16.5 | 10.7
2007-06-04 | 18:45 | 10.91 | 7.42 | 10.96 | 99.2 | 16.5 | 10.7
2007-06-04 | 19:00 | 10.81 | 7.42 | 10.98 | 99.1 | 16.5 | 10.8
2007-06-04 | 19:15 | 10.66 | 7.42 | 11.03 | 99.3 | 16.7 | 10.8
2007-06-04 | 19:30 | 10.52 | 7.39 | 11.08 | 99.3 | 16.5 | 10.7
2007-06-04 | 19:45 | 10.38 | 7.4 | 11.11 | 99.3 | 16.5 | 10.7
2007-06-04 | 20:00 | 10.22 | 7.39 | 11.16 | 99.3 | 16.4 | 10.7
2007-06-04 | 20:15 | 10.02 | 7.39 | 11.2 | 99.3 | 16.4 | 10.7
2007-06-04 | 20:30 | 9.82 | 7.38 | 11.27 | 99.4 | 16.3 | 10.6
2007-06-04 | 20:45 | 9.64 | 7.38 | 11.31 | 99.4 | 16.2 | 10.5
2007-06-04 | 21:00 | 9.46 | 7.38 | 11.36 | 99.4 | 16.2 | 10.5
2007-06-04 | 21:15 | 9.28 | 7.38 | 11.39 | 99.2 | 16.1 | 10.5
2007-06-04 | 21:30 | 9.11 | 7.36 | 11.45 | 99.3 | 16 | 10.4
2007-06-04 | 21:45 | 8.95 | 7.36 | 11.5 | 99.3 | 15.9 | 10.3
2007-06-04 | 22:00 | 8.8 | 7.34 | 11.55 | 99.4 | 15.8 | 10.2
2007-06-04 | 22:15 | 8.64 | 7.35 | 11.55 | 99 | 15.6 | 10.1
2007-06-04 | 22:30 | 8.46 | 7.35 | 11.64 | 99.4 | 15.5 | 10.1
2007-06-04 | 22:45 | 8.27 | 7.35 | 11.72 | 99.6 | 16.3 | 10.6
2007-06-04 | 23:00 | 8.06 | 7.33 | 11.74 | 99.3 | 16.1 | 10.5
2007-06-04 | 23:15 | 7.87 | 7.33 | 11.77 | 99.1 | 16.1 | 10.5
2007-06-04 | 23:30 | 7.7 | 7.33 | 11.86 | 99.4 | 16 | 10.4
2007-06-04 | 23:45 | 7.53 | 7.32 | 11.9 | 99.4 | 15.9 | 10.4
2007-06-05 | 0:00 | 7.38 | 7.32 | 11.95 | 99.4 | 15.6 | 10.1
2007-06-05 | 0:15 | 7.23 | 7.32 | 12 | 99.4 | 15.9 | 10.3
2007-06-05 | 0:30 | 7.11 | 7.32 | 12.04 | 99.4 | 15.6 | 10.2
2007-06-05 | 0:45 | 7 | 7.32 | 12.06 | 99.3 | 15.5 | 10.1
2007-06-05 | 1:01 | 6.9 | 7.31 | 12.08 | 99.2 | 15.4 | 10
2007-06-05 | 1:16 | 6.8 | 7.31 | 12.12 | 99.3 | 15.2 | 9.9
2007-06-05 | 1:31 | 6.71 | 7.31 | 12.13 | 99.2 | 15.7 | 10.2
2007-06-05 | 1:46 | 6.62 | 7.31 | 12.15 | 99.2 | 15.4 | 10
2007-06-05 | 2:01 | 6.52 | 7.31 | 12.18 | 99.2 | 15.3 | 9.9
2007-06-05 | 2:16 | 6.41 | 7.3 | 12.21 | 99.1 | 15 | 9.7
2007-06-05 | 2:31 | 6.31 | 7.3 | 12.24 | 99.1 | 15.5 | 10
2007-06-05 | 2:46 | 6.21 | 7.31 | 12.27 | 99.1 | 15.3 | 10
2007-06-05 | 3:01 | 6.13 | 7.31 | 12.29 | 99.1 | 15.2 | 9.9
2007-06-05 | 3:16 | 6.05 | 7.3 | 12.32 | 99.1 | 15.4 | 10
2007-06-05 | 3:31 | 5.99 | 7.3 | 12.34 | 99.1 | 15.3 | 9.9
2007-06-05 | 3:46 | 5.93 | 7.3 | 12.37 | 99.2 | 15.3 | 9.9
2007-06-05 | 4:01 | 5.88 | 7.31 | 12.37 | 99.1 | 15.1 | 9.8
2007-06-05 | 4:16 | 5.83 | 7.31 | 12.39 | 99.1 | 15.1 | 9.8
2007-06-05 | 4:31 | 5.77 | 7.31 | 12.4 | 99 | 15.1 | 9.8
2007-06-05 | 4:46 | 5.71 | 7.32 | 12.42 | 99 | 14.9 | 9.7
2007-06-05 | 5:01 | 5.64 | 7.31 | 12.43 | 99 | 15.5 | 10.1
2007-06-05 | 5:16 | 5.58 | 7.31 | 12.44 | 98.9 | 15.2 | 9.9
2007-06-05 | 5:31 | 5.52 | 7.31 | 12.48 | 99.1 | 15.1 | 9.8
2007-06-05 | 5:46 | 5.47 | 7.32 | 12.49 | 99 | 15.5 | 10.1
2007-06-05 | 6:01 | 5.42 | 7.31 | 12.5 | 99 | 15.3 | 10
2007-06-05 | 6:16 | 5.38 | 7.32 | 12.48 | 98.7 | 15.2 | 9.9
2007-06-05 | 6:31 | 5.34 | 7.33 | 12.55 | 99.1 | 15.6 | 10.1
2007-06-05 | 6:46 | 5.31 | 7.32 | 12.55 | 99.1 | 15.5 | 10.1
2007-06-05 | 7:01 | 5.27 | 7.32 | 12.57 | 99.1 | 15.5 | 10.1
2007-06-05 | 7:16 | 5.24 | 7.32 | 12.58 | 99.1 | 15.4 | 10
2007-06-05 | 7:31 | 5.2 | 7.31 | 12.57 | 99 | 15.4 | 10
2007-06-05 | 7:46 | 5.16 | 7.33 | 12.6 | 99 | 15.8 | 10.3
2007-06-05 | 8:01 | 5.13 | 7.32 | 12.61 | 99 | 15.9 | 10.3
2007-06-05 | 8:16 | 5.1 | 7.33 | 12.63 | 99.1 | 15.8 | 10.3
2007-06-05 | 8:31 | 5.1 | 7.32 | 12.63 | 99.1 | 15.8 | 10.3
2007-06-05 | 8:46 | 5.11 | 7.34 | 12.63 | 99.1 | 15.7 | 10.2
2007-06-05 | 9:01 | 5.1 | 7.34 | 12.63 | 99.1 | 16.1 | 10.4
2007-06-05 | 9:16 | 5.1 | 7.34 | 12.62 | 99 | 16.1 | 10.5
2007-06-05 | 9:31 | 5.1 | 7.34 | 12.63 | 99.1 | 16.1 | 10.5
2007-06-05 | 9:46 | 5.11 | 7.35 | 12.64 | 99.2 | 16.2 | 10.5
2007-06-05 | 10:01 | 5.12 | 7.35 | 12.63 | 99.2 | 16.2 | 10.6
2007-06-05 | 10:16 | 5.13 | 7.35 | 12.64 | 99.3 | 16.2 | 10.5
2007-06-05 | 10:31 | 5.14 | 7.35 | 12.62 | 99.1 | 16.2 | 10.5
2007-06-05 | 10:46 | 5.15 | 7.35 | 12.61 | 99.1 | 16.2 | 10.5
2007-06-05 | 11:01 | 5.22 | 7.35 | 12.61 | 99.3 | 16.2 | 10.5
2007-06-05 | 11:16 | 5.27 | 7.36 | 12.6 | 99.3 | 16.2 | 10.5
2007-06-05 | 11:31 | 5.36 | 7.36 | 12.57 | 99.3 | 16.3 | 10.6
2007-06-05 | 11:46 | 5.49 | 7.36 | 12.52 | 99.3 | 16.3 | 10.6
2007-06-05 | 12:01 | 5.61 | 7.37 | 12.48 | 99.3 | 16.4 | 10.7
2007-06-05 | 12:16 | 5.67 | 7.36 | 12.46 | 99.3 | 16.4 | 10.7
2007-06-05 | 12:31 | 5.76 | 7.38 | 12.42 | 99.2 | 16.4 | 10.6
2007-06-05 | 12:46 | 5.87 | 7.38 | 12.39 | 99.2 | 16.4 | 10.6
2007-06-05 | 13:00 | 6.05 | 7.39 | 12.36 | 99.4 | 16.4 | 10.7
2007-06-05 | 13:15 | 6.26 | 7.38 | 12.27 | 99.2 | 16.4 | 10.6
2007-06-05 | 13:30 | 6.49 | 7.4 | 12.22 | 99.4 | 16.4 | 10.7
2007-06-05 | 13:45 | 6.68 | 7.38 | 12.17 | 99.4 | 16.4 | 10.7
2007-06-05 | 14:00 | 6.78 | 7.4 | 12.14 | 99.4 | 16.4 | 10.7
2007-06-05 | 14:15 | 6.88 | 7.4 | 12.07 | 99.1 | 16.4 | 10.7
2007-06-05 | 14:30 | 6.96 | 7.4 | 12.05 | 99.2 | 16.4 | 10.6
2007-06-05 | 14:45 | 7.22 | 7.4 | 12.01 | 99.5 | 16.2 | 10.5
2007-06-05 | 15:00 | 7.36 | 7.4 | 11.96 | 99.4 | 16.2 | 10.5
2007-06-05 | 15:15 | 7.49 | 7.4 | 11.94 | 99.6 | 16.2 | 10.5
2007-06-05 | 15:30 | 7.48 | 7.41 | 11.92 | 99.3 | 16.2 | 10.5
2007-06-05 | 15:45 | 7.48 | 7.4 | 11.9 | 99.2 | 16.1 | 10.5
2007-06-05 | 16:00 | 7.66 | 7.4 | 11.85 | 99.2 | 16.1 | 10.4
2007-06-05 | 16:15 | 7.94 | 7.42 | 11.78 | 99.3 | 16 | 10.4
2007-06-05 | 16:30 | 8.12 | 7.42 | 11.73 | 99.3 | 16 | 10.4
2007-06-05 | 16:45 | 8.1 | 7.41 | 11.73 | 99.2 | 16 | 10.4
2007-06-05 | 17:00 | 7.9 | 7.42 | 11.78 | 99.2 | 16 | 10.4
2007-06-05 | 17:15 | 7.8 | 7.42 | 11.81 | 99.2 | 16 | 10.4
2007-06-05 | 17:30 | 7.79 | 7.41 | 11.79 | 99 | 16.9 | 11
2007-06-05 | 17:45 | 7.85 | 7.41 | 11.8 | 99.2 | 16.8 | 10.9
2007-06-05 | 18:00 | 7.77 | 7.4 | 11.8 | 99.1 | 16.8 | 10.9
2007-06-05 | 18:15 | 7.71 | 7.4 | 11.82 | 99.1 | 16.8 | 10.9
2007-06-05 | 18:30 | 7.66 | 7.41 | 11.85 | 99.2 | 16.8 | 10.9
2007-06-05 | 18:45 | 7.6 | 7.4 | 11.86 | 99.1 | 16.7 | 10.9
2007-06-05 | 19:00 | 7.56 | 7.41 | 11.87 | 99.2 | 16.8 | 10.9
2007-06-05 | 19:15 | 7.47 | 7.41 | 11.89 | 99.1 | 16.8 | 10.9
2007-06-05 | 19:30 | 7.37 | 7.41 | 11.92 | 99.1 | 16.7 | 10.9
2007-06-05 | 19:45 | 7.27 | 7.39 | 11.95 | 99 | 16.7 | 10.8
2007-06-05 | 20:00 | 7.15 | 7.4 | 11.99 | 99.1 | 16.7 | 10.8
2007-06-05 | 20:15 | 7 | 7.4 | 12.04 | 99.2 | 16.7 | 10.9
2007-06-05 | 20:30 | 6.82 | 7.39 | 12.1 | 99.2 | 16.6 | 10.8
2007-06-05 | 20:45 | 6.65 | 7.39 | 12.13 | 99.1 | 16.6 | 10.8
2007-06-05 | 21:00 | 6.51 | 7.39 | 12.19 | 99.2 | 16.6 | 10.8
2007-06-05 | 21:15 | 6.4 | 7.39 | 12.21 | 99.1 | 16.5 | 10.7
2007-06-05 | 21:30 | 6.3 | 7.39 | 12.24 | 99.1 | 16.5 | 10.7
2007-06-05 | 21:45 | 6.18 | 7.39 | 12.28 | 99.1 | 16.5 | 10.7
2007-06-05 | 22:00 | 6.07 | 7.38 | 12.32 | 99.1 | 16.6 | 10.8
2007-06-05 | 22:15 | 5.96 | 7.39 | 12.35 | 99.1 | 16.5 | 10.7
2007-06-05 | 22:30 | 5.85 | 7.38 | 12.38 | 99.1 | 16.4 | 10.7
2007-06-05 | 22:45 | 5.74 | 7.39 | 12.42 | 99.1 | 16.5 | 10.7
2007-06-05 | 23:00 | 5.64 | 7.38 | 12.44 | 99 | 16.5 | 10.7
2007-06-05 | 23:15 | 5.53 | 7.38 | 12.48 | 99 | 16.4 | 10.7
2007-06-05 | 23:30 | 5.42 | 7.38 | 12.53 | 99.2 | 16.5 | 10.7
2007-06-05 | 23:45 | 5.32 | 7.38 | 12.55 | 99 | 16.5 | 10.7
2007-06-06 | 0:00 | 5.21 | 7.37 | 12.59 | 99.1 | 16.5 | 10.7
2007-06-06 | 0:15 | 5.09 | 7.38 | 12.62 | 99.1 | 16.4 | 10.7
2007-06-06 | 0:30 | 4.98 | 7.38 | 12.66 | 99.1 | 17.1 | 11.1
2007-06-06 | 0:45 | 4.87 | 7.38 | 12.7 | 99.1 | 17.2 | 11.2
2007-06-06 | 1:01 | 4.76 | 7.37 | 12.73 | 99 | 17.2 | 11.2
2007-06-06 | 1:16 | 4.65 | 7.38 | 12.77 | 99.1 | 17.2 | 11.2
2007-06-06 | 1:31 | 4.56 | 7.38 | 12.81 | 99.2 | 17.2 | 11.2
2007-06-06 | 1:46 | 4.47 | 7.38 | 12.83 | 99.1 | 17.2 | 11.2
2007-06-06 | 2:01 | 4.39 | 7.37 | 12.87 | 99.1 | 17.4 | 11.3
2007-06-06 | 2:16 | 4.3 | 7.37 | 12.89 | 99.1 | 17.4 | 11.3
2007-06-06 | 2:31 | 4.22 | 7.37 | 12.92 | 99.1 | 17.3 | 11.3
2007-06-06 | 2:46 | 4.15 | 7.38 | 12.94 | 99.1 | 17.3 | 11.3
2007-06-06 | 3:01 | 4.08 | 7.37 | 12.97 | 99.2 | 17.4 | 11.3
2007-06-06 | 3:16 | 4.02 | 7.38 | 12.98 | 99.1 | 17.4 | 11.3
2007-06-06 | 3:31 | 3.96 | 7.38 | 13 | 99.1 | 17.5 | 11.4
2007-06-06 | 3:46 | 3.9 | 7.38 | 13.03 | 99.1 | 17.5 | 11.4
2007-06-06 | 4:01 | 3.83 | 7.38 | 13.05 | 99.1 | 17.5 | 11.4
2007-06-06 | 4:16 | 3.77 | 7.38 | 13.08 | 99.1 | 17.5 | 11.4
2007-06-06 | 4:31 | 3.7 | 7.37 | 13.1 | 99.1 | 17.5 | 11.4
2007-06-06 | 4:46 | 3.65 | 7.37 | 13.12 | 99.1 | 17.5 | 11.4
2007-06-06 | 5:01 | 3.59 | 7.38 | 13.14 | 99.1 | 17.6 | 11.4
2007-06-06 | 5:16 | 3.53 | 7.38 | 13.16 | 99.1 | 17.6 | 11.4
2007-06-06 | 5:31 | 3.48 | 7.37 | 13.17 | 99.1 | 17.5 | 11.4
2007-06-06 | 5:46 | 3.43 | 7.37 | 13.2 | 99.2 | 17.5 | 11.4
2007-06-06 | 6:01 | 3.38 | 7.37 | 13.22 | 99.2 | 17.5 | 11.4
2007-06-06 | 6:16 | 3.34 | 7.37 | 13.23 | 99.1 | 17.5 | 11.4
2007-06-06 | 6:31 | 3.3 | 7.37 | 13.25 | 99.2 | 17.5 | 11.4
2007-06-06 | 6:46 | 3.27 | 7.38 | 13.28 | 99.4 | 17.5 | 11.4
2007-06-06 | 7:01 | 3.24 | 7.38 | 13.28 | 99.3 | 17.5 | 11.4
2007-06-06 | 7:16 | 3.23 | 7.38 | 13.29 | 99.3 | 17.6 | 11.4
2007-06-06 | 7:31 | 3.22 | 7.38 | 13.3 | 99.4 | 17.6 | 11.4
2007-06-06 | 7:46 | 3.21 | 7.38 | 13.3 | 99.3 | 17.6 | 11.4
2007-06-06 | 8:01 | 3.22 | 7.39 | 13.3 | 99.4 | 17.6 | 11.5
2007-06-06 | 8:16 | 3.22 | 7.39 | 13.3 | 99.4 | 17.6 | 11.5
2007-06-06 | 8:31 | 3.24 | 7.39 | 13.29 | 99.4 | 17.6 | 11.5
2007-06-06 | 8:46 | 3.26 | 7.39 | 13.29 | 99.4 | 17.6 | 11.5
2007-06-06 | 9:01 | 3.31 | 7.39 | 13.27 | 99.4 | 17.7 | 11.5
2007-06-06 | 9:16 | 3.37 | 7.4 | 13.26 | 99.5 | 17.6 | 11.4
2007-06-06 | 9:31 | 3.45 | 7.4 | 13.22 | 99.4 | 17.6 | 11.4
2007-06-06 | 9:46 | 3.54 | 7.4 | 13.2 | 99.5 | 17.6 | 11.5
2007-06-06 | 10:01 | 3.63 | 7.4 | 13.17 | 99.5 | 17.6 | 11.4
2007-06-06 | 10:16 | 3.74 | 7.4 | 13.13 | 99.5 | 17.6 | 11.4
2007-06-06 | 10:31 | 3.87 | 7.42 | 13.09 | 99.5 | 17.6 | 11.4
2007-06-06 | 10:46 | 4 | 7.41 | 13.04 | 99.5 | 17.6 | 11.4
2007-06-06 | 11:01 | 4.17 | 7.42 | 12.99 | 99.5 | 17.5 | 11.4
2007-06-06 | 11:16 | 4.35 | 7.42 | 12.92 | 99.5 | 17.5 | 11.4
2007-06-06 | 11:31 | 4.54 | 7.42 | 12.88 | 99.7 | 17.5 | 11.4
2007-06-06 | 11:46 | 4.74 | 7.42 | 12.83 | 99.7 | 17.4 | 11.3
2007-06-06 | 12:01 | 4.96 | 7.42 | 12.76 | 99.8 | 17.5 | 11.4
2007-06-06 | 12:16 | 5.2 | 7.43 | 12.68 | 99.8 | 17.5 | 11.4
2007-06-06 | 12:31 | 5.45 | 7.42 | 12.6 | 99.8 | 17.5 | 11.4
2007-06-06 | 12:46 | 5.71 | 7.44 | 12.55 | 100 | 17.4 | 11.3
2007-06-06 | 13:00 | 5.98 | 7.44 | 12.44 | 99.9 | 17.5 | 11.4
2007-06-06 | 13:15 | 6.26 | 7.44 | 12.34 | 99.8 | 17.5 | 11.4
2007-06-06 | 13:30 | 6.54 | 7.44 | 12.26 | 99.9 | 17.5 | 11.4
2007-06-06 | 13:45 | 6.81 | 7.44 | 12.19 | 99.9 | 17.5 | 11.4
2007-06-06 | 14:00 | 7.06 | 7.45 | 12.13 | 100 | 17.6 | 11.4
2007-06-06 | 14:15 | 7.33 | 7.45 | 12.04 | 100 | 17.7 | 11.5
2007-06-06 | 14:30 | 7.59 | 7.45 | 11.96 | 99.9 | 17.7 | 11.5
2007-06-06 | 14:45 | 7.83 | 7.45 | 11.89 | 100 | 17.8 | 11.5
2007-06-06 | 15:00 | 8.07 | 7.46 | 11.82 | 99.9 | 17.8 | 11.6
2007-06-06 | 15:15 | 8.3 | 7.46 | 11.76 | 100 | 17.8 | 11.6
2007-06-06 | 15:30 | 8.43 | 7.45 | 11.72 | 100 | 17.8 | 11.6
2007-06-06 | 15:45 | 8.61 | 7.46 | 11.68 | 100.1 | 17.9 | 11.6
2007-06-06 | 16:00 | 8.83 | 7.46 | 11.6 | 99.9 | 18 | 11.7
2007-06-06 | 16:15 | 8.99 | 7.46 | 11.58 | 100.1 | 17.9 | 11.6
2007-06-06 | 16:30 | 9.04 | 7.46 | 11.54 | 99.9 | 17.8 | 11.6
2007-06-06 | 16:45 | 9.11 | 7.47 | 11.52 | 99.9 | 17.8 | 11.6
2007-06-06 | 17:00 | 9.19 | 7.47 | 11.5 | 100 | 17.8 | 11.6
2007-06-06 | 17:15 | 9.21 | 7.46 | 11.5 | 100 | 17.8 | 11.6
2007-06-06 | 17:30 | 9.15 | 7.47 | 11.51 | 99.9 | 17.8 | 11.6
2007-06-06 | 17:45 | 9.14 | 7.47 | 11.51 | 99.9 | 17.8 | 11.6
2007-06-06 | 18:00 | 9.12 | 7.48 | 11.51 | 99.8 | 17.8 | 11.5
2007-06-06 | 18:15 | 9.04 | 7.46 | 11.5 | 99.6 | 17.7 | 11.5
2007-06-06 | 18:30 | 8.91 | 7.46 | 11.54 | 99.6 | 17.7 | 11.5
2007-06-06 | 18:45 | 8.78 | 7.46 | 11.58 | 99.6 | 17.7 | 11.5
2007-06-06 | 19:00 | 8.62 | 7.46 | 11.63 | 99.6 | 17.7 | 11.5
2007-06-06 | 19:15 | 8.45 | 7.46 | 11.68 | 99.7 | 17.7 | 11.5
2007-06-06 | 19:30 | 8.26 | 7.46 | 11.72 | 99.5 | 17.6 | 11.5
2007-06-06 | 19:45 | 8.08 | 7.46 | 11.78 | 99.6 | 17.6 | 11.4
2007-06-06 | 20:00 | 7.94 | 7.46 | 11.82 | 99.7 | 17.6 | 11.4
2007-06-06 | 20:15 | 7.79 | 7.45 | 11.86 | 99.6 | 17.6 | 11.4
2007-06-06 | 20:30 | 7.62 | 7.45 | 11.9 | 99.6 | 17.6 | 11.4
2007-06-06 | 20:45 | 7.45 | 7.45 | 11.95 | 99.5 | 17.5 | 11.4
2007-06-06 | 21:00 | 7.27 | 7.44 | 12.01 | 99.6 | 17.6 | 11.4
2007-06-06 | 21:15 | 7.09 | 7.44 | 12.06 | 99.6 | 17.5 | 11.4
2007-06-06 | 21:30 | 6.92 | 7.43 | 12.12 | 99.6 | 17.5 | 11.4
2007-06-06 | 21:45 | 6.75 | 7.43 | 12.17 | 99.6 | 17.4 | 11.3
2007-06-06 | 22:00 | 6.59 | 7.43 | 12.21 | 99.6 | 17.4 | 11.3
2007-06-06 | 22:15 | 6.43 | 7.44 | 12.26 | 99.5 | 17.4 | 11.3
2007-06-06 | 22:30 | 6.27 | 7.43 | 12.33 | 99.7 | 17.4 | 11.3
2007-06-06 | 22:45 | 6.12 | 7.43 | 12.36 | 99.6 | 17.4 | 11.3
2007-06-06 | 23:00 | 5.98 | 7.43 | 12.41 | 99.6 | 17.4 | 11.3
2007-06-06 | 23:15 | 5.85 | 7.42 | 12.45 | 99.7 | 17.3 | 11.3
2007-06-06 | 23:30 | 5.73 | 7.43 | 12.49 | 99.7 | 17.4 | 11.3
2007-06-06 | 23:45 | 5.62 | 7.42 | 12.51 | 99.5 | 17.4 | 11.3
2007-06-07 | 0:00 | 5.5 | 7.42 | 12.55 | 99.6 | 17.4 | 11.3
2007-06-07 | 0:15 | 5.39 | 7.42 | 12.59 | 99.6 | 17.4 | 11.3
2007-06-07 | 0:30 | 5.28 | 7.42 | 12.64 | 99.7 | 17.4 | 11.3
2007-06-07 | 0:45 | 5.18 | 7.42 | 12.66 | 99.6 | 17.4 | 11.3
2007-06-07 | 1:01 | 5.08 | 7.41 | 12.71 | 99.7 | 17.4 | 11.3
2007-06-07 | 1:16 | 4.98 | 7.41 | 12.75 | 99.8 | 17.4 | 11.3
2007-06-07 | 1:31 | 4.89 | 7.41 | 12.76 | 99.6 | 17.4 | 11.3
2007-06-07 | 1:46 | 4.8 | 7.42 | 12.79 | 99.7 | 17.4 | 11.3
2007-06-07 | 2:01 | 4.71 | 7.41 | 12.82 | 99.7 | 17.4 | 11.3
2007-06-07 | 2:16 | 4.62 | 7.41 | 12.85 | 99.6 | 17.4 | 11.3
2007-06-07 | 2:31 | 4.52 | 7.41 | 12.87 | 99.6 | 17.4 | 11.3
2007-06-07 | 2:46 | 4.42 | 7.4 | 12.91 | 99.6 | 17.5 | 11.3
2007-06-07 | 3:01 | 4.32 | 7.4 | 12.92 | 99.4 | 17.5 | 11.4
2007-06-07 | 3:16 | 4.22 | 7.4 | 12.99 | 99.7 | 17.5 | 11.4
2007-06-07 | 3:31 | 4.12 | 7.4 | 13.02 | 99.7 | 17.5 | 11.4
2007-06-07 | 3:46 | 4.03 | 7.39 | 13.06 | 99.7 | 17.5 | 11.4
2007-06-07 | 4:01 | 3.95 | 7.39 | 13.06 | 99.5 | 17.5 | 11.4
2007-06-07 | 4:16 | 3.88 | 7.38 | 13.11 | 99.7 | 17.6 | 11.4
2007-06-07 | 4:31 | 3.82 | 7.39 | 13.14 | 99.7 | 17.6 | 11.4
2007-06-07 | 4:46 | 3.75 | 7.38 | 13.15 | 99.7 | 17.6 | 11.5
2007-06-07 | 5:01 | 3.67 | 7.37 | 13.19 | 99.7 | 17.7 | 11.5
2007-06-07 | 5:16 | 3.59 | 7.38 | 13.22 | 99.8 | 17.7 | 11.5
2007-06-07 | 5:31 | 3.52 | 7.38 | 13.25 | 99.8 | 17.7 | 11.5
2007-06-07 | 5:46 | 3.45 | 7.37 | 13.27 | 99.8 | 17.7 | 11.5
2007-06-07 | 6:01 | 3.4 | 7.38 | 13.24 | 99.4 | 17.7 | 11.5
2007-06-07 | 6:16 | 3.36 | 7.39 | 13.32 | 99.9 | 17.7 | 11.5
2007-06-07 | 6:31 | 3.33 | 7.38 | 13.32 | 99.8 | 17.7 | 11.5
2007-06-07 | 6:46 | 3.32 | 7.39 | 13.33 | 99.9 | 17.8 | 11.5
2007-06-07 | 7:01 | 3.31 | 7.39 | 13.35 | 100 | 17.8 | 11.5
2007-06-07 | 7:16 | 3.32 | 7.4 | 13.33 | 99.9 | 17.8 | 11.6
2007-06-07 | 7:31 | 3.34 | 7.41 | 13.34 | 100 | 17.7 | 11.5
2007-06-07 | 7:46 | 3.36 | 7.42 | 13.32 | 99.9 | 17.7 | 11.5
2007-06-07 | 8:01 | 3.39 | 7.42 | 13.32 | 100 | 17.7 | 11.5
2007-06-07 | 8:16 | 3.42 | 7.42 | 13.31 | 100 | 17.8 | 11.6
2007-06-07 | 8:31 | 3.48 | 7.42 | 13.3 | 100.1 | 17.8 | 11.5
2007-06-07 | 8:46 | 3.54 | 7.42 | 13.27 | 100 | 17.8 | 11.6
2007-06-07 | 9:01 | 3.61 | 7.43 | 13.24 | 100 | 17.8 | 11.6
2007-06-07 | 9:16 | 3.7 | 7.43 | 13.21 | 100 | 17.8 | 11.6
2007-06-07 | 9:31 | 3.79 | 7.42 | 13.19 | 100.1 | 17.8 | 11.6
2007-06-07 | 9:46 | 3.89 | 7.44 | 13.15 | 100 | 17.8 | 11.6
2007-06-07 | 10:01 | 3.98 | 7.43 | 13.11 | 99.9 | 17.8 | 11.6
2007-06-07 | 10:16 | 4.12 | 7.44 | 13.08 | 100.1 | 17.9 | 11.6
2007-06-07 | 10:31 | 4.28 | 7.44 | 13.03 | 100.1 | 17.9 | 11.6
2007-06-07 | 10:46 | 4.32 | 7.43 | 13.01 | 100.1 | 17.8 | 11.6
2007-06-07 | 11:01 | 4.34 | 7.44 | 13.01 | 100.1 | 17.9 | 11.6
2007-06-07 | 11:16 | 4.41 | 7.45 | 12.98 | 100.1 | 17.8 | 11.6
2007-06-07 | 11:31 | 4.48 | 7.45 | 12.94 | 100 | 17.8 | 11.6
2007-06-07 | 11:46 | 4.59 | 7.45 | 12.92 | 100.1 | 17.8 | 11.6
2007-06-07 | 12:01 | 4.76 | 7.45 | 12.86 | 100 | 17.9 | 11.6
2007-06-07 | 12:16 | 4.93 | 7.45 | 12.82 | 100.2 | 17.9 | 11.6
2007-06-07 | 12:31 | 5.06 | 7.45 | 12.75 | 100 | 17.9 | 11.6
2007-06-07 | 12:46 | 5.18 | 7.45 | 12.7 | 99.9 | 17.9 | 11.6
2007-06-07 | 13:00 | 5.35 | 7.46 | 12.68 | 100.2 | 18 | 11.7
2007-06-07 | 13:15 | 5.53 | 7.46 | 12.62 | 100.1 | 18 | 11.7
2007-06-07 | 13:30 | 5.63 | 7.46 | 12.57 | 100 | 18 | 11.7
2007-06-07 | 13:45 | 5.68 | 7.45 | 12.55 | 100 | 18.1 | 11.7
2007-06-07 | 14:00 | 5.73 | 7.46 | 12.55 | 100.1 | 18 | 11.7
2007-06-07 | 14:15 | 5.76 | 7.46 | 12.52 | 100 | 18 | 11.7
2007-06-07 | 14:30 | 5.81 | 7.46 | 12.51 | 100 | 18.1 | 11.7
2007-06-07 | 14:45 | 5.9 | 7.5 | 12.48 | 100 | 18.1 | 11.8