## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, City Park
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA audited in situ against NIST, COB Pro Series, and Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-05 |  | 2009-05-05 |  | A
Time | Computer set to atomic time |  | 15:18 |  | 15:18 |  | A
Temp | VWR NIST traceable surface water |  | 8.4 |  | 8.2 |  | A
DO | In Situ |  | 11.49 |  | 11.48 |  | A
SpC | In Situ |  | 37.7 |  | 36 |  | A
Ph | in Situ |  | 7.62 |  | 7.69 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-05-05 19:15 | 8.1 | 7.51 | 10.91 | 92.4 | 29.7 | 19
Sonde | 2009-05-05 19:15 | 8.13 | 7.57 | 11.28 | 95.6 | 29 | 19 | Passing with COB Pro-Series meter
Grade |  | A | A | B |  | A
Audit 2 | 2009-05-12 12:01 | 6.44 | 6.9 | 17.07 | 138.7 | 20 | 20
Sonde | 2009-05-12 12:00 | 6.23 | 7.68 | 12.98 | 104.9 | 21 | 21 | Audit with questionable 556
Grade |  | A |  |  |  | A
Audit 3
Sonde
Grade |  | A | A | A |  | #DIV/0!
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA audited in situ against NIST and Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-12 |  | 2009-05-12 |  | A
Time | Computer set to atomic time |  | 12:10 |  | 12:10 |  | A
Temp | VWR NIST traceable surface water |  | 6.8 |  | 6.6 |  | A
DO | In Situ |  | 12.88 |  | 12.9 |  | A
SpC | In Situ |  | 32 |  | 32 |  | A
Ph | In Situ |  | 7.7 |  | 7.61 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Was not working for pH and DO. Sonde data can be used based on representation of historic site data given the time of year. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, City Park
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | COB Pro Series
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-05 |  | 2009-05-05 |  | A
Time | Computer set to atomic time |  | 15:18 |  | 15:18 |  | A
Temp | VWR NIST traceable surface water |  | 8.4 |  | 8.2 |  | A
DO | Winkler titration of surface water |  | 11.48 |  | 10.95 |  | B
SpC | 148 µS/mL |  | 36 |  | 37.7 |  | A
Ph | pH 7 standard solution |  | 7.66 |  | 7.62 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-05 19:15 | 8.1 | 7.51 | 10.91 | 92.4 | 29.7 | 19
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-18 |  | 2009-05-18 |  | A
Time | Computer set to atomic time |  | 10:53 |  | 10:53 |  | A
Temp | VWR NIST traceable surface water |  | 14.16 |  | 14.2 |  | A
DO | Winkler titration of surface water |  | 9 |  | 9.51 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 152.3 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.04 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
5/18/09 Audit is from COB after we sent their meter back (both of these audits are for the COB Pro-Sereis meter) |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, City Park
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-12 |  | 2009-05-12 |  | A
Time | Computer set to atomic time |  | 12:10 |  | 12:10 |  | A
Temp | VWR NIST traceable surface water |  | 6.8 |  | 6.68 |  | A
DO | Winkler titration of surface water |  | 12.88 |  | 15.88 |  | 0
SpC | 148 µS/mL |  | 32 |  | 31 |  | A
Ph | pH 7 standard solution |  | 7.7 |  | 6.98 |  | Fail
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-12 12:01 | 6.44 | 6.9 | 17.07 | 138.7 | 20 | 20
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-15 |  | 2009-05-15 |  | A
Time | Computer set to atomic time |  | 15:21 |  | 15:21 |  | A
Temp | VWR NIST traceable surface water |  | 14.9 |  | 14.7 |  | A
DO | Winkler titration of surface water |  | 10.8 |  | 9.23 |  | Fail
SpC | 148 µS/mL |  | 61 |  | 61.3 |  | A
Ph | pH 7 standard solution |  | 8.45 |  | 8.34 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Paramters that failed cannot be trusted |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-05-05 | 19:15 | 8.13 | 19 | 29 | 11.28 | 95.6 | 7.57
2009-05-05 | 19:30 | 7.94 | 19 | 29 | 11.66 | 98.3 | 7.6
2009-05-05 | 19:45 | 7.73 | 19 | 29 | 11.7 | 98.2 | 7.6
2009-05-05 | 20:00 | 7.53 | 19 | 29 | 11.84 | 98.8 | 7.6
2009-05-05 | 20:15 | 7.32 | 19 | 29 | 11.88 | 98.6 | 7.6
2009-05-05 | 20:30 | 7.12 | 19 | 29 | 11.92 | 98.5 | 7.6
2009-05-05 | 20:45 | 6.93 | 19 | 29 | 11.95 | 98.3 | 7.59
2009-05-05 | 21:00 | 6.75 | 19 | 29 | 12 | 98.2 | 7.59
2009-05-05 | 21:15 | 6.58 | 19 | 29 | 12.06 | 98.3 | 7.6
2009-05-05 | 21:30 | 6.41 | 19 | 29 | 12.17 | 98.8 | 7.59
2009-05-05 | 21:45 | 6.24 | 19 | 29 | 12.26 | 99.1 | 7.59
2009-05-05 | 22:00 | 6.08 | 19 | 29 | 12.36 | 99.5 | 7.59
2009-05-05 | 22:15 | 5.93 | 19 | 29 | 12.35 | 99 | 7.59
2009-05-05 | 22:30 | 5.78 | 19 | 29 | 12.39 | 99 | 7.59
2009-05-05 | 22:45 | 5.64 | 19 | 29 | 12.42 | 98.9 | 7.59
2009-05-05 | 23:00 | 5.51 | 19 | 29 | 12.52 | 99.3 | 7.59
2009-05-05 | 23:15 | 5.37 | 19 | 29 | 12.53 | 99.1 | 7.59
2009-05-05 | 23:30 | 5.24 | 19 | 29 | 12.59 | 99.1 | 7.58
2009-05-05 | 23:45 | 5.11 | 19 | 29 | 12.71 | 99.8 | 7.58
2009-05-06 | 0:00 | 4.99 | 19 | 29 | 12.71 | 99.5 | 7.58
2009-05-06 | 0:15 | 4.88 | 19 | 29 | 12.77 | 99.6 | 7.57
2009-05-06 | 0:30 | 4.77 | 19 | 29 | 12.77 | 99.3 | 7.57
2009-05-06 | 0:45 | 4.66 | 19 | 29 | 12.85 | 99.7 | 7.56
2009-05-06 | 1:01 | 4.57 | 19 | 29 | 12.83 | 99.3 | 7.57
2009-05-06 | 1:16 | 4.48 | 19 | 29 | 12.86 | 99.3 | 7.57
2009-05-06 | 1:31 | 4.42 | 19 | 29 | 12.93 | 99.8 | 7.57
2009-05-06 | 1:46 | 4.36 | 19 | 29 | 12.93 | 99.5 | 7.57
2009-05-06 | 2:01 | 4.31 | 19 | 29 | 12.99 | 99.9 | 7.56
2009-05-06 | 2:16 | 4.26 | 19 | 29 | 12.97 | 99.6 | 7.57
2009-05-06 | 2:31 | 4.21 | 19 | 29 | 12.97 | 99.5 | 7.56
2009-05-06 | 2:46 | 4.16 | 19 | 29 | 12.98 | 99.4 | 7.56
2009-05-06 | 3:01 | 4.12 | 19 | 29 | 13 | 99.4 | 7.56
2009-05-06 | 3:16 | 4.08 | 19 | 29 | 13.06 | 99.8 | 7.56
2009-05-06 | 3:31 | 4.04 | 19 | 29 | 13.06 | 99.7 | 7.56
2009-05-06 | 3:46 | 4 | 19 | 29 | 13.05 | 99.6 | 7.56
2009-05-06 | 4:01 | 3.97 | 19 | 29 | 13.11 | 99.9 | 7.56
2009-05-06 | 4:16 | 3.94 | 19 | 29 | 13.08 | 99.6 | 7.57
2009-05-06 | 4:31 | 3.91 | 19 | 29 | 13.09 | 99.7 | 7.56
2009-05-06 | 4:46 | 3.89 | 19 | 29 | 13.1 | 99.6 | 7.56
2009-05-06 | 5:01 | 3.87 | 19 | 29 | 13.15 | 100 | 7.57
2009-05-06 | 5:16 | 3.84 | 19 | 29 | 13.13 | 99.7 | 7.57
2009-05-06 | 5:31 | 3.83 | 19 | 29 | 13.15 | 99.8 | 7.57
2009-05-06 | 5:46 | 3.81 | 19 | 29 | 13.19 | 100.1 | 7.57
2009-05-06 | 6:01 | 3.8 | 19 | 29 | 13.2 | 100.2 | 7.58
2009-05-06 | 6:16 | 3.79 | 19 | 29 | 13.17 | 99.9 | 7.58
2009-05-06 | 6:31 | 3.79 | 19 | 29 | 13.13 | 99.6 | 7.58
2009-05-06 | 6:46 | 3.79 | 19 | 30 | 13.19 | 100.1 | 7.58
2009-05-06 | 7:01 | 3.8 | 19 | 30 | 13.22 | 100.3 | 7.59
2009-05-06 | 7:16 | 3.81 | 19 | 30 | 13.23 | 100.4 | 7.59
2009-05-06 | 7:31 | 3.83 | 19 | 30 | 13.24 | 100.6 | 7.59
2009-05-06 | 7:46 | 3.87 | 19 | 30 | 13.18 | 100.2 | 7.59
2009-05-06 | 8:01 | 3.92 | 19 | 30 | 13.16 | 100.2 | 7.6
2009-05-06 | 8:16 | 3.97 | 19 | 30 | 13.13 | 100.1 | 7.6
2009-05-06 | 8:31 | 4.03 | 19 | 30 | 13.12 | 100.1 | 7.61
2009-05-06 | 8:46 | 4.1 | 19 | 30 | 13.08 | 100.1 | 7.61
2009-05-06 | 9:01 | 4.18 | 20 | 30 | 13.14 | 100.7 | 7.62
2009-05-06 | 9:16 | 4.26 | 20 | 30 | 13.16 | 101.1 | 7.62
2009-05-06 | 9:31 | 4.36 | 20 | 30 | 13.07 | 100.7 | 7.62
2009-05-06 | 9:46 | 4.47 | 20 | 30 | 13.09 | 101.1 | 7.62
2009-05-06 | 10:01 | 4.57 | 20 | 30 | 13.09 | 101.3 | 7.62
2009-05-06 | 10:16 | 4.68 | 20 | 30 | 13.02 | 101.1 | 7.63
2009-05-06 | 10:31 | 4.8 | 20 | 31 | 12.95 | 100.9 | 7.63
2009-05-06 | 10:46 | 4.91 | 20 | 30 | 12.91 | 100.8 | 7.63
2009-05-06 | 11:01 | 5.02 | 20 | 30 | 12.93 | 101.3 | 7.63
2009-05-06 | 11:16 | 5.14 | 20 | 30 | 12.91 | 101.5 | 7.64
2009-05-06 | 11:31 | 5.27 | 20 | 30 | 12.9 | 101.7 | 7.64
2009-05-06 | 11:46 | 5.39 | 20 | 30 | 12.8 | 101.2 | 7.64
2009-05-06 | 12:01 | 5.51 | 20 | 30 | 12.73 | 101 | 7.64
2009-05-06 | 12:16 | 5.65 | 20 | 30 | 12.68 | 101 | 7.64
2009-05-06 | 12:31 | 5.82 | 20 | 30 | 12.65 | 101.2 | 7.64
2009-05-06 | 12:46 | 6 | 20 | 30 | 12.6 | 101.2 | 7.65
2009-05-06 | 13:00 | 6.22 | 20 | 30 | 12.57 | 101.5 | 7.65
2009-05-06 | 13:15 | 6.62 | 20 | 30 | 12.52 | 102.2 | 7.65
2009-05-06 | 13:30 | 6.99 | 20 | 30 | 12.4 | 102.1 | 7.66
2009-05-06 | 13:45 | 7.37 | 20 | 30 | 12.33 | 102.5 | 7.66
2009-05-06 | 14:00 | 7.73 | 20 | 30 | 12.23 | 102.6 | 7.66
2009-05-06 | 14:15 | 8.12 | 20 | 30 | 12.06 | 102.1 | 7.66
2009-05-06 | 14:30 | 8.44 | 20 | 30 | 11.95 | 102 | 7.68
2009-05-06 | 14:45 | 8.61 | 20 | 30 | 11.93 | 102.2 | 7.68
2009-05-06 | 15:00 | 8.69 | 20 | 30 | 11.85 | 101.7 | 7.67
2009-05-06 | 15:15 | 8.6 | 20 | 30 | 11.77 | 100.9 | 7.67
2009-05-06 | 15:30 | 8.47 | 20 | 30 | 11.82 | 101 | 7.67
2009-05-06 | 15:45 | 8.48 | 20 | 30 | 11.85 | 101.3 | 7.67
2009-05-06 | 16:00 | 8.52 | 20 | 30 | 11.8 | 100.9 | 7.67
2009-05-06 | 16:15 | 8.6 | 20 | 30 | 11.76 | 100.7 | 7.67
2009-05-06 | 16:30 | 8.6 | 20 | 30 | 11.77 | 100.8 | 7.67
2009-05-06 | 16:45 | 8.52 | 20 | 30 | 11.72 | 100.2 | 7.67
2009-05-06 | 17:00 | 8.4 | 20 | 30 | 11.78 | 100.4 | 7.67
2009-05-06 | 17:15 | 8.25 | 20 | 30 | 11.78 | 100.1 | 7.67
2009-05-06 | 17:30 | 8.13 | 20 | 30 | 11.86 | 100.5 | 7.66
2009-05-06 | 17:45 | 8.01 | 20 | 30 | 11.87 | 100.2 | 7.66
2009-05-06 | 18:00 | 7.91 | 20 | 30 | 11.86 | 99.9 | 7.66
2009-05-06 | 18:15 | 7.84 | 20 | 30 | 11.93 | 100.3 | 7.66
2009-05-06 | 18:30 | 7.78 | 20 | 30 | 11.96 | 100.4 | 7.65
2009-05-06 | 18:45 | 7.7 | 20 | 30 | 11.98 | 100.4 | 7.65
2009-05-06 | 19:00 | 7.56 | 20 | 30 | 11.97 | 100 | 7.65
2009-05-06 | 19:15 | 7.41 | 20 | 30 | 12.05 | 100.3 | 7.65
2009-05-06 | 19:30 | 7.26 | 20 | 30 | 12.11 | 100.4 | 7.64
2009-05-06 | 19:45 | 7.11 | 19 | 30 | 12.11 | 100 | 7.64
2009-05-06 | 20:00 | 6.97 | 19 | 30 | 12.13 | 99.8 | 7.63
2009-05-06 | 20:15 | 6.82 | 19 | 30 | 12.17 | 99.8 | 7.64
2009-05-06 | 20:30 | 6.64 | 19 | 30 | 12.27 | 100.2 | 7.63
2009-05-06 | 20:45 | 6.46 | 19 | 30 | 12.35 | 100.4 | 7.62
2009-05-06 | 21:00 | 6.3 | 19 | 29 | 12.42 | 100.5 | 7.62
2009-05-06 | 21:15 | 6.13 | 19 | 29 | 12.46 | 100.4 | 7.61
2009-05-06 | 21:30 | 5.99 | 19 | 29 | 12.48 | 100.2 | 7.61
2009-05-06 | 21:45 | 5.84 | 19 | 29 | 12.51 | 100.1 | 7.61
2009-05-06 | 22:00 | 5.68 | 19 | 29 | 12.6 | 100.4 | 7.6
2009-05-06 | 22:15 | 5.52 | 19 | 29 | 12.65 | 100.3 | 7.6
2009-05-06 | 22:30 | 5.37 | 19 | 29 | 12.75 | 100.8 | 7.59
2009-05-06 | 22:45 | 5.23 | 19 | 29 | 12.83 | 101.1 | 7.59
2009-05-06 | 23:00 | 5.1 | 19 | 29 | 12.82 | 100.7 | 7.59
2009-05-06 | 23:15 | 4.97 | 19 | 29 | 12.93 | 101.2 | 7.59
2009-05-06 | 23:30 | 4.86 | 19 | 29 | 12.99 | 101.3 | 7.58
2009-05-06 | 23:45 | 4.73 | 19 | 29 | 13.08 | 101.7 | 7.58
2009-05-07 | 0:00 | 4.61 | 19 | 29 | 13.11 | 101.6 | 7.58
2009-05-07 | 0:15 | 4.49 | 19 | 29 | 13.1 | 101.2 | 7.58
2009-05-07 | 0:30 | 4.37 | 19 | 29 | 13.11 | 101 | 7.58
2009-05-07 | 0:45 | 4.25 | 19 | 29 | 13.21 | 101.4 | 7.57
2009-05-07 | 1:01 | 4.15 | 19 | 29 | 13.22 | 101.2 | 7.57
2009-05-07 | 1:16 | 4.06 | 19 | 29 | 13.24 | 101.2 | 7.57
2009-05-07 | 1:31 | 3.97 | 19 | 29 | 13.34 | 101.7 | 7.58
2009-05-07 | 1:46 | 3.89 | 19 | 29 | 13.33 | 101.4 | 7.58
2009-05-07 | 2:01 | 3.8 | 19 | 29 | 13.41 | 101.7 | 7.57
2009-05-07 | 2:16 | 3.71 | 19 | 29 | 13.4 | 101.4 | 7.57
2009-05-07 | 2:31 | 3.61 | 19 | 29 | 13.42 | 101.4 | 7.57
2009-05-07 | 2:46 | 3.52 | 19 | 29 | 13.53 | 101.9 | 7.57
2009-05-07 | 3:01 | 3.45 | 19 | 29 | 13.51 | 101.5 | 7.57
2009-05-07 | 3:16 | 3.39 | 19 | 29 | 13.51 | 101.4 | 7.57
2009-05-07 | 3:31 | 3.33 | 19 | 29 | 13.53 | 101.4 | 7.57
2009-05-07 | 3:46 | 3.29 | 19 | 29 | 13.54 | 101.4 | 7.58
2009-05-07 | 4:01 | 3.24 | 19 | 29 | 13.56 | 101.4 | 7.57
2009-05-07 | 4:16 | 3.2 | 19 | 29 | 13.58 | 101.4 | 7.58
2009-05-07 | 4:31 | 3.15 | 19 | 29 | 13.67 | 101.9 | 7.58
2009-05-07 | 4:46 | 3.1 | 19 | 29 | 13.72 | 102.2 | 7.57
2009-05-07 | 5:01 | 3.05 | 19 | 29 | 13.76 | 102.4 | 7.57
2009-05-07 | 5:16 | 3.01 | 19 | 29 | 13.72 | 102 | 7.57
2009-05-07 | 5:31 | 2.96 | 19 | 29 | 13.74 | 102 | 7.58
2009-05-07 | 5:46 | 2.92 | 19 | 29 | 13.72 | 101.7 | 7.58
2009-05-07 | 6:01 | 2.88 | 19 | 29 | 13.73 | 101.7 | 7.58
2009-05-07 | 6:16 | 2.84 | 19 | 29 | 13.76 | 101.8 | 7.58
2009-05-07 | 6:31 | 2.84 | 19 | 29 | 13.76 | 101.8 | 7.58
2009-05-07 | 6:46 | 2.83 | 19 | 29 | 13.77 | 101.8 | 7.58
2009-05-07 | 7:01 | 2.83 | 19 | 29 | 13.84 | 102.4 | 7.58
2009-05-07 | 7:16 | 2.82 | 19 | 29 | 13.88 | 102.6 | 7.58
2009-05-07 | 7:31 | 2.81 | 19 | 29 | 13.88 | 102.6 | 7.58
2009-05-07 | 7:46 | 2.8 | 19 | 29 | 13.92 | 102.8 | 7.58
2009-05-07 | 8:01 | 2.8 | 19 | 29 | 13.94 | 103 | 7.58
2009-05-07 | 8:16 | 2.82 | 19 | 29 | 13.89 | 102.7 | 7.58
2009-05-07 | 8:31 | 2.86 | 19 | 29 | 13.89 | 102.8 | 7.59
2009-05-07 | 8:46 | 2.92 | 19 | 29 | 13.89 | 103 | 7.59
2009-05-07 | 9:01 | 2.99 | 19 | 29 | 13.92 | 103.4 | 7.59
2009-05-07 | 9:16 | 3.1 | 19 | 29 | 13.91 | 103.6 | 7.59
2009-05-07 | 9:31 | 3.21 | 19 | 29 | 13.88 | 103.7 | 7.59
2009-05-07 | 9:46 | 3.37 | 19 | 29 | 13.77 | 103.3 | 7.6
2009-05-07 | 10:01 | 3.53 | 19 | 29 | 13.75 | 103.6 | 7.6
2009-05-07 | 10:16 | 3.71 | 19 | 29 | 13.71 | 103.8 | 7.61
2009-05-07 | 10:31 | 3.91 | 19 | 29 | 13.65 | 103.9 | 7.61
2009-05-07 | 10:46 | 4.12 | 19 | 29 | 13.53 | 103.6 | 7.61
2009-05-07 | 11:01 | 4.38 | 19 | 29 | 13.41 | 103.3 | 7.61
2009-05-07 | 11:16 | 4.6 | 19 | 29 | 13.31 | 103.2 | 7.61
2009-05-07 | 11:31 | 4.84 | 19 | 29 | 13.24 | 103.2 | 7.62
2009-05-07 | 11:46 | 5.14 | 19 | 30 | 13.19 | 103.6 | 7.62
2009-05-07 | 12:01 | 5.48 | 19 | 30 | 13.11 | 103.9 | 7.63
2009-05-07 | 12:16 | 5.87 | 19 | 30 | 12.93 | 103.6 | 7.63
2009-05-07 | 12:31 | 6.24 | 19 | 30 | 12.82 | 103.6 | 7.63
2009-05-07 | 12:46 | 6.58 | 19 | 30 | 12.73 | 103.8 | 7.65
2009-05-07 | 13:00 | 6.81 | 19 | 30 | 12.73 | 104.4 | 7.64
2009-05-07 | 13:15 | 7.07 | 19 | 30 | 12.63 | 104.2 | 7.65
2009-05-07 | 13:30 | 7.37 | 19 | 30 | 12.53 | 104.2 | 7.65
2009-05-07 | 13:45 | 7.71 | 19 | 30 | 12.45 | 104.4 | 7.66
2009-05-07 | 14:00 | 7.91 | 19 | 30 | 12.43 | 104.7 | 7.66
2009-05-07 | 14:15 | 8.18 | 19 | 30 | 12.36 | 104.8 | 7.67
2009-05-07 | 14:30 | 8.4 | 19 | 30 | 12.32 | 105.1 | 7.66
2009-05-07 | 14:45 | 8.61 | 19 | 30 | 12.18 | 104.3 | 7.67
2009-05-07 | 15:00 | 8.7 | 20 | 30 | 12.16 | 104.5 | 7.67
2009-05-07 | 15:15 | 8.75 | 20 | 30 | 12.09 | 103.9 | 7.68
2009-05-07 | 15:30 | 8.63 | 20 | 30 | 12.09 | 103.7 | 7.67
2009-05-07 | 15:45 | 8.64 | 20 | 30 | 12.03 | 103.2 | 7.68
2009-05-07 | 16:00 | 8.73 | 20 | 30 | 12.02 | 103.3 | 7.68
2009-05-07 | 16:15 | 8.76 | 20 | 30 | 12.03 | 103.5 | 7.67
2009-05-07 | 16:30 | 8.68 | 20 | 30 | 12.03 | 103.3 | 7.68
2009-05-07 | 16:45 | 8.67 | 20 | 30 | 11.95 | 102.6 | 7.68
2009-05-07 | 17:00 | 8.66 | 20 | 30 | 11.99 | 102.9 | 7.68
2009-05-07 | 17:15 | 8.58 | 20 | 30 | 11.99 | 102.6 | 7.68
2009-05-07 | 17:30 | 8.47 | 20 | 30 | 11.98 | 102.3 | 7.67
2009-05-07 | 17:45 | 8.31 | 20 | 30 | 12.03 | 102.4 | 7.67
2009-05-07 | 18:00 | 8.18 | 20 | 30 | 12 | 101.8 | 7.67
2009-05-07 | 18:15 | 8.06 | 20 | 30 | 12.02 | 101.6 | 7.66
2009-05-07 | 18:30 | 8.01 | 20 | 30 | 12.02 | 101.5 | 7.66
2009-05-07 | 18:45 | 7.9 | 20 | 30 | 12.04 | 101.4 | 7.66
2009-05-07 | 19:00 | 7.75 | 20 | 30 | 12.13 | 101.8 | 7.66
2009-05-07 | 19:15 | 7.57 | 20 | 30 | 12.15 | 101.5 | 7.66
2009-05-07 | 19:30 | 7.4 | 20 | 30 | 12.19 | 101.4 | 7.65
2009-05-07 | 19:45 | 7.24 | 20 | 30 | 12.22 | 101.2 | 7.65
2009-05-07 | 20:00 | 7.07 | 20 | 30 | 12.27 | 101.3 | 7.65
2009-05-07 | 20:15 | 6.94 | 20 | 30 | 12.35 | 101.6 | 7.65
2009-05-07 | 20:30 | 6.84 | 20 | 30 | 12.35 | 101.3 | 7.64
2009-05-07 | 20:45 | 6.74 | 20 | 30 | 12.35 | 101.1 | 7.64
2009-05-07 | 21:00 | 6.61 | 20 | 30 | 12.49 | 101.9 | 7.64
2009-05-07 | 21:15 | 6.47 | 20 | 30 | 12.55 | 102 | 7.63
2009-05-07 | 21:30 | 6.34 | 20 | 30 | 12.53 | 101.5 | 7.64
2009-05-07 | 21:45 | 6.19 | 20 | 30 | 12.61 | 101.8 | 7.63
2009-05-07 | 22:00 | 6.03 | 20 | 31 | 12.69 | 102 | 7.63
2009-05-07 | 22:15 | 5.86 | 20 | 31 | 12.7 | 101.6 | 7.63
2009-05-07 | 22:30 | 5.66 | 20 | 31 | 12.76 | 101.6 | 7.62
2009-05-07 | 22:45 | 5.47 | 20 | 31 | 12.82 | 101.6 | 7.62
2009-05-07 | 23:00 | 5.28 | 20 | 31 | 12.89 | 101.6 | 7.61
2009-05-07 | 23:15 | 5.09 | 20 | 31 | 13.01 | 102.1 | 7.61
2009-05-07 | 23:30 | 4.89 | 20 | 31 | 13.03 | 101.8 | 7.6
2009-05-07 | 23:45 | 4.71 | 20 | 31 | 13.2 | 102.6 | 7.6
2009-05-08 | 0:00 | 4.53 | 20 | 31 | 13.28 | 102.8 | 7.6
2009-05-08 | 0:15 | 4.34 | 20 | 31 | 13.35 | 102.8 | 7.59
2009-05-08 | 0:30 | 4.16 | 20 | 31 | 13.36 | 102.4 | 7.59
2009-05-08 | 0:45 | 4 | 20 | 31 | 13.5 | 103 | 7.59
2009-05-08 | 1:01 | 3.84 | 20 | 31 | 13.58 | 103.1 | 7.59
2009-05-08 | 1:16 | 3.69 | 20 | 31 | 13.58 | 102.7 | 7.58
2009-05-08 | 1:31 | 3.54 | 20 | 31 | 13.69 | 103.2 | 7.58
2009-05-08 | 1:46 | 3.41 | 20 | 31 | 13.75 | 103.3 | 7.57
2009-05-08 | 2:01 | 3.28 | 20 | 31 | 13.77 | 103.1 | 7.58
2009-05-08 | 2:16 | 3.17 | 20 | 31 | 13.78 | 102.8 | 7.58
2009-05-08 | 2:31 | 3.07 | 20 | 31 | 13.82 | 102.8 | 7.59
2009-05-08 | 2:46 | 2.97 | 20 | 31 | 13.88 | 103.1 | 7.58
2009-05-08 | 3:01 | 2.86 | 20 | 31 | 13.94 | 103.1 | 7.58
2009-05-08 | 3:16 | 2.76 | 20 | 31 | 13.95 | 103 | 7.58
2009-05-08 | 3:31 | 2.68 | 20 | 31 | 13.98 | 103 | 7.57
2009-05-08 | 3:46 | 2.59 | 20 | 31 | 14.01 | 102.9 | 7.58
2009-05-08 | 4:01 | 2.51 | 20 | 31 | 14.1 | 103.4 | 7.58
2009-05-08 | 4:16 | 2.45 | 20 | 31 | 14.16 | 103.6 | 7.58
2009-05-08 | 4:31 | 2.38 | 20 | 31 | 14.22 | 103.9 | 7.57
2009-05-08 | 4:46 | 2.31 | 20 | 31 | 14.2 | 103.6 | 7.57
2009-05-08 | 5:01 | 2.24 | 20 | 31 | 14.3 | 104 | 7.57
2009-05-08 | 5:16 | 2.18 | 20 | 31 | 14.26 | 103.6 | 7.57
2009-05-08 | 5:31 | 2.12 | 20 | 31 | 14.32 | 103.8 | 7.57
2009-05-08 | 5:46 | 2.06 | 20 | 31 | 14.28 | 103.5 | 7.57
2009-05-08 | 6:01 | 2.01 | 20 | 31 | 14.31 | 103.5 | 7.58
2009-05-08 | 6:16 | 1.96 | 20 | 31 | 14.38 | 103.8 | 7.57
2009-05-08 | 6:31 | 1.91 | 20 | 31 | 14.34 | 103.4 | 7.57
2009-05-08 | 6:46 | 1.87 | 20 | 31 | 14.46 | 104.2 | 7.57
2009-05-08 | 7:01 | 1.83 | 20 | 32 | 14.48 | 104.2 | 7.56
2009-05-08 | 7:16 | 1.8 | 20 | 32 | 14.54 | 104.5 | 7.57
2009-05-08 | 7:31 | 1.78 | 20 | 32 | 14.55 | 104.6 | 7.58
2009-05-08 | 7:46 | 1.79 | 21 | 32 | 14.51 | 104.3 | 7.58
2009-05-08 | 8:01 | 1.81 | 21 | 32 | 14.49 | 104.2 | 7.59
2009-05-08 | 8:16 | 1.85 | 21 | 32 | 14.47 | 104.2 | 7.6
2009-05-08 | 8:31 | 1.91 | 21 | 32 | 14.45 | 104.3 | 7.6
2009-05-08 | 8:46 | 2 | 21 | 32 | 14.48 | 104.7 | 7.61
2009-05-08 | 9:01 | 2.12 | 21 | 32 | 14.48 | 105 | 7.62
2009-05-08 | 9:16 | 2.26 | 21 | 32 | 14.38 | 104.7 | 7.62
2009-05-08 | 9:31 | 2.42 | 21 | 32 | 14.37 | 105.1 | 7.63
2009-05-08 | 9:46 | 2.62 | 21 | 32 | 14.31 | 105.2 | 7.63
2009-05-08 | 10:01 | 2.83 | 21 | 32 | 14.26 | 105.5 | 7.63
2009-05-08 | 10:16 | 3.07 | 21 | 32 | 14.1 | 105 | 7.64
2009-05-08 | 10:31 | 3.35 | 21 | 32 | 14.02 | 105.1 | 7.65
2009-05-08 | 10:46 | 3.63 | 21 | 32 | 13.95 | 105.4 | 7.65
2009-05-08 | 11:01 | 3.94 | 21 | 32 | 13.83 | 105.4 | 7.65
2009-05-08 | 11:16 | 4.25 | 21 | 32 | 13.66 | 104.9 | 7.66
2009-05-08 | 11:31 | 4.59 | 21 | 32 | 13.58 | 105.2 | 7.66
2009-05-08 | 11:46 | 4.96 | 21 | 32 | 13.37 | 104.6 | 7.67
2009-05-08 | 12:01 | 5.33 | 21 | 32 | 13.21 | 104.4 | 7.67
2009-05-08 | 12:16 | 5.73 | 21 | 32 | 13.14 | 104.9 | 7.67
2009-05-08 | 12:31 | 6.13 | 21 | 32 | 13.08 | 105.5 | 7.67
2009-05-08 | 12:46 | 6.55 | 21 | 32 | 13 | 105.9 | 7.68
2009-05-08 | 13:00 | 6.94 | 21 | 32 | 12.85 | 105.7 | 7.68
2009-05-08 | 13:15 | 7.34 | 21 | 32 | 12.78 | 106.2 | 7.69
2009-05-08 | 13:30 | 7.7 | 21 | 32 | 12.72 | 106.6 | 7.69
2009-05-08 | 13:45 | 8.03 | 21 | 32 | 12.57 | 106.2 | 7.69
2009-05-08 | 14:00 | 8.35 | 21 | 32 | 12.43 | 105.9 | 7.7
2009-05-08 | 14:15 | 8.65 | 21 | 32 | 12.4 | 106.4 | 7.7
2009-05-08 | 14:30 | 8.9 | 21 | 32 | 12.29 | 106.1 | 7.7
2009-05-08 | 14:45 | 9.12 | 21 | 32 | 12.15 | 105.4 | 7.7
2009-05-08 | 15:00 | 9.29 | 21 | 32 | 12.03 | 104.8 | 7.71
2009-05-08 | 15:15 | 9.43 | 21 | 32 | 11.96 | 104.5 | 7.71
2009-05-08 | 15:30 | 9.52 | 21 | 32 | 11.94 | 104.6 | 7.71
2009-05-08 | 15:45 | 9.56 | 21 | 32 | 11.87 | 104.1 | 7.71
2009-05-08 | 16:00 | 9.57 | 21 | 32 | 11.9 | 104.3 | 7.71
2009-05-08 | 16:15 | 9.52 | 21 | 32 | 11.89 | 104.2 | 7.7
2009-05-08 | 16:30 | 9.46 | 21 | 32 | 11.82 | 103.4 | 7.71
2009-05-08 | 16:45 | 9.38 | 21 | 32 | 11.8 | 103 | 7.7
2009-05-08 | 17:00 | 9.29 | 21 | 32 | 11.77 | 102.6 | 7.7
2009-05-08 | 17:15 | 9.18 | 21 | 32 | 11.76 | 102.2 | 7.7
2009-05-08 | 17:30 | 9.08 | 21 | 32 | 11.84 | 102.6 | 7.7
2009-05-08 | 17:45 | 8.97 | 21 | 32 | 11.82 | 102.2 | 7.7
2009-05-08 | 18:00 | 8.86 | 21 | 32 | 11.82 | 101.9 | 7.7
2009-05-08 | 18:15 | 8.75 | 21 | 32 | 11.83 | 101.7 | 7.7
2009-05-08 | 18:30 | 8.65 | 21 | 32 | 11.89 | 102 | 7.69
2009-05-08 | 18:45 | 8.54 | 21 | 32 | 11.88 | 101.6 | 7.69
2009-05-08 | 19:00 | 8.44 | 21 | 32 | 11.93 | 101.8 | 7.69
2009-05-08 | 19:15 | 8.34 | 21 | 32 | 11.98 | 102 | 7.69
2009-05-08 | 19:30 | 8.23 | 21 | 32 | 12 | 101.9 | 7.68
2009-05-08 | 19:45 | 8.12 | 21 | 32 | 11.97 | 101.4 | 7.68
2009-05-08 | 20:00 | 8.03 | 21 | 32 | 12.04 | 101.7 | 7.68
2009-05-08 | 20:15 | 7.94 | 21 | 32 | 12.05 | 101.6 | 7.67
2009-05-08 | 20:30 | 7.83 | 21 | 32 | 12.1 | 101.8 | 7.67
2009-05-08 | 20:45 | 7.73 | 21 | 32 | 12.06 | 101.2 | 7.66
2009-05-08 | 21:00 | 7.62 | 21 | 32 | 12.11 | 101.3 | 7.65
2009-05-08 | 21:15 | 7.52 | 21 | 32 | 12.09 | 100.9 | 7.64
2009-05-08 | 21:30 | 7.41 | 21 | 32 | 12.17 | 101.2 | 7.64
2009-05-08 | 21:45 | 7.27 | 21 | 32 | 12.17 | 101 | 7.62
2009-05-08 | 22:00 | 7.13 | 21 | 32 | 12.2 | 100.8 | 7.62
2009-05-08 | 22:15 | 6.97 | 21 | 32 | 12.25 | 100.9 | 7.61
2009-05-08 | 22:30 | 6.82 | 21 | 32 | 12.29 | 100.8 | 7.61
2009-05-08 | 22:45 | 6.63 | 21 | 32 | 12.36 | 100.9 | 7.61
2009-05-08 | 23:00 | 6.44 | 21 | 32 | 12.48 | 101.4 | 7.61
2009-05-08 | 23:15 | 6.24 | 21 | 32 | 12.58 | 101.7 | 7.61
2009-05-08 | 23:30 | 6.05 | 21 | 32 | 12.66 | 101.9 | 7.61
2009-05-08 | 23:45 | 5.85 | 21 | 32 | 12.76 | 102.1 | 7.59
2009-05-09 | 0:00 | 5.65 | 21 | 32 | 12.79 | 101.8 | 7.6
2009-05-09 | 0:15 | 5.46 | 21 | 32 | 12.82 | 101.6 | 7.6
2009-05-09 | 0:30 | 5.28 | 21 | 32 | 12.88 | 101.6 | 7.6
2009-05-09 | 0:45 | 5.12 | 21 | 32 | 12.99 | 102 | 7.6
2009-05-09 | 1:01 | 4.96 | 21 | 32 | 13.02 | 101.8 | 7.6
2009-05-09 | 1:16 | 4.8 | 21 | 32 | 13.14 | 102.4 | 7.6
2009-05-09 | 1:31 | 4.66 | 21 | 32 | 13.21 | 102.6 | 7.59
2009-05-09 | 1:46 | 4.51 | 21 | 32 | 13.23 | 102.2 | 7.59
2009-05-09 | 2:01 | 4.37 | 21 | 32 | 13.31 | 102.5 | 7.59
2009-05-09 | 2:16 | 4.24 | 21 | 32 | 13.38 | 102.7 | 7.59
2009-05-09 | 2:31 | 4.12 | 21 | 32 | 13.44 | 102.9 | 7.59
2009-05-09 | 2:46 | 4.01 | 21 | 32 | 13.43 | 102.5 | 7.59
2009-05-09 | 3:01 | 3.89 | 21 | 32 | 13.45 | 102.3 | 7.59
2009-05-09 | 3:16 | 3.79 | 21 | 32 | 13.58 | 103 | 7.59
2009-05-09 | 3:31 | 3.68 | 21 | 32 | 13.61 | 103 | 7.58
2009-05-09 | 3:46 | 3.59 | 21 | 32 | 13.6 | 102.7 | 7.59
2009-05-09 | 4:01 | 3.49 | 21 | 32 | 13.67 | 102.9 | 7.58
2009-05-09 | 4:16 | 3.39 | 21 | 32 | 13.73 | 103 | 7.58
2009-05-09 | 4:31 | 3.3 | 21 | 32 | 13.72 | 102.8 | 7.58
2009-05-09 | 4:46 | 3.2 | 21 | 32 | 13.81 | 103.2 | 7.58
2009-05-09 | 5:01 | 3.12 | 21 | 32 | 13.8 | 102.8 | 7.58
2009-05-09 | 5:16 | 3.03 | 21 | 32 | 13.81 | 102.7 | 7.58
2009-05-09 | 5:31 | 2.94 | 21 | 32 | 13.9 | 103.1 | 7.57
2009-05-09 | 5:46 | 2.86 | 21 | 32 | 13.92 | 103.1 | 7.58
2009-05-09 | 6:01 | 2.79 | 21 | 32 | 13.92 | 102.8 | 7.57
2009-05-09 | 6:16 | 2.71 | 21 | 32 | 14 | 103.2 | 7.57
2009-05-09 | 6:31 | 2.65 | 21 | 32 | 14.07 | 103.5 | 7.58
2009-05-09 | 6:46 | 2.6 | 21 | 32 | 14.07 | 103.4 | 7.59
2009-05-09 | 7:01 | 2.56 | 21 | 32 | 14.07 | 103.3 | 7.58
2009-05-09 | 7:16 | 2.52 | 21 | 32 | 14.08 | 103.3 | 7.59
2009-05-09 | 7:31 | 2.51 | 21 | 32 | 14.17 | 103.9 | 7.59
2009-05-09 | 7:46 | 2.5 | 21 | 32 | 14.2 | 104.1 | 7.59
2009-05-09 | 8:01 | 2.52 | 21 | 32 | 14.17 | 103.9 | 7.59
2009-05-09 | 8:16 | 2.55 | 21 | 32 | 14.16 | 103.9 | 7.6
2009-05-09 | 8:31 | 2.62 | 21 | 32 | 14.2 | 104.4 | 7.61
2009-05-09 | 8:46 | 2.7 | 21 | 32 | 14.15 | 104.3 | 7.62
2009-05-09 | 9:01 | 2.82 | 21 | 32 | 14.09 | 104.2 | 7.63
2009-05-09 | 9:16 | 2.97 | 21 | 32 | 14.04 | 104.2 | 7.64
2009-05-09 | 9:31 | 3.14 | 21 | 32 | 13.98 | 104.2 | 7.65
2009-05-09 | 9:46 | 3.33 | 21 | 32 | 13.97 | 104.7 | 7.65
2009-05-09 | 10:01 | 3.54 | 21 | 32 | 13.92 | 104.9 | 7.65
2009-05-09 | 10:16 | 3.78 | 21 | 32 | 13.79 | 104.6 | 7.66
2009-05-09 | 10:31 | 4.05 | 21 | 32 | 13.66 | 104.4 | 7.67
2009-05-09 | 10:46 | 4.34 | 21 | 32 | 13.56 | 104.4 | 7.67
2009-05-09 | 11:01 | 4.67 | 21 | 32 | 13.44 | 104.3 | 7.67
2009-05-09 | 11:16 | 5 | 21 | 32 | 13.3 | 104.2 | 7.68
2009-05-09 | 11:31 | 5.36 | 21 | 32 | 13.16 | 104 | 7.68
2009-05-09 | 11:46 | 5.75 | 21 | 32 | 13.03 | 104 | 7.68
2009-05-09 | 12:01 | 6.15 | 21 | 32 | 12.89 | 103.9 | 7.69
2009-05-09 | 12:16 | 6.57 | 21 | 32 | 12.75 | 103.9 | 7.69
2009-05-09 | 12:31 | 6.97 | 21 | 32 | 12.72 | 104.7 | 7.7
2009-05-09 | 12:46 | 7.42 | 21 | 32 | 12.65 | 105.3 | 7.7
2009-05-09 | 13:00 | 7.86 | 21 | 32 | 12.56 | 105.7 | 7.71
2009-05-09 | 13:15 | 8.25 | 21 | 32 | 12.41 | 105.4 | 7.71
2009-05-09 | 13:30 | 8.63 | 21 | 32 | 12.34 | 105.8 | 7.72
2009-05-09 | 13:45 | 8.98 | 21 | 32 | 12.27 | 106.1 | 7.72
2009-05-09 | 14:00 | 9.31 | 21 | 32 | 12.12 | 105.6 | 7.72
2009-05-09 | 14:15 | 9.59 | 21 | 32 | 12 | 105.3 | 7.73
2009-05-09 | 14:30 | 9.82 | 21 | 32 | 11.95 | 105.4 | 7.73
2009-05-09 | 14:45 | 10.02 | 21 | 32 | 11.88 | 105.3 | 7.73
2009-05-09 | 15:00 | 10.17 | 21 | 32 | 11.75 | 104.5 | 7.73
2009-05-09 | 15:15 | 10.2 | 21 | 32 | 11.64 | 103.6 | 7.73
2009-05-09 | 15:30 | 10.21 | 21 | 32 | 11.61 | 103.3 | 7.73
2009-05-09 | 15:45 | 9.99 | 21 | 32 | 11.67 | 103.3 | 7.73
2009-05-09 | 16:00 | 10.02 | 21 | 32 | 11.73 | 103.9 | 7.73
2009-05-09 | 16:15 | 9.78 | 21 | 32 | 11.73 | 103.4 | 7.72
2009-05-09 | 16:30 | 9.75 | 21 | 32 | 11.67 | 102.8 | 7.72
2009-05-09 | 16:45 | 9.77 | 21 | 32 | 11.7 | 103.1 | 7.72
2009-05-09 | 17:00 | 9.64 | 21 | 32 | 11.71 | 102.9 | 7.72
2009-05-09 | 17:15 | 9.53 | 21 | 32 | 11.71 | 102.6 | 7.71
2009-05-09 | 17:30 | 9.39 | 21 | 32 | 11.7 | 102.1 | 7.72
2009-05-09 | 17:45 | 9.18 | 21 | 32 | 11.72 | 101.9 | 7.72
2009-05-09 | 18:00 | 8.99 | 21 | 32 | 11.83 | 102.3 | 7.71
2009-05-09 | 18:15 | 8.85 | 21 | 32 | 11.83 | 101.9 | 7.71
2009-05-09 | 18:30 | 8.82 | 21 | 32 | 11.85 | 102.1 | 7.71
2009-05-09 | 18:45 | 8.78 | 21 | 32 | 11.87 | 102.2 | 7.7
2009-05-09 | 19:00 | 8.69 | 21 | 32 | 11.91 | 102.2 | 7.7
2009-05-09 | 19:15 | 8.61 | 21 | 32 | 11.9 | 102 | 7.7
2009-05-09 | 19:30 | 8.51 | 21 | 32 | 11.89 | 101.6 | 7.7
2009-05-09 | 19:45 | 8.39 | 21 | 32 | 11.94 | 101.8 | 7.7
2009-05-09 | 20:00 | 8.26 | 21 | 32 | 12 | 102 | 7.69
2009-05-09 | 20:15 | 8.1 | 21 | 32 | 12.04 | 101.9 | 7.68
2009-05-09 | 20:30 | 7.94 | 21 | 32 | 12.02 | 101.4 | 7.68
2009-05-09 | 20:45 | 7.79 | 21 | 32 | 12.11 | 101.8 | 7.66
2009-05-09 | 21:00 | 7.66 | 21 | 32 | 12.16 | 101.8 | 7.65
2009-05-09 | 21:15 | 7.53 | 21 | 32 | 12.14 | 101.4 | 7.65
2009-05-09 | 21:30 | 7.43 | 21 | 32 | 12.2 | 101.6 | 7.65
2009-05-09 | 21:45 | 7.31 | 21 | 32 | 12.26 | 101.8 | 7.64
2009-05-09 | 22:00 | 7.18 | 21 | 32 | 12.25 | 101.4 | 7.64
2009-05-09 | 22:15 | 7.06 | 21 | 32 | 12.34 | 101.8 | 7.64
2009-05-09 | 22:30 | 6.93 | 21 | 32 | 12.41 | 102 | 7.63
2009-05-09 | 22:45 | 6.79 | 21 | 32 | 12.43 | 101.9 | 7.63
2009-05-09 | 23:00 | 6.63 | 21 | 32 | 12.5 | 102.1 | 7.63
2009-05-09 | 23:15 | 6.45 | 21 | 32 | 12.56 | 102.1 | 7.62
2009-05-09 | 23:30 | 6.28 | 21 | 32 | 12.58 | 101.8 | 7.61
2009-05-09 | 23:45 | 6.13 | 21 | 32 | 12.66 | 102.1 | 7.62
2009-05-10 | 0:00 | 5.98 | 21 | 32 | 12.74 | 102.3 | 7.6
2009-05-10 | 0:15 | 5.84 | 21 | 32 | 12.74 | 102 | 7.6
2009-05-10 | 0:30 | 5.71 | 21 | 32 | 12.76 | 101.7 | 7.6
2009-05-10 | 0:45 | 5.56 | 21 | 32 | 12.88 | 102.3 | 7.59
2009-05-10 | 1:01 | 5.42 | 21 | 32 | 12.9 | 102.1 | 7.59
2009-05-10 | 1:16 | 5.29 | 21 | 32 | 12.95 | 102.1 | 7.59
2009-05-10 | 1:31 | 5.16 | 21 | 32 | 12.99 | 102.2 | 7.58
2009-05-10 | 1:46 | 5.04 | 21 | 32 | 13.07 | 102.5 | 7.58
2009-05-10 | 2:01 | 4.91 | 21 | 32 | 13.09 | 102.2 | 7.58
2009-05-10 | 2:16 | 4.79 | 21 | 32 | 13.12 | 102.1 | 7.58
2009-05-10 | 2:31 | 4.67 | 21 | 32 | 13.22 | 102.7 | 7.58
2009-05-10 | 2:46 | 4.56 | 21 | 32 | 13.28 | 102.8 | 7.58
2009-05-10 | 3:01 | 4.44 | 21 | 32 | 13.28 | 102.5 | 7.58
2009-05-10 | 3:16 | 4.34 | 21 | 32 | 13.3 | 102.4 | 7.58
2009-05-10 | 3:31 | 4.23 | 21 | 32 | 13.33 | 102.3 | 7.57
2009-05-10 | 3:46 | 4.14 | 21 | 32 | 13.36 | 102.3 | 7.58
2009-05-10 | 4:01 | 4.03 | 21 | 32 | 13.43 | 102.5 | 7.58
2009-05-10 | 4:16 | 3.93 | 21 | 32 | 13.46 | 102.5 | 7.59
2009-05-10 | 4:31 | 3.83 | 21 | 32 | 13.49 | 102.4 | 7.59
2009-05-10 | 4:46 | 3.73 | 21 | 32 | 13.53 | 102.5 | 7.59
2009-05-10 | 5:01 | 3.64 | 21 | 32 | 13.56 | 102.5 | 7.59
2009-05-10 | 5:16 | 3.55 | 21 | 32 | 13.67 | 103.1 | 7.59
2009-05-10 | 5:31 | 3.47 | 21 | 32 | 13.71 | 103.1 | 7.58
2009-05-10 | 5:46 | 3.39 | 21 | 32 | 13.76 | 103.3 | 7.59
2009-05-10 | 6:01 | 3.3 | 21 | 32 | 13.83 | 103.5 | 7.58
2009-05-10 | 6:16 | 3.23 | 21 | 32 | 13.82 | 103.3 | 7.59
2009-05-10 | 6:31 | 3.16 | 21 | 32 | 13.88 | 103.6 | 7.59
2009-05-10 | 6:46 | 3.1 | 21 | 32 | 13.92 | 103.7 | 7.59
2009-05-10 | 7:01 | 3.05 | 21 | 32 | 13.9 | 103.4 | 7.59
2009-05-10 | 7:16 | 3.01 | 21 | 32 | 13.97 | 103.8 | 7.59
2009-05-10 | 7:31 | 2.99 | 21 | 32 | 14.02 | 104.1 | 7.6
2009-05-10 | 7:46 | 2.99 | 21 | 32 | 14.05 | 104.3 | 7.6
2009-05-10 | 8:01 | 3.01 | 21 | 33 | 14 | 104 | 7.62
2009-05-10 | 8:16 | 3.05 | 21 | 33 | 14.05 | 104.5 | 7.62
2009-05-10 | 8:31 | 3.12 | 21 | 33 | 14 | 104.3 | 7.63
2009-05-10 | 8:46 | 3.21 | 21 | 33 | 13.95 | 104.2 | 7.64
2009-05-10 | 9:01 | 3.34 | 21 | 33 | 13.91 | 104.3 | 7.64
2009-05-10 | 9:16 | 3.49 | 21 | 33 | 13.92 | 104.7 | 7.65
2009-05-10 | 9:31 | 3.66 | 21 | 33 | 13.82 | 104.5 | 7.66
2009-05-10 | 9:46 | 3.85 | 21 | 33 | 13.8 | 104.9 | 7.67
2009-05-10 | 10:01 | 4.06 | 21 | 33 | 13.76 | 105.1 | 7.67
2009-05-10 | 10:16 | 4.31 | 21 | 33 | 13.61 | 104.7 | 7.68
2009-05-10 | 10:31 | 4.59 | 21 | 33 | 13.5 | 104.6 | 7.68
2009-05-10 | 10:46 | 4.88 | 21 | 33 | 13.45 | 105 | 7.68
2009-05-10 | 11:01 | 5.17 | 21 | 33 | 13.31 | 104.6 | 7.69
2009-05-10 | 11:16 | 5.49 | 21 | 33 | 13.23 | 104.9 | 7.69
2009-05-10 | 11:31 | 5.81 | 21 | 33 | 13.16 | 105.2 | 7.69
2009-05-10 | 11:46 | 6.19 | 21 | 33 | 13.03 | 105.1 | 7.7
2009-05-10 | 12:01 | 6.59 | 21 | 33 | 12.81 | 104.5 | 7.7
2009-05-10 | 12:16 | 7.02 | 21 | 33 | 12.69 | 104.6 | 7.71
2009-05-10 | 12:31 | 7.42 | 21 | 33 | 12.58 | 104.7 | 7.71
2009-05-10 | 12:46 | 7.73 | 21 | 33 | 12.49 | 104.8 | 7.72
2009-05-10 | 13:00 | 8.11 | 21 | 33 | 12.43 | 105.3 | 7.72
2009-05-10 | 13:15 | 8.4 | 21 | 33 | 12.33 | 105.1 | 7.73
2009-05-10 | 13:30 | 8.66 | 21 | 33 | 12.24 | 105 | 7.73
2009-05-10 | 13:45 | 8.93 | 21 | 33 | 12.21 | 105.4 | 7.73
2009-05-10 | 14:00 | 9.23 | 21 | 33 | 12.09 | 105.2 | 7.74
2009-05-10 | 14:15 | 9.47 | 21 | 33 | 12.03 | 105.2 | 7.75
2009-05-10 | 14:30 | 9.62 | 21 | 33 | 11.92 | 104.6 | 7.74
2009-05-10 | 14:45 | 9.81 | 21 | 33 | 11.86 | 104.6 | 7.74
2009-05-10 | 15:00 | 9.81 | 21 | 33 | 11.81 | 104.2 | 7.75
2009-05-10 | 15:15 | 9.75 | 21 | 33 | 11.78 | 103.7 | 7.75
2009-05-10 | 15:30 | 9.61 | 21 | 33 | 11.76 | 103.2 | 7.74
2009-05-10 | 15:45 | 9.46 | 21 | 33 | 11.73 | 102.6 | 7.74
2009-05-10 | 16:00 | 9.34 | 21 | 33 | 11.75 | 102.5 | 7.74
2009-05-10 | 16:15 | 9.22 | 21 | 33 | 11.75 | 102.2 | 7.74
2009-05-10 | 16:30 | 9.08 | 21 | 33 | 11.78 | 102.1 | 7.74
2009-05-10 | 16:45 | 8.96 | 21 | 33 | 11.8 | 102 | 7.73
2009-05-10 | 17:00 | 8.84 | 21 | 33 | 11.88 | 102.4 | 7.73
2009-05-10 | 17:15 | 8.76 | 21 | 33 | 11.88 | 102.1 | 7.73
2009-05-10 | 17:30 | 8.7 | 21 | 33 | 11.93 | 102.5 | 7.73
2009-05-10 | 17:45 | 8.69 | 21 | 33 | 11.95 | 102.6 | 7.73
2009-05-10 | 18:00 | 8.71 | 21 | 33 | 11.94 | 102.6 | 7.73
2009-05-10 | 18:15 | 8.8 | 21 | 33 | 11.85 | 102 | 7.73
2009-05-10 | 18:30 | 8.91 | 21 | 33 | 11.77 | 101.6 | 7.73
2009-05-10 | 18:45 | 8.98 | 21 | 33 | 11.76 | 101.7 | 7.73
2009-05-10 | 19:00 | 9.06 | 21 | 33 | 11.75 | 101.8 | 7.73
2009-05-10 | 19:15 | 9.14 | 21 | 33 | 11.66 | 101.2 | 7.73
2009-05-10 | 19:30 | 9.24 | 21 | 33 | 11.58 | 100.8 | 7.73
2009-05-10 | 19:45 | 9.3 | 21 | 33 | 11.53 | 100.5 | 7.74
2009-05-10 | 20:00 | 9.3 | 21 | 33 | 11.57 | 100.8 | 7.73
2009-05-10 | 20:15 | 9.2 | 21 | 33 | 11.56 | 100.5 | 7.73
2009-05-10 | 20:30 | 9.04 | 21 | 33 | 11.58 | 100.3 | 7.72
2009-05-10 | 20:45 | 8.83 | 21 | 33 | 11.63 | 100.2 | 7.72
2009-05-10 | 21:00 | 8.61 | 21 | 33 | 11.76 | 100.8 | 7.71
2009-05-10 | 21:15 | 8.36 | 21 | 33 | 11.86 | 101 | 7.7
2009-05-10 | 21:30 | 8.14 | 21 | 33 | 11.95 | 101.2 | 7.69
2009-05-10 | 21:45 | 7.92 | 21 | 33 | 11.98 | 100.9 | 7.69
2009-05-10 | 22:00 | 7.73 | 21 | 33 | 12.02 | 100.8 | 7.69
2009-05-10 | 22:15 | 7.55 | 21 | 33 | 12.13 | 101.3 | 7.68
2009-05-10 | 22:30 | 7.39 | 21 | 33 | 12.2 | 101.5 | 7.69
2009-05-10 | 22:45 | 7.24 | 21 | 32 | 12.28 | 101.7 | 7.68
2009-05-10 | 23:00 | 7.11 | 21 | 32 | 12.25 | 101.2 | 7.67
2009-05-10 | 23:15 | 7.01 | 21 | 32 | 12.26 | 101 | 7.67
2009-05-10 | 23:30 | 6.94 | 21 | 32 | 12.27 | 101 | 7.67
2009-05-10 | 23:45 | 6.9 | 21 | 32 | 12.28 | 100.9 | 7.67
2009-05-11 | 0:00 | 6.85 | 21 | 32 | 12.35 | 101.3 | 7.66
2009-05-11 | 0:15 | 6.8 | 21 | 32 | 12.32 | 101 | 7.66
2009-05-11 | 0:30 | 6.73 | 21 | 32 | 12.33 | 100.9 | 7.66
2009-05-11 | 0:45 | 6.64 | 21 | 32 | 12.41 | 101.3 | 7.65
2009-05-11 | 1:01 | 6.53 | 21 | 32 | 12.42 | 101.1 | 7.64
2009-05-11 | 1:16 | 6.4 | 21 | 32 | 12.51 | 101.5 | 7.64
2009-05-11 | 1:31 | 6.25 | 21 | 32 | 12.58 | 101.7 | 7.64
2009-05-11 | 1:46 | 6.09 | 21 | 32 | 12.59 | 101.4 | 7.64
2009-05-11 | 2:01 | 5.94 | 21 | 32 | 12.62 | 101.2 | 7.64
2009-05-11 | 2:16 | 5.79 | 21 | 32 | 12.74 | 101.8 | 7.63
2009-05-11 | 2:31 | 5.64 | 20 | 31 | 12.82 | 102 | 7.63
2009-05-11 | 2:46 | 5.49 | 20 | 31 | 12.84 | 101.8 | 7.63
2009-05-11 | 3:01 | 5.36 | 20 | 31 | 12.92 | 102.1 | 7.63
2009-05-11 | 3:16 | 5.23 | 20 | 31 | 13 | 102.4 | 7.62
2009-05-11 | 3:31 | 5.12 | 20 | 31 | 12.99 | 102.1 | 7.63
2009-05-11 | 3:46 | 5.01 | 20 | 31 | 13.01 | 101.9 | 7.63
2009-05-11 | 4:01 | 4.92 | 20 | 31 | 13.04 | 101.9 | 7.64
2009-05-11 | 4:16 | 4.84 | 20 | 31 | 13.06 | 101.9 | 7.63
2009-05-11 | 4:31 | 4.75 | 20 | 31 | 13.1 | 101.9 | 7.64
2009-05-11 | 4:46 | 4.66 | 20 | 31 | 13.19 | 102.4 | 7.63
2009-05-11 | 5:01 | 4.56 | 20 | 31 | 13.19 | 102.1 | 7.63
2009-05-11 | 5:16 | 4.47 | 20 | 31 | 13.28 | 102.5 | 7.63
2009-05-11 | 5:31 | 4.37 | 20 | 31 | 13.28 | 102.3 | 7.63
2009-05-11 | 5:46 | 4.29 | 20 | 31 | 13.29 | 102.2 | 7.63
2009-05-11 | 6:01 | 4.21 | 20 | 31 | 13.32 | 102.2 | 7.63
2009-05-11 | 6:16 | 4.14 | 20 | 31 | 13.35 | 102.2 | 7.63
2009-05-11 | 6:31 | 4.08 | 20 | 31 | 13.39 | 102.4 | 7.63
2009-05-11 | 6:46 | 4.02 | 20 | 31 | 13.41 | 102.4 | 7.63
2009-05-11 | 7:01 | 3.98 | 20 | 31 | 13.5 | 102.9 | 7.63
2009-05-11 | 7:16 | 3.94 | 20 | 31 | 13.57 | 103.4 | 7.63
2009-05-11 | 7:31 | 3.91 | 20 | 31 | 13.6 | 103.5 | 7.64
2009-05-11 | 7:46 | 3.91 | 20 | 31 | 13.61 | 103.5 | 7.64
2009-05-11 | 8:01 | 3.93 | 20 | 31 | 13.64 | 103.9 | 7.64
2009-05-11 | 8:16 | 3.96 | 20 | 31 | 13.6 | 103.6 | 7.65
2009-05-11 | 8:31 | 4.03 | 21 | 32 | 13.56 | 103.6 | 7.65
2009-05-11 | 8:46 | 4.12 | 21 | 32 | 13.51 | 103.4 | 7.65
2009-05-11 | 9:01 | 4.24 | 21 | 32 | 13.53 | 103.9 | 7.66
2009-05-11 | 9:16 | 4.38 | 21 | 32 | 13.51 | 104.1 | 7.66
2009-05-11 | 9:31 | 4.54 | 21 | 32 | 13.42 | 103.8 | 7.67
2009-05-11 | 9:46 | 4.73 | 21 | 32 | 13.39 | 104.1 | 7.67
2009-05-11 | 10:01 | 4.93 | 21 | 32 | 13.28 | 103.8 | 7.67
2009-05-11 | 10:16 | 5.19 | 21 | 32 | 13.24 | 104.2 | 7.67
2009-05-11 | 10:31 | 5.47 | 21 | 32 | 13.11 | 103.9 | 7.68
2009-05-11 | 10:46 | 5.76 | 21 | 32 | 13 | 103.8 | 7.68
2009-05-11 | 11:01 | 6.09 | 21 | 32 | 12.88 | 103.7 | 7.69
2009-05-11 | 11:16 | 6.44 | 21 | 32 | 12.8 | 104 | 7.69
2009-05-11 | 11:31 | 6.84 | 21 | 32 | 12.64 | 103.7 | 7.69
2009-05-11 | 11:46 | 7.24 | 21 | 32 | 12.49 | 103.5 | 7.7
2009-05-11 | 12:01 | 7.66 | 21 | 32 | 12.33 | 103.3 | 7.7
2009-05-11 | 12:16 | 8.08 | 21 | 32 | 12.27 | 103.8 | 7.71
2009-05-11 | 12:31 | 8.46 | 21 | 32 | 12.2 | 104.2 | 7.71
2009-05-11 | 12:46 | 8.87 | 21 | 32 | 12.13 | 104.6 | 7.72
2009-05-11 | 13:00 | 9.24 | 21 | 32 | 11.99 | 104.3 | 7.72
2009-05-11 | 13:15 | 9.62 | 21 | 32 | 11.88 | 104.3 | 7.73
2009-05-11 | 13:30 | 9.93 | 21 | 32 | 11.8 | 104.3 | 7.73
2009-05-11 | 13:45 | 10.23 | 21 | 32 | 11.71 | 104.3 | 7.73
2009-05-11 | 14:00 | 10.54 | 21 | 32 | 11.62 | 104.3 | 7.74
2009-05-11 | 14:15 | 10.8 | 21 | 32 | 11.59 | 104.6 | 7.74
2009-05-11 | 14:30 | 11.01 | 21 | 32 | 11.54 | 104.6 | 7.74
2009-05-11 | 14:45 | 11.12 | 21 | 32 | 11.48 | 104.4 | 7.74
2009-05-11 | 15:00 | 11.21 | 21 | 32 | 11.44 | 104.3 | 7.74
2009-05-11 | 15:15 | 11.33 | 21 | 32 | 11.4 | 104.2 | 7.74
2009-05-11 | 15:30 | 11.5 | 21 | 32 | 11.32 | 103.9 | 7.74
2009-05-11 | 15:45 | 11.62 | 21 | 32 | 11.26 | 103.6 | 7.74
2009-05-11 | 16:00 | 11.66 | 21 | 32 | 11.19 | 103 | 7.74
2009-05-11 | 16:15 | 11.64 | 21 | 32 | 11.14 | 102.5 | 7.74
2009-05-11 | 16:30 | 11.58 | 21 | 32 | 11.11 | 102.1 | 7.74
2009-05-11 | 16:45 | 11.47 | 21 | 32 | 11.1 | 101.7 | 7.74
2009-05-11 | 17:00 | 11.34 | 21 | 32 | 11.16 | 102 | 7.73
2009-05-11 | 17:15 | 11.18 | 21 | 32 | 11.14 | 101.5 | 7.73
2009-05-11 | 17:30 | 10.99 | 21 | 32 | 11.17 | 101.3 | 7.73
2009-05-11 | 17:45 | 10.76 | 21 | 32 | 11.25 | 101.5 | 7.72
2009-05-11 | 18:00 | 10.56 | 21 | 32 | 11.26 | 101.1 | 7.72
2009-05-11 | 18:15 | 10.33 | 21 | 32 | 11.29 | 100.8 | 7.71
2009-05-11 | 18:30 | 10.11 | 21 | 32 | 11.33 | 100.7 | 7.71
2009-05-11 | 18:45 | 9.88 | 21 | 32 | 11.39 | 100.6 | 7.7
2009-05-11 | 19:00 | 9.65 | 21 | 32 | 11.5 | 101 | 7.7
2009-05-11 | 19:15 | 9.45 | 21 | 32 | 11.58 | 101.2 | 7.69
2009-05-11 | 19:30 | 9.28 | 21 | 32 | 11.58 | 100.8 | 7.69
2009-05-11 | 19:45 | 9.13 | 21 | 32 | 11.57 | 100.4 | 7.69
2009-05-11 | 20:00 | 8.98 | 21 | 32 | 11.66 | 100.8 | 7.69
2009-05-11 | 20:15 | 8.85 | 21 | 32 | 11.66 | 100.5 | 7.69
2009-05-11 | 20:30 | 8.73 | 21 | 32 | 11.74 | 100.9 | 7.68
2009-05-11 | 20:45 | 8.59 | 21 | 32 | 11.72 | 100.4 | 7.68
2009-05-11 | 21:00 | 8.46 | 21 | 32 | 11.82 | 100.9 | 7.68
2009-05-11 | 21:15 | 8.35 | 20 | 32 | 11.82 | 100.6 | 7.68
2009-05-11 | 21:30 | 8.21 | 20 | 31 | 11.85 | 100.6 | 7.67
2009-05-11 | 21:45 | 8.05 | 20 | 31 | 11.93 | 100.8 | 7.67
2009-05-11 | 22:00 | 7.89 | 20 | 31 | 11.97 | 100.8 | 7.67
2009-05-11 | 22:15 | 7.72 | 20 | 31 | 12.08 | 101.3 | 7.66
2009-05-11 | 22:30 | 7.56 | 20 | 31 | 12.12 | 101.3 | 7.66
2009-05-11 | 22:45 | 7.38 | 20 | 31 | 12.13 | 100.9 | 7.66
2009-05-11 | 23:00 | 7.19 | 20 | 31 | 12.17 | 100.7 | 7.65
2009-05-11 | 23:15 | 7.04 | 20 | 31 | 12.27 | 101.1 | 7.65
2009-05-11 | 23:30 | 6.87 | 20 | 31 | 12.36 | 101.5 | 7.65
2009-05-11 | 23:45 | 6.72 | 20 | 31 | 12.41 | 101.5 | 7.65
2009-05-12 | 0:00 | 6.57 | 20 | 31 | 12.43 | 101.3 | 7.64
2009-05-12 | 0:15 | 6.43 | 20 | 31 | 12.52 | 101.7 | 7.64
2009-05-12 | 0:30 | 6.31 | 20 | 31 | 12.52 | 101.4 | 7.64
2009-05-12 | 0:45 | 6.17 | 20 | 31 | 12.54 | 101.2 | 7.64
2009-05-12 | 1:01 | 6.03 | 20 | 31 | 12.66 | 101.8 | 7.64
2009-05-12 | 1:16 | 5.87 | 20 | 31 | 12.67 | 101.5 | 7.63
2009-05-12 | 1:31 | 5.7 | 20 | 31 | 12.78 | 101.9 | 7.63
2009-05-12 | 1:46 | 5.52 | 20 | 30 | 12.8 | 101.6 | 7.63
2009-05-12 | 2:01 | 5.32 | 20 | 30 | 12.92 | 102 | 7.62
2009-05-12 | 2:16 | 5.14 | 20 | 30 | 12.95 | 101.8 | 7.62
2009-05-12 | 2:31 | 4.96 | 20 | 30 | 13.07 | 102.3 | 7.62
2009-05-12 | 2:46 | 4.79 | 20 | 30 | 13.1 | 102 | 7.62
2009-05-12 | 3:01 | 4.61 | 20 | 30 | 13.16 | 102 | 7.61
2009-05-12 | 3:16 | 4.46 | 20 | 30 | 13.23 | 102.1 | 7.62
2009-05-12 | 3:31 | 4.32 | 20 | 30 | 13.27 | 102.1 | 7.62
2009-05-12 | 3:46 | 4.19 | 20 | 31 | 13.32 | 102.1 | 7.62
2009-05-12 | 4:01 | 4.1 | 20 | 30 | 13.42 | 102.6 | 7.62
2009-05-12 | 4:16 | 3.99 | 20 | 31 | 13.43 | 102.4 | 7.62
2009-05-12 | 4:31 | 3.9 | 20 | 31 | 13.45 | 102.4 | 7.62
2009-05-12 | 4:46 | 3.82 | 20 | 31 | 13.54 | 102.8 | 7.62
2009-05-12 | 5:01 | 3.74 | 20 | 31 | 13.6 | 103 | 7.62
2009-05-12 | 5:16 | 3.65 | 20 | 31 | 13.59 | 102.7 | 7.62
2009-05-12 | 5:31 | 3.56 | 20 | 31 | 13.6 | 102.6 | 7.62
2009-05-12 | 5:46 | 3.47 | 20 | 31 | 13.68 | 102.9 | 7.62
2009-05-12 | 6:01 | 3.38 | 20 | 31 | 13.76 | 103.3 | 7.62
2009-05-12 | 6:16 | 3.31 | 20 | 31 | 13.75 | 103 | 7.62
2009-05-12 | 6:31 | 3.25 | 20 | 31 | 13.84 | 103.5 | 7.62
2009-05-12 | 6:46 | 3.19 | 20 | 31 | 13.81 | 103.1 | 7.62
2009-05-12 | 7:01 | 3.15 | 20 | 31 | 13.83 | 103.2 | 7.62
2009-05-12 | 7:16 | 3.11 | 20 | 31 | 13.93 | 103.8 | 7.63
2009-05-12 | 7:31 | 3.09 | 20 | 31 | 13.91 | 103.6 | 7.63
2009-05-12 | 7:46 | 3.09 | 20 | 31 | 13.93 | 103.7 | 7.63
2009-05-12 | 8:01 | 3.1 | 20 | 31 | 13.91 | 103.6 | 7.64
2009-05-12 | 8:16 | 3.13 | 20 | 31 | 13.92 | 103.8 | 7.63
2009-05-12 | 8:31 | 3.19 | 20 | 31 | 13.97 | 104.3 | 7.64
2009-05-12 | 8:46 | 3.28 | 20 | 32 | 14.01 | 104.8 | 7.65
2009-05-12 | 9:01 | 3.39 | 21 | 32 | 13.98 | 104.9 | 7.65
2009-05-12 | 9:16 | 3.52 | 21 | 32 | 13.96 | 105.1 | 7.66
2009-05-12 | 9:31 | 3.66 | 21 | 32 | 13.93 | 105.3 | 7.66
2009-05-12 | 9:46 | 3.81 | 21 | 32 | 13.86 | 105.2 | 7.66
2009-05-12 | 10:01 | 3.99 | 21 | 32 | 13.83 | 105.5 | 7.66
2009-05-12 | 10:16 | 4.11 | 21 | 32 | 13.72 | 104.9 | 7.67
2009-05-12 | 10:31 | 4.35 | 21 | 32 | 13.68 | 105.3 | 7.67
2009-05-12 | 10:46 | 4.58 | 21 | 32 | 13.56 | 105 | 7.68
2009-05-12 | 11:01 | 4.86 | 21 | 32 | 13.44 | 104.8 | 7.67
2009-05-12 | 11:16 | 5.17 | 21 | 32 | 13.31 | 104.7 | 7.68
2009-05-12 | 11:31 | 5.49 | 21 | 32 | 13.19 | 104.6 | 7.68
2009-05-12 | 11:46 | 5.87 | 21 | 32 | 13.05 | 104.5 | 7.68
2009-05-12 | 12:01 | 6.23 | 21 | 32 | 12.98 | 104.9 | 7.68