## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited against Sonde AA, Pro-Series (loner), & NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-26 |  | 2009-08-26 |  | A
Time | Computer set to atomic time |  | 9:23 |  | 9:23 |  | A
Temp | VWR NIST traceable surface water |  | 8.8 |  | 8.9 |  | A
DO | In Situ |  | 10.89 |  | 10.85 |  | A
SpC | In Situ |  | 44.4 |  | 47 |  | A
Ph | In Situ |  | 7.62 |  | 8.09 |  | B
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-08-26 10:45 | 12.3 | 8.05 | 10.09 | 94.4 | 71 | 46.2
Sonde | 2009-08-26 10:45 | 12.43 | 8.17 | 10.49 | 98.3 | 72 | 47
Grade |  | A | A | B |  | A
Audit 2 | 2009-09-02 10:01 | 13.2 | 7.97 | 9.83 | 93.7 | 60.4 | 39
Sonde | 2009-09-02 10:00 | 13.22 | 8.11 | 10.3 | 98.3 | 59 | 38
Grade |  | A | A | B |  | A
Audit 3 | 2009-09-11 09:39 | 11.9 | 7.94 | 10.03 | 92.9 | 77.9 | 50.7
Sonde | 2009-09-11 09:37 | 11.91 | 8.18 | 10.75 | 99.6 | 81 | 53
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited against Sonde AA, Pro-Series (loner), & NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-09-11 |  | 2009-09-11 |  | A
Time | Computer set to atomic time |  | 10:53 |  | 10:55 |  | A
Temp | VWR NIST traceable surface water |  | 16.6 |  | 16.49 |  | A
DO | In Situ |  | 8.79 |  | 10.59 |  | Fail
SpC | In Situ |  | 70 |  | 69 |  | A
Ph | In Situ |  | 8.17 |  | 8.34 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
DO data can be considered acceptable because it passed the final audit of deployment within hours of failing the final post sampling audit. Also, no DO standards were violated |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-24 |  | 2009-08-24 |  | A
Time | Computer set to atomic time |  | 12:47 |  | 12:47 |  | A
Temp | VWR NIST traceable surface water |  | 13.9 |  | 14 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 9.86 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 146.4 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.16 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-08-26 10:45 | 12.3 | 8.05 | 10.09 | 94.4 | 71 | 46.2
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-26 |  | 2009-08-26 |  | A
Time | Computer set to atomic time |  | 11:50 |  | 11:50 |  | A
Temp | VWR NIST traceable surface water |  | 18.9 |  | 18.7 |  | A
DO | Winkler titration of surface water |  | 9.4 |  | 9.19 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 138.7 |  | B
Ph | pH 7 standard solution |  | 7.02 |  | 7.18 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-26 |  | 2009-08-26 |  | A
Time | Computer set to atomic time |  | 11:50 |  | 11:50 |  | A
Temp | VWR NIST traceable surface water |  | 18.9 |  | 18.7 |  | A
DO | Winkler titration of surface water |  | 9.4 |  | 9.19 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 138.7 |  | B
Ph | pH 7 standard solution |  | 7.02 |  | 7.18 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-09-02 10:01 | 13.2 | 7.97 | 9.83 | 93.7 | 60.4 | 39
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-09-02 |  | 2009-09-02 |  | A
Time | Computer set to atomic time |  | 12:00 |  | 12:00 |  | A
Temp | VWR NIST traceable surface water |  | 13.19 |  | 13.22 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 9.83 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 152.5 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.18 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-09-08 |  | 2009-09-08 |  | A
Time | Computer set to atomic time |  | 14:00 |  | 14:00 |  | A
Temp | VWR NIST traceable surface water |  | 7.8 |  | 7.5 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.43 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 171.4 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-09-11 09:39 | 11.9 | 7.94 | 10.03 | 92.9 | 77.9 | 50.7
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-09-11 |  | 2009-09-11 |  | A
Time | Computer set to atomic time |  | 14:00 |  | 14:00 |  | A
Temp | VWR NIST traceable surface water |  | 16.6 |  | 16.4 |  | A
DO | Winkler titration of surface water |  | 9.6 |  | 8.79 |  | B
SpC | 148 µS/mL |  | 156.7 |  | 111.1 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.18 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
On the Post-Sampling Audit, SpC Can be trusted (was working in accordance with historic station record in field) |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-08-26 | 10:45 | 12.43 | 47 | 72 | 10.49 | 98.3 | 8.17
2009-08-26 | 11:01 | 12.58 | 47 | 71 | 10.98 | 103.2 | 8.21
2009-08-26 | 11:16 | 12.75 | 46 | 70 | 10.82 | 102.1 | 8.21
2009-08-26 | 11:31 | 12.97 | 45 | 68 | 10.62 | 100.8 | 8.23
2009-08-26 | 11:46 | 13.28 | 44 | 67 | 10.49 | 100.2 | 8.24
2009-08-26 | 12:01 | 13.64 | 44 | 66 | 10.35 | 99.7 | 8.25
2009-08-26 | 12:16 | 14.08 | 43 | 66 | 10.17 | 98.9 | 8.26
2009-08-26 | 12:31 | 14.55 | 43 | 65 | 10.01 | 98.4 | 8.27
2009-08-26 | 12:46 | 15.05 | 42 | 64 | 9.91 | 98.4 | 8.28
2009-08-26 | 13:01 | 15.58 | 42 | 64 | 9.79 | 98.3 | 8.3
2009-08-26 | 13:16 | 16.14 | 41 | 63 | 9.65 | 98.1 | 8.32
2009-08-26 | 13:31 | 16.68 | 41 | 63 | 9.53 | 98 | 8.34
2009-08-26 | 13:46 | 17.24 | 41 | 62 | 9.42 | 97.9 | 8.34
2009-08-26 | 14:01 | 17.8 | 41 | 62 | 9.26 | 97.4 | 8.36
2009-08-26 | 14:16 | 18.37 | 40 | 62 | 9.13 | 97.1 | 8.38
2009-08-26 | 14:31 | 18.92 | 40 | 62 | 9.03 | 97.2 | 8.39
2009-08-26 | 14:46 | 19.39 | 40 | 62 | 8.9 | 96.8 | 8.4
2009-08-26 | 15:01 | 19.82 | 40 | 62 | 8.83 | 96.8 | 8.41
2009-08-26 | 15:16 | 20.18 | 40 | 62 | 8.73 | 96.3 | 8.42
2009-08-26 | 15:31 | 20.51 | 40 | 62 | 8.64 | 96 | 8.42
2009-08-26 | 15:46 | 20.82 | 40 | 62 | 8.57 | 95.7 | 8.42
2009-08-26 | 16:01 | 21.03 | 40 | 62 | 8.51 | 95.5 | 8.42
2009-08-26 | 16:16 | 21.15 | 40 | 62 | 8.46 | 95.2 | 8.41
2009-08-26 | 16:31 | 21.18 | 40 | 62 | 8.45 | 95.2 | 8.39
2009-08-26 | 16:46 | 21.12 | 40 | 62 | 8.44 | 94.9 | 8.37
2009-08-26 | 17:01 | 21.01 | 40 | 62 | 8.43 | 94.6 | 8.35
2009-08-26 | 17:16 | 20.85 | 40 | 62 | 8.4 | 93.9 | 8.32
2009-08-26 | 17:31 | 20.7 | 40 | 62 | 8.41 | 93.8 | 8.29
2009-08-26 | 17:46 | 20.57 | 40 | 62 | 8.4 | 93.5 | 8.25
2009-08-26 | 18:01 | 20.39 | 40 | 62 | 8.42 | 93.3 | 8.22
2009-08-26 | 18:16 | 20.25 | 40 | 62 | 8.39 | 92.7 | 8.19
2009-08-26 | 18:31 | 20.12 | 40 | 62 | 8.38 | 92.4 | 8.17
2009-08-26 | 18:46 | 20.02 | 40 | 62 | 8.4 | 92.5 | 8.15
2009-08-26 | 19:01 | 19.92 | 41 | 62 | 8.39 | 92.1 | 8.13
2009-08-26 | 19:16 | 19.84 | 41 | 62 | 8.4 | 92.1 | 8.12
2009-08-26 | 19:31 | 19.78 | 41 | 63 | 8.41 | 92.1 | 8.1
2009-08-26 | 19:46 | 19.72 | 41 | 63 | 8.4 | 91.9 | 8.09
2009-08-26 | 20:01 | 19.64 | 41 | 63 | 8.37 | 91.4 | 8.08
2009-08-26 | 20:16 | 19.58 | 41 | 63 | 8.39 | 91.5 | 8.07
2009-08-26 | 20:31 | 19.5 | 41 | 63 | 8.4 | 91.5 | 8.07
2009-08-26 | 20:46 | 19.42 | 41 | 63 | 8.41 | 91.5 | 8.06
2009-08-26 | 21:01 | 19.35 | 41 | 63 | 8.42 | 91.4 | 8.06
2009-08-26 | 21:16 | 19.25 | 41 | 63 | 8.44 | 91.5 | 8.05
2009-08-26 | 21:31 | 19.14 | 41 | 63 | 8.46 | 91.5 | 8.05
2009-08-26 | 21:46 | 19.02 | 41 | 63 | 8.49 | 91.6 | 8.05
2009-08-26 | 22:01 | 18.88 | 41 | 64 | 8.5 | 91.4 | 8.05
2009-08-26 | 22:16 | 18.73 | 41 | 64 | 8.52 | 91.3 | 8.04
2009-08-26 | 22:31 | 18.56 | 41 | 64 | 8.57 | 91.6 | 8.04
2009-08-26 | 22:46 | 18.38 | 42 | 64 | 8.58 | 91.4 | 8.04
2009-08-26 | 23:01 | 18.21 | 42 | 64 | 8.64 | 91.7 | 8.03
2009-08-26 | 23:16 | 18.04 | 42 | 64 | 8.65 | 91.5 | 8.03
2009-08-26 | 23:31 | 17.87 | 42 | 64 | 8.68 | 91.5 | 8.03
2009-08-26 | 23:46 | 17.69 | 42 | 65 | 8.74 | 91.8 | 8.03
2009-08-27 | 0:01 | 17.51 | 42 | 65 | 8.79 | 92 | 8.04
2009-08-27 | 0:16 | 17.33 | 42 | 65 | 8.8 | 91.7 | 8.04
2009-08-27 | 0:31 | 17.15 | 42 | 65 | 8.86 | 92 | 8.04
2009-08-27 | 0:46 | 16.96 | 42 | 65 | 8.88 | 91.8 | 8.04
2009-08-27 | 1:01 | 16.77 | 42 | 65 | 8.93 | 92 | 8.04
2009-08-27 | 1:16 | 16.6 | 42 | 65 | 8.98 | 92.1 | 8.04
2009-08-27 | 1:31 | 16.42 | 42 | 65 | 8.99 | 91.9 | 8.04
2009-08-27 | 1:46 | 16.23 | 42 | 65 | 9.03 | 91.9 | 8.04
2009-08-27 | 2:01 | 16.05 | 42 | 65 | 9.09 | 92.2 | 8.04
2009-08-27 | 2:16 | 15.88 | 42 | 65 | 9.1 | 92 | 8.04
2009-08-27 | 2:31 | 15.71 | 42 | 65 | 9.13 | 91.9 | 8.03
2009-08-27 | 2:46 | 15.54 | 43 | 66 | 9.19 | 92.2 | 8.04
2009-08-27 | 3:01 | 15.38 | 43 | 66 | 9.24 | 92.4 | 8.04
2009-08-27 | 3:16 | 15.23 | 43 | 66 | 9.28 | 92.5 | 8.03
2009-08-27 | 3:31 | 15.08 | 43 | 66 | 9.29 | 92.3 | 8.04
2009-08-27 | 3:46 | 14.94 | 43 | 66 | 9.34 | 92.6 | 8.04
2009-08-27 | 4:01 | 14.79 | 43 | 66 | 9.39 | 92.7 | 8.04
2009-08-27 | 4:16 | 14.65 | 43 | 66 | 9.41 | 92.6 | 8.04
2009-08-27 | 4:31 | 14.51 | 43 | 67 | 9.46 | 92.8 | 8.05
2009-08-27 | 4:46 | 14.38 | 43 | 67 | 9.48 | 92.7 | 8.05
2009-08-27 | 5:01 | 14.25 | 43 | 67 | 9.53 | 93 | 8.05
2009-08-27 | 5:16 | 14.13 | 44 | 67 | 9.53 | 92.8 | 8.05
2009-08-27 | 5:31 | 14 | 44 | 68 | 9.59 | 93.1 | 8.05
2009-08-27 | 5:46 | 13.9 | 44 | 68 | 9.6 | 93 | 8.06
2009-08-27 | 6:01 | 13.81 | 44 | 69 | 9.64 | 93.2 | 8.06
2009-08-27 | 6:16 | 13.72 | 45 | 69 | 9.67 | 93.2 | 8.07
2009-08-27 | 6:31 | 13.63 | 45 | 70 | 9.66 | 93 | 8.07
2009-08-27 | 6:46 | 13.53 | 45 | 70 | 9.72 | 93.3 | 8.08
2009-08-27 | 7:01 | 13.43 | 46 | 71 | 9.77 | 93.6 | 8.08
2009-08-27 | 7:16 | 13.31 | 46 | 71 | 9.81 | 93.8 | 8.1
2009-08-27 | 7:31 | 13.2 | 46 | 72 | 9.86 | 94 | 8.1
2009-08-27 | 7:46 | 13.1 | 47 | 72 | 9.87 | 93.9 | 8.11
2009-08-27 | 8:01 | 13.01 | 47 | 73 | 9.92 | 94.2 | 8.12
2009-08-27 | 8:16 | 12.94 | 47 | 73 | 9.97 | 94.5 | 8.13
2009-08-27 | 8:31 | 12.88 | 47 | 73 | 10 | 94.7 | 8.14
2009-08-27 | 8:46 | 12.83 | 47 | 73 | 10 | 94.5 | 8.14
2009-08-27 | 9:01 | 12.83 | 47 | 72 | 10.05 | 95 | 8.15
2009-08-27 | 9:16 | 12.83 | 47 | 71 | 10.04 | 95 | 8.16
2009-08-27 | 9:31 | 12.85 | 46 | 70 | 10.08 | 95.4 | 8.16
2009-08-27 | 9:46 | 12.9 | 46 | 69 | 10.11 | 95.7 | 8.16
2009-08-27 | 10:01 | 12.98 | 45 | 68 | 10.09 | 95.7 | 8.17
2009-08-27 | 10:16 | 13.05 | 44 | 66 | 10.08 | 95.8 | 8.18
2009-08-27 | 10:31 | 13.13 | 43 | 65 | 10.07 | 95.9 | 8.18
2009-08-27 | 10:46 | 13.24 | 42 | 64 | 10.07 | 96.1 | 8.18
2009-08-27 | 11:01 | 13.38 | 42 | 63 | 10.06 | 96.3 | 8.18
2009-08-27 | 11:16 | 13.56 | 41 | 63 | 10.04 | 96.5 | 8.19
2009-08-27 | 11:31 | 13.78 | 41 | 62 | 10.04 | 96.9 | 8.19
2009-08-27 | 11:46 | 14.08 | 40 | 61 | 10 | 97.2 | 8.21
2009-08-27 | 12:01 | 14.45 | 40 | 61 | 9.94 | 97.4 | 8.21
2009-08-27 | 12:16 | 14.87 | 40 | 61 | 9.84 | 97.4 | 8.23
2009-08-27 | 12:31 | 15.32 | 39 | 60 | 9.75 | 97.4 | 8.24
2009-08-27 | 12:46 | 15.79 | 39 | 60 | 9.63 | 97.1 | 8.25
2009-08-27 | 13:01 | 16.29 | 39 | 60 | 9.51 | 96.9 | 8.27
2009-08-27 | 13:16 | 16.82 | 39 | 60 | 9.4 | 96.9 | 8.29
2009-08-27 | 13:31 | 17.35 | 39 | 60 | 9.29 | 96.8 | 8.29
2009-08-27 | 13:46 | 17.87 | 39 | 59 | 9.19 | 96.8 | 8.33
2009-08-27 | 14:01 | 18.41 | 39 | 59 | 9.11 | 97.1 | 8.33
2009-08-27 | 14:16 | 18.94 | 39 | 59 | 9.01 | 97 | 8.34
2009-08-27 | 14:31 | 19.43 | 39 | 59 | 8.89 | 96.6 | 8.36
2009-08-27 | 14:46 | 19.88 | 39 | 60 | 8.82 | 96.7 | 8.38
2009-08-27 | 15:01 | 20.25 | 39 | 60 | 8.74 | 96.6 | 8.38
2009-08-27 | 15:16 | 20.58 | 39 | 60 | 8.64 | 96.1 | 8.39
2009-08-27 | 15:31 | 20.88 | 39 | 60 | 8.57 | 95.9 | 8.4
2009-08-27 | 15:46 | 21.12 | 39 | 60 | 8.51 | 95.7 | 8.4
2009-08-27 | 16:01 | 21.32 | 39 | 60 | 8.46 | 95.5 | 8.4
2009-08-27 | 16:16 | 21.4 | 39 | 60 | 8.45 | 95.5 | 8.39
2009-08-27 | 16:31 | 21.4 | 39 | 60 | 8.43 | 95.3 | 8.37
2009-08-27 | 16:46 | 21.31 | 39 | 60 | 8.39 | 94.7 | 8.35
2009-08-27 | 17:01 | 21.2 | 39 | 60 | 8.37 | 94.3 | 8.32
2009-08-27 | 17:16 | 21.05 | 39 | 60 | 8.36 | 93.9 | 8.28
2009-08-27 | 17:31 | 20.86 | 39 | 60 | 8.39 | 93.8 | 8.25
2009-08-27 | 17:46 | 20.73 | 39 | 60 | 8.39 | 93.7 | 8.22
2009-08-27 | 18:01 | 20.55 | 39 | 60 | 8.38 | 93.2 | 8.19
2009-08-27 | 18:16 | 20.41 | 39 | 61 | 8.4 | 93.1 | 8.16
2009-08-27 | 18:31 | 20.27 | 39 | 61 | 8.38 | 92.7 | 8.14
2009-08-27 | 18:46 | 20.17 | 39 | 61 | 8.41 | 92.8 | 8.12
2009-08-27 | 19:01 | 20.09 | 40 | 61 | 8.41 | 92.7 | 8.11
2009-08-27 | 19:16 | 20.01 | 40 | 61 | 8.42 | 92.7 | 8.1
2009-08-27 | 19:31 | 19.97 | 40 | 61 | 8.42 | 92.5 | 8.08
2009-08-27 | 19:46 | 19.92 | 40 | 61 | 8.38 | 92 | 8.06
2009-08-27 | 20:01 | 19.85 | 40 | 61 | 8.37 | 91.8 | 8.05
2009-08-27 | 20:16 | 19.74 | 40 | 61 | 8.39 | 91.8 | 8.04
2009-08-27 | 20:31 | 19.63 | 40 | 61 | 8.41 | 91.8 | 8.03
2009-08-27 | 20:46 | 19.52 | 40 | 61 | 8.43 | 91.8 | 8.02
2009-08-27 | 21:01 | 19.39 | 40 | 62 | 8.44 | 91.6 | 8.01
2009-08-27 | 21:16 | 19.24 | 40 | 62 | 8.45 | 91.6 | 8.01
2009-08-27 | 21:31 | 19.08 | 40 | 62 | 8.5 | 91.8 | 8.01
2009-08-27 | 21:46 | 18.91 | 40 | 62 | 8.54 | 91.9 | 8.01
2009-08-27 | 22:01 | 18.73 | 40 | 62 | 8.59 | 92.1 | 8.01
2009-08-27 | 22:16 | 18.53 | 40 | 62 | 8.61 | 91.9 | 8
2009-08-27 | 22:31 | 18.35 | 40 | 62 | 8.63 | 91.9 | 8
2009-08-27 | 22:46 | 18.17 | 40 | 62 | 8.66 | 91.8 | 8
2009-08-27 | 23:01 | 17.98 | 41 | 63 | 8.73 | 92.1 | 8
2009-08-27 | 23:16 | 17.8 | 41 | 63 | 8.77 | 92.2 | 8
2009-08-27 | 23:31 | 17.62 | 41 | 63 | 8.81 | 92.3 | 8
2009-08-27 | 23:46 | 17.43 | 41 | 63 | 8.83 | 92.2 | 8
2009-08-28 | 0:01 | 17.24 | 41 | 63 | 8.86 | 92.1 | 8
2009-08-28 | 0:16 | 17.07 | 41 | 63 | 8.91 | 92.3 | 8
2009-08-28 | 0:31 | 16.89 | 41 | 63 | 8.93 | 92.2 | 8
2009-08-28 | 0:46 | 16.72 | 41 | 63 | 8.99 | 92.5 | 8
2009-08-28 | 1:01 | 16.56 | 41 | 63 | 9.03 | 92.6 | 8
2009-08-28 | 1:16 | 16.4 | 41 | 63 | 9.07 | 92.7 | 8
2009-08-28 | 1:31 | 16.24 | 41 | 63 | 9.11 | 92.8 | 8
2009-08-28 | 1:46 | 16.09 | 41 | 63 | 9.11 | 92.5 | 8
2009-08-28 | 2:01 | 15.94 | 41 | 64 | 9.13 | 92.4 | 8
2009-08-28 | 2:16 | 15.8 | 41 | 64 | 9.19 | 92.7 | 8
2009-08-28 | 2:31 | 15.67 | 41 | 64 | 9.2 | 92.5 | 8
2009-08-28 | 2:46 | 15.54 | 41 | 64 | 9.22 | 92.5 | 8
2009-08-28 | 3:01 | 15.42 | 41 | 64 | 9.26 | 92.7 | 8.01
2009-08-28 | 3:16 | 15.28 | 42 | 64 | 9.28 | 92.6 | 8
2009-08-28 | 3:31 | 15.15 | 42 | 64 | 9.33 | 92.8 | 8.01
2009-08-28 | 3:46 | 15.02 | 42 | 64 | 9.34 | 92.7 | 8
2009-08-28 | 4:01 | 14.89 | 42 | 64 | 9.4 | 93 | 8.01
2009-08-28 | 4:16 | 14.77 | 42 | 65 | 9.42 | 93 | 8.01
2009-08-28 | 4:31 | 14.65 | 42 | 65 | 9.45 | 93 | 8.02
2009-08-28 | 4:46 | 14.54 | 42 | 65 | 9.47 | 93 | 8.02
2009-08-28 | 5:01 | 14.42 | 42 | 66 | 9.53 | 93.3 | 8.02
2009-08-28 | 5:16 | 14.31 | 43 | 66 | 9.56 | 93.4 | 8.02
2009-08-28 | 5:31 | 14.2 | 43 | 66 | 9.59 | 93.5 | 8.02
2009-08-28 | 5:46 | 14.1 | 43 | 67 | 9.59 | 93.3 | 8.03
2009-08-28 | 6:01 | 14.01 | 43 | 67 | 9.6 | 93.2 | 8.03
2009-08-28 | 6:16 | 13.91 | 44 | 68 | 9.65 | 93.5 | 8.04
2009-08-28 | 6:31 | 13.82 | 44 | 68 | 9.66 | 93.4 | 8.04
2009-08-28 | 6:46 | 13.73 | 45 | 69 | 9.68 | 93.4 | 8.05
2009-08-28 | 7:01 | 13.64 | 45 | 70 | 9.74 | 93.8 | 8.06
2009-08-28 | 7:16 | 13.55 | 45 | 70 | 9.8 | 94.1 | 8.07
2009-08-28 | 7:31 | 13.46 | 46 | 71 | 9.85 | 94.4 | 8.08
2009-08-28 | 7:46 | 13.38 | 46 | 71 | 9.89 | 94.7 | 8.1
2009-08-28 | 8:01 | 13.32 | 46 | 72 | 9.94 | 95.1 | 8.11
2009-08-28 | 8:16 | 13.27 | 47 | 72 | 9.96 | 95.1 | 8.14
2009-08-28 | 8:31 | 13.24 | 47 | 72 | 10.02 | 95.6 | 8.16
2009-08-28 | 8:46 | 13.22 | 47 | 72 | 10.06 | 96 | 8.17
2009-08-28 | 9:01 | 13.22 | 47 | 72 | 10.06 | 95.9 | 8.19
2009-08-28 | 9:16 | 13.24 | 47 | 71 | 10.1 | 96.4 | 8.21
2009-08-28 | 9:31 | 13.28 | 46 | 70 | 10.08 | 96.3 | 8.22
2009-08-28 | 9:46 | 13.34 | 46 | 69 | 10.12 | 96.8 | 8.23
2009-08-28 | 10:01 | 13.44 | 45 | 68 | 10.13 | 97.1 | 8.25
2009-08-28 | 10:16 | 13.55 | 44 | 66 | 10.13 | 97.3 | 8.26
2009-08-28 | 10:31 | 13.69 | 43 | 65 | 10.08 | 97.2 | 8.27
2009-08-28 | 10:46 | 13.84 | 42 | 64 | 10.08 | 97.5 | 8.28
2009-08-28 | 11:01 | 14 | 42 | 63 | 10.07 | 97.7 | 8.29
2009-08-28 | 11:16 | 14.17 | 41 | 62 | 10.01 | 97.5 | 8.3
2009-08-28 | 11:31 | 14.32 | 41 | 62 | 9.98 | 97.5 | 8.29
2009-08-28 | 11:46 | 14.46 | 40 | 61 | 9.95 | 97.5 | 8.29
2009-08-28 | 12:01 | 14.62 | 40 | 61 | 9.9 | 97.4 | 8.3
2009-08-28 | 12:16 | 14.78 | 39 | 60 | 9.88 | 97.6 | 8.3
2009-08-28 | 12:31 | 14.96 | 39 | 60 | 9.82 | 97.4 | 8.32
2009-08-28 | 12:46 | 15.14 | 39 | 60 | 9.81 | 97.6 | 8.32
2009-08-28 | 13:01 | 15.3 | 39 | 59 | 9.75 | 97.3 | 8.33
2009-08-28 | 13:16 | 15.45 | 39 | 59 | 9.7 | 97.1 | 8.33
2009-08-28 | 13:31 | 15.57 | 38 | 59 | 9.65 | 96.9 | 8.33
2009-08-28 | 13:46 | 15.69 | 38 | 59 | 9.63 | 96.9 | 8.33
2009-08-28 | 14:01 | 15.82 | 38 | 59 | 9.63 | 97.2 | 8.34
2009-08-28 | 14:16 | 15.89 | 38 | 59 | 9.55 | 96.5 | 8.33
2009-08-28 | 14:31 | 15.97 | 38 | 59 | 9.54 | 96.5 | 8.32
2009-08-28 | 14:46 | 16.09 | 38 | 59 | 9.56 | 97 | 8.32
2009-08-28 | 15:01 | 16.19 | 38 | 58 | 9.53 | 97 | 8.34
2009-08-28 | 15:16 | 16.27 | 38 | 58 | 9.53 | 97.1 | 8.34
2009-08-28 | 15:31 | 16.43 | 38 | 58 | 9.49 | 97 | 8.34
2009-08-28 | 15:46 | 16.57 | 38 | 58 | 9.49 | 97.3 | 8.36
2009-08-28 | 16:01 | 16.69 | 38 | 59 | 9.49 | 97.5 | 8.36
2009-08-28 | 16:16 | 16.76 | 38 | 59 | 9.42 | 97 | 8.37
2009-08-28 | 16:31 | 16.87 | 38 | 59 | 9.38 | 96.8 | 8.37
2009-08-28 | 16:46 | 16.89 | 38 | 59 | 9.37 | 96.8 | 8.36
2009-08-28 | 17:01 | 16.88 | 38 | 59 | 9.36 | 96.6 | 8.34
2009-08-28 | 17:16 | 16.85 | 38 | 59 | 9.27 | 95.7 | 8.3
2009-08-28 | 17:31 | 16.8 | 38 | 59 | 9.23 | 95.1 | 8.25
2009-08-28 | 17:46 | 16.73 | 38 | 59 | 9.19 | 94.5 | 8.19
2009-08-28 | 18:01 | 16.7 | 39 | 59 | 9.13 | 93.9 | 8.15
2009-08-28 | 18:16 | 16.65 | 39 | 59 | 9.15 | 94 | 8.12
2009-08-28 | 18:31 | 16.62 | 39 | 59 | 9.16 | 94.1 | 8.11
2009-08-28 | 18:46 | 16.61 | 39 | 60 | 9.18 | 94.2 | 8.11
2009-08-28 | 19:01 | 16.61 | 39 | 60 | 9.16 | 94 | 8.1
2009-08-28 | 19:16 | 16.6 | 39 | 60 | 9.14 | 93.8 | 8.1
2009-08-28 | 19:31 | 16.56 | 39 | 60 | 9.11 | 93.4 | 8.08
2009-08-28 | 19:46 | 16.52 | 39 | 60 | 9.12 | 93.4 | 8.06
2009-08-28 | 20:01 | 16.46 | 39 | 60 | 9.11 | 93.2 | 8.04
2009-08-28 | 20:16 | 16.38 | 39 | 60 | 9.1 | 92.9 | 8.02
2009-08-28 | 20:31 | 16.28 | 39 | 61 | 9.12 | 93 | 8.01
2009-08-28 | 20:46 | 16.17 | 39 | 61 | 9.13 | 92.8 | 8
2009-08-28 | 21:01 | 16.05 | 39 | 61 | 9.14 | 92.7 | 8
2009-08-28 | 21:16 | 15.92 | 40 | 61 | 9.19 | 93 | 7.99
2009-08-28 | 21:31 | 15.8 | 40 | 61 | 9.21 | 92.9 | 7.99
2009-08-28 | 21:46 | 15.68 | 40 | 61 | 9.25 | 93.1 | 7.99
2009-08-28 | 22:01 | 15.57 | 40 | 61 | 9.28 | 93.2 | 7.98
2009-08-28 | 22:16 | 15.45 | 40 | 62 | 9.31 | 93.2 | 7.98
2009-08-28 | 22:31 | 15.34 | 40 | 62 | 9.34 | 93.3 | 7.98
2009-08-28 | 22:46 | 15.23 | 40 | 62 | 9.36 | 93.3 | 7.98
2009-08-28 | 23:01 | 15.12 | 40 | 62 | 9.39 | 93.4 | 7.99
2009-08-28 | 23:16 | 15.01 | 40 | 62 | 9.42 | 93.4 | 7.98
2009-08-28 | 23:31 | 14.89 | 40 | 62 | 9.42 | 93.2 | 7.98
2009-08-28 | 23:46 | 14.78 | 40 | 62 | 9.46 | 93.3 | 7.98
2009-08-29 | 0:01 | 14.67 | 40 | 62 | 9.49 | 93.4 | 7.98
2009-08-29 | 0:16 | 14.57 | 40 | 62 | 9.49 | 93.2 | 7.98
2009-08-29 | 0:31 | 14.46 | 40 | 62 | 9.53 | 93.4 | 7.99
2009-08-29 | 0:46 | 14.35 | 41 | 62 | 9.56 | 93.5 | 7.98
2009-08-29 | 1:01 | 14.24 | 41 | 63 | 9.56 | 93.3 | 7.98
2009-08-29 | 1:16 | 14.14 | 41 | 63 | 9.61 | 93.5 | 7.98
2009-08-29 | 1:31 | 14.04 | 41 | 63 | 9.64 | 93.6 | 7.98
2009-08-29 | 1:46 | 13.96 | 41 | 63 | 9.66 | 93.7 | 7.98
2009-08-29 | 2:01 | 13.88 | 41 | 63 | 9.65 | 93.4 | 7.99
2009-08-29 | 2:16 | 13.79 | 41 | 63 | 9.66 | 93.3 | 7.99
2009-08-29 | 2:31 | 13.71 | 41 | 63 | 9.71 | 93.6 | 7.99
2009-08-29 | 2:46 | 13.63 | 41 | 63 | 9.74 | 93.7 | 7.99
2009-08-29 | 3:01 | 13.54 | 41 | 63 | 9.76 | 93.8 | 7.99
2009-08-29 | 3:16 | 13.44 | 41 | 64 | 9.76 | 93.5 | 7.99
2009-08-29 | 3:31 | 13.34 | 41 | 64 | 9.8 | 93.7 | 7.99
2009-08-29 | 3:46 | 13.24 | 41 | 64 | 9.83 | 93.8 | 7.99
2009-08-29 | 4:01 | 13.14 | 42 | 64 | 9.82 | 93.5 | 7.99
2009-08-29 | 4:16 | 13.04 | 42 | 64 | 9.87 | 93.8 | 7.99
2009-08-29 | 4:31 | 12.94 | 42 | 64 | 9.91 | 93.9 | 7.99
2009-08-29 | 4:46 | 12.84 | 42 | 65 | 9.91 | 93.7 | 8
2009-08-29 | 5:01 | 12.75 | 42 | 65 | 9.95 | 93.9 | 8
2009-08-29 | 5:16 | 12.65 | 42 | 65 | 9.99 | 94.1 | 8
2009-08-29 | 5:31 | 12.57 | 42 | 65 | 10.02 | 94.1 | 8
2009-08-29 | 5:46 | 12.48 | 42 | 65 | 10.04 | 94.2 | 8
2009-08-29 | 6:01 | 12.4 | 43 | 66 | 10.07 | 94.3 | 8
2009-08-29 | 6:16 | 12.33 | 43 | 66 | 10.09 | 94.3 | 8
2009-08-29 | 6:31 | 12.28 | 43 | 66 | 10.07 | 94.1 | 8
2009-08-29 | 6:46 | 12.23 | 43 | 66 | 10.12 | 94.4 | 8.01
2009-08-29 | 7:01 | 12.19 | 43 | 66 | 10.16 | 94.7 | 8.02
2009-08-29 | 7:16 | 12.16 | 43 | 67 | 10.16 | 94.7 | 8.03
2009-08-29 | 7:31 | 12.12 | 43 | 67 | 10.18 | 94.7 | 8.04
2009-08-29 | 7:46 | 12.08 | 44 | 67 | 10.22 | 95.1 | 8.05
2009-08-29 | 8:01 | 12.05 | 44 | 68 | 10.25 | 95.3 | 8.06
2009-08-29 | 8:16 | 12.03 | 44 | 68 | 10.24 | 95.1 | 8.07
2009-08-29 | 8:31 | 12.02 | 44 | 68 | 10.25 | 95.1 | 8.08
2009-08-29 | 8:46 | 12.03 | 44 | 69 | 10.29 | 95.5 | 8.08
2009-08-29 | 9:01 | 12.07 | 45 | 69 | 10.28 | 95.5 | 8.09
2009-08-29 | 9:16 | 12.12 | 45 | 69 | 10.28 | 95.6 | 8.11
2009-08-29 | 9:31 | 12.17 | 45 | 69 | 10.31 | 96.1 | 8.12
2009-08-29 | 9:46 | 12.26 | 45 | 70 | 10.33 | 96.5 | 8.14
2009-08-29 | 10:01 | 12.35 | 45 | 70 | 10.31 | 96.5 | 8.15
2009-08-29 | 10:16 | 12.44 | 46 | 70 | 10.33 | 96.8 | 8.16
2009-08-29 | 10:31 | 12.51 | 46 | 70 | 10.34 | 97.1 | 8.17
2009-08-29 | 10:46 | 12.63 | 46 | 71 | 10.31 | 97.1 | 8.19
2009-08-29 | 11:01 | 12.78 | 46 | 71 | 10.33 | 97.5 | 8.2
2009-08-29 | 11:16 | 12.88 | 46 | 70 | 10.32 | 97.7 | 8.2
2009-08-29 | 11:31 | 13.21 | 46 | 70 | 10.29 | 98.1 | 8.23
2009-08-29 | 11:46 | 13.29 | 45 | 70 | 10.17 | 97.2 | 8.23
2009-08-29 | 12:01 | 13.51 | 45 | 69 | 10.15 | 97.4 | 8.25
2009-08-29 | 12:16 | 13.86 | 45 | 69 | 10.1 | 97.7 | 8.26
2009-08-29 | 12:31 | 14.33 | 45 | 68 | 10.02 | 97.9 | 8.28
2009-08-29 | 12:46 | 14.8 | 44 | 68 | 9.9 | 97.8 | 8.3
2009-08-29 | 13:01 | 15.26 | 44 | 68 | 9.75 | 97.3 | 8.31
2009-08-29 | 13:16 | 15.6 | 44 | 67 | 9.67 | 97.1 | 8.3
2009-08-29 | 13:31 | 15.97 | 44 | 67 | 9.64 | 97.6 | 8.31
2009-08-29 | 13:46 | 16.46 | 44 | 67 | 9.56 | 97.8 | 8.33
2009-08-29 | 14:01 | 16.96 | 43 | 66 | 9.44 | 97.6 | 8.35
2009-08-29 | 14:16 | 17.65 | 43 | 66 | 9.32 | 97.8 | 8.38
2009-08-29 | 14:31 | 18.28 | 43 | 65 | 9.19 | 97.6 | 8.39
2009-08-29 | 14:46 | 18.72 | 42 | 65 | 9.09 | 97.4 | 8.4
2009-08-29 | 15:01 | 19.03 | 42 | 64 | 8.99 | 97 | 8.4
2009-08-29 | 15:16 | 19.26 | 42 | 64 | 8.96 | 97.1 | 8.4
2009-08-29 | 15:31 | 19.4 | 42 | 63 | 8.89 | 96.6 | 8.39
2009-08-29 | 15:46 | 19.56 | 41 | 63 | 8.85 | 96.5 | 8.39
2009-08-29 | 16:01 | 19.72 | 41 | 62 | 8.85 | 96.8 | 8.39
2009-08-29 | 16:16 | 19.8 | 41 | 62 | 8.82 | 96.6 | 8.39
2009-08-29 | 16:31 | 19.72 | 40 | 62 | 8.8 | 96.3 | 8.38
2009-08-29 | 16:46 | 19.5 | 40 | 62 | 8.8 | 95.9 | 8.36
2009-08-29 | 17:01 | 19.28 | 40 | 61 | 8.77 | 95.1 | 8.32
2009-08-29 | 17:16 | 19.14 | 40 | 61 | 8.77 | 94.8 | 8.28
2009-08-29 | 17:31 | 19.02 | 40 | 61 | 8.76 | 94.4 | 8.24
2009-08-29 | 17:46 | 18.87 | 40 | 61 | 8.76 | 94.2 | 8.21
2009-08-29 | 18:01 | 18.76 | 40 | 61 | 8.77 | 94 | 8.18
2009-08-29 | 18:16 | 18.66 | 40 | 61 | 8.78 | 94 | 8.17
2009-08-29 | 18:31 | 18.57 | 40 | 61 | 8.8 | 94 | 8.16
2009-08-29 | 18:46 | 18.5 | 40 | 61 | 8.78 | 93.7 | 8.14
2009-08-29 | 19:01 | 18.45 | 40 | 61 | 8.77 | 93.5 | 8.12
2009-08-29 | 19:16 | 18.42 | 40 | 61 | 8.74 | 93.1 | 8.11
2009-08-29 | 19:31 | 18.4 | 40 | 61 | 8.73 | 93 | 8.1
2009-08-29 | 19:46 | 18.36 | 40 | 62 | 8.73 | 92.9 | 8.08
2009-08-29 | 20:01 | 18.33 | 40 | 62 | 8.72 | 92.8 | 8.07
2009-08-29 | 20:16 | 18.27 | 40 | 62 | 8.7 | 92.4 | 8.05
2009-08-29 | 20:31 | 18.19 | 40 | 62 | 8.72 | 92.5 | 8.04
2009-08-29 | 20:46 | 18.08 | 40 | 62 | 8.75 | 92.6 | 8.03
2009-08-29 | 21:01 | 17.96 | 40 | 62 | 8.78 | 92.7 | 8.02
2009-08-29 | 21:16 | 17.81 | 40 | 62 | 8.79 | 92.4 | 8.02
2009-08-29 | 21:31 | 17.64 | 40 | 62 | 8.81 | 92.4 | 8.02
2009-08-29 | 21:46 | 17.48 | 41 | 62 | 8.87 | 92.7 | 8.01
2009-08-29 | 22:01 | 17.31 | 41 | 62 | 8.88 | 92.5 | 8
2009-08-29 | 22:16 | 17.14 | 41 | 62 | 8.93 | 92.7 | 8
2009-08-29 | 22:31 | 16.96 | 41 | 63 | 8.98 | 92.8 | 8
2009-08-29 | 22:46 | 16.78 | 41 | 63 | 9.02 | 92.9 | 8
2009-08-29 | 23:01 | 16.6 | 41 | 63 | 9.06 | 93 | 8
2009-08-29 | 23:16 | 16.43 | 41 | 63 | 9.1 | 93 | 8
2009-08-29 | 23:31 | 16.27 | 41 | 63 | 9.1 | 92.8 | 8
2009-08-29 | 23:46 | 16.12 | 41 | 63 | 9.16 | 93 | 8
2009-08-30 | 0:01 | 15.95 | 41 | 63 | 9.2 | 93.1 | 8
2009-08-30 | 0:16 | 15.77 | 41 | 63 | 9.21 | 92.9 | 8
2009-08-30 | 0:31 | 15.62 | 41 | 63 | 9.27 | 93.2 | 8
2009-08-30 | 0:46 | 15.46 | 41 | 63 | 9.31 | 93.2 | 8
2009-08-30 | 1:01 | 15.29 | 41 | 63 | 9.32 | 93 | 8
2009-08-30 | 1:16 | 15.13 | 41 | 63 | 9.37 | 93.2 | 8
2009-08-30 | 1:31 | 14.97 | 41 | 63 | 9.39 | 93 | 8
2009-08-30 | 1:46 | 14.81 | 41 | 63 | 9.41 | 93 | 7.99
2009-08-30 | 2:01 | 14.66 | 41 | 63 | 9.47 | 93.3 | 7.99
2009-08-30 | 2:16 | 14.52 | 41 | 63 | 9.52 | 93.4 | 7.99
2009-08-30 | 2:31 | 14.38 | 41 | 63 | 9.52 | 93.2 | 8
2009-08-30 | 2:46 | 14.24 | 41 | 63 | 9.55 | 93.1 | 8
2009-08-30 | 3:01 | 14.1 | 41 | 63 | 9.56 | 93 | 8
2009-08-30 | 3:16 | 13.95 | 41 | 63 | 9.64 | 93.5 | 7.99
2009-08-30 | 3:31 | 13.81 | 41 | 64 | 9.65 | 93.3 | 8
2009-08-30 | 3:46 | 13.69 | 41 | 64 | 9.71 | 93.6 | 7.99
2009-08-30 | 4:01 | 13.55 | 41 | 64 | 9.73 | 93.5 | 7.99
2009-08-30 | 4:16 | 13.44 | 41 | 64 | 9.75 | 93.5 | 8
2009-08-30 | 4:31 | 13.33 | 42 | 64 | 9.82 | 93.8 | 8
2009-08-30 | 4:46 | 13.23 | 42 | 65 | 9.85 | 94 | 8.01
2009-08-30 | 5:01 | 13.12 | 42 | 65 | 9.89 | 94.1 | 8.01
2009-08-30 | 5:16 | 13.01 | 42 | 65 | 9.92 | 94.2 | 8.01
2009-08-30 | 5:31 | 12.9 | 42 | 65 | 9.92 | 93.9 | 8.01
2009-08-30 | 5:46 | 12.8 | 42 | 66 | 9.93 | 93.9 | 8.01
2009-08-30 | 6:01 | 12.7 | 43 | 66 | 9.99 | 94.1 | 8.02
2009-08-30 | 6:16 | 12.59 | 43 | 67 | 9.99 | 93.9 | 8.02
2009-08-30 | 6:31 | 12.5 | 43 | 67 | 10.01 | 93.9 | 8.02
2009-08-30 | 6:46 | 12.4 | 44 | 68 | 10.04 | 94 | 8.03
2009-08-30 | 7:01 | 12.31 | 44 | 68 | 10.11 | 94.4 | 8.04
2009-08-30 | 7:16 | 12.22 | 44 | 68 | 10.12 | 94.4 | 8.05
2009-08-30 | 7:31 | 12.13 | 44 | 69 | 10.16 | 94.5 | 8.06
2009-08-30 | 7:46 | 12.04 | 45 | 69 | 10.22 | 94.9 | 8.06
2009-08-30 | 8:01 | 11.96 | 45 | 70 | 10.23 | 94.9 | 8.07
2009-08-30 | 8:16 | 11.89 | 45 | 70 | 10.29 | 95.2 | 8.08
2009-08-30 | 8:31 | 11.84 | 46 | 71 | 10.29 | 95.2 | 8.09
2009-08-30 | 8:46 | 11.81 | 46 | 71 | 10.34 | 95.5 | 8.1
2009-08-30 | 9:01 | 11.81 | 46 | 71 | 10.38 | 95.9 | 8.11
2009-08-30 | 9:16 | 11.83 | 46 | 71 | 10.4 | 96.1 | 8.12
2009-08-30 | 9:31 | 11.86 | 46 | 71 | 10.38 | 96 | 8.13
2009-08-30 | 9:46 | 11.93 | 46 | 71 | 10.42 | 96.5 | 8.14
2009-08-30 | 10:01 | 12.01 | 46 | 70 | 10.39 | 96.5 | 8.16
2009-08-30 | 10:16 | 12.09 | 46 | 69 | 10.42 | 96.9 | 8.16
2009-08-30 | 10:31 | 12.17 | 45 | 68 | 10.39 | 96.8 | 8.16
2009-08-30 | 10:46 | 12.27 | 44 | 67 | 10.38 | 96.9 | 8.16
2009-08-30 | 11:01 | 12.41 | 44 | 66 | 10.36 | 97 | 8.16
2009-08-30 | 11:16 | 12.59 | 43 | 65 | 10.33 | 97.2 | 8.17
2009-08-30 | 11:31 | 12.79 | 43 | 65 | 10.29 | 97.3 | 8.18
2009-08-30 | 11:46 | 13.06 | 42 | 64 | 10.27 | 97.6 | 8.18
2009-08-30 | 12:01 | 13.41 | 42 | 64 | 10.17 | 97.5 | 8.19
2009-08-30 | 12:16 | 13.83 | 41 | 63 | 10.1 | 97.7 | 8.2
2009-08-30 | 12:31 | 14.27 | 41 | 63 | 9.98 | 97.4 | 8.2
2009-08-30 | 12:46 | 14.74 | 41 | 63 | 9.9 | 97.6 | 8.21
2009-08-30 | 13:01 | 15.23 | 41 | 62 | 9.8 | 97.7 | 8.24
2009-08-30 | 13:16 | 15.74 | 41 | 62 | 9.69 | 97.6 | 8.25
2009-08-30 | 13:31 | 16.25 | 40 | 62 | 9.55 | 97.3 | 8.26
2009-08-30 | 13:46 | 16.76 | 40 | 62 | 9.43 | 97.1 | 8.27
2009-08-30 | 14:01 | 17.27 | 40 | 62 | 9.35 | 97.3 | 8.29
2009-08-30 | 14:16 | 17.78 | 40 | 62 | 9.25 | 97.3 | 8.3
2009-08-30 | 14:31 | 18.28 | 40 | 61 | 9.12 | 96.8 | 8.31
2009-08-30 | 14:46 | 18.72 | 40 | 61 | 9.01 | 96.6 | 8.33
2009-08-30 | 15:01 | 19.07 | 40 | 61 | 8.96 | 96.7 | 8.34
2009-08-30 | 15:16 | 19.36 | 40 | 61 | 8.9 | 96.6 | 8.35
2009-08-30 | 15:31 | 19.64 | 40 | 61 | 8.81 | 96.2 | 8.34
2009-08-30 | 15:46 | 19.89 | 40 | 61 | 8.74 | 95.9 | 8.35
2009-08-30 | 16:01 | 20.08 | 40 | 61 | 8.71 | 96 | 8.35
2009-08-30 | 16:16 | 20.15 | 40 | 61 | 8.69 | 95.8 | 8.34
2009-08-30 | 16:31 | 20.12 | 40 | 62 | 8.64 | 95.3 | 8.33
2009-08-30 | 16:46 | 20.06 | 40 | 62 | 8.65 | 95.2 | 8.3
2009-08-30 | 17:01 | 19.94 | 40 | 62 | 8.64 | 94.9 | 8.28
2009-08-30 | 17:16 | 19.76 | 40 | 62 | 8.64 | 94.6 | 8.24
2009-08-30 | 17:31 | 19.6 | 40 | 62 | 8.64 | 94.2 | 8.21
2009-08-30 | 17:46 | 19.45 | 40 | 62 | 8.64 | 93.9 | 8.18
2009-08-30 | 18:01 | 19.3 | 40 | 62 | 8.65 | 93.8 | 8.17
2009-08-30 | 18:16 | 19.17 | 40 | 62 | 8.65 | 93.5 | 8.13
2009-08-30 | 18:31 | 19.05 | 40 | 62 | 8.65 | 93.4 | 8.12
2009-08-30 | 18:46 | 18.95 | 40 | 62 | 8.66 | 93.2 | 8.11
2009-08-30 | 19:01 | 18.88 | 40 | 62 | 8.66 | 93.1 | 8.09
2009-08-30 | 19:16 | 18.83 | 40 | 62 | 8.65 | 93 | 8.08
2009-08-30 | 19:31 | 18.77 | 41 | 62 | 8.65 | 92.8 | 8.07
2009-08-30 | 19:46 | 18.71 | 41 | 62 | 8.62 | 92.4 | 8.06
2009-08-30 | 20:01 | 18.64 | 41 | 63 | 8.61 | 92.1 | 8.05
2009-08-30 | 20:16 | 18.56 | 41 | 63 | 8.6 | 91.9 | 8.03
2009-08-30 | 20:31 | 18.48 | 41 | 63 | 8.61 | 91.8 | 8.03
2009-08-30 | 20:46 | 18.37 | 41 | 63 | 8.62 | 91.8 | 8.02
2009-08-30 | 21:01 | 18.26 | 41 | 63 | 8.64 | 91.8 | 8.02
2009-08-30 | 21:16 | 18.15 | 41 | 63 | 8.69 | 92.1 | 8.01
2009-08-30 | 21:31 | 18.02 | 41 | 63 | 8.7 | 92 | 8.01
2009-08-30 | 21:46 | 17.86 | 41 | 64 | 8.73 | 92 | 8.01
2009-08-30 | 22:01 | 17.7 | 41 | 64 | 8.79 | 92.3 | 8.01
2009-08-30 | 22:16 | 17.53 | 41 | 64 | 8.81 | 92.1 | 8.01
2009-08-30 | 22:31 | 17.36 | 41 | 64 | 8.86 | 92.4 | 8.01
2009-08-30 | 22:46 | 17.19 | 42 | 64 | 8.88 | 92.3 | 8.01
2009-08-30 | 23:01 | 17.01 | 42 | 64 | 8.94 | 92.5 | 8.01
2009-08-30 | 23:16 | 16.83 | 42 | 64 | 8.96 | 92.4 | 8.01
2009-08-30 | 23:31 | 16.65 | 42 | 64 | 9.02 | 92.6 | 8.01
2009-08-30 | 23:46 | 16.48 | 42 | 65 | 9.04 | 92.5 | 8.01
2009-08-31 | 0:01 | 16.31 | 42 | 65 | 9.09 | 92.7 | 8.01
2009-08-31 | 0:16 | 16.14 | 42 | 65 | 9.13 | 92.8 | 8
2009-08-31 | 0:31 | 15.98 | 42 | 65 | 9.15 | 92.7 | 8
2009-08-31 | 0:46 | 15.82 | 42 | 65 | 9.18 | 92.6 | 8
2009-08-31 | 1:01 | 15.66 | 42 | 65 | 9.2 | 92.6 | 8
2009-08-31 | 1:16 | 15.51 | 42 | 65 | 9.27 | 92.9 | 8.01
2009-08-31 | 1:31 | 15.35 | 42 | 65 | 9.31 | 93.1 | 8
2009-08-31 | 1:46 | 15.21 | 42 | 65 | 9.32 | 92.9 | 8
2009-08-31 | 2:01 | 15.07 | 42 | 65 | 9.35 | 92.9 | 8
2009-08-31 | 2:16 | 14.94 | 43 | 65 | 9.37 | 92.8 | 8.01
2009-08-31 | 2:31 | 14.81 | 43 | 66 | 9.43 | 93.1 | 8
2009-08-31 | 2:46 | 14.69 | 43 | 66 | 9.46 | 93.2 | 8
2009-08-31 | 3:01 | 14.56 | 43 | 66 | 9.47 | 93 | 8
2009-08-31 | 3:16 | 14.44 | 43 | 66 | 9.49 | 92.9 | 8.01
2009-08-31 | 3:31 | 14.32 | 43 | 66 | 9.51 | 92.9 | 8.01
2009-08-31 | 3:46 | 14.2 | 43 | 66 | 9.57 | 93.3 | 8.01
2009-08-31 | 4:01 | 14.09 | 43 | 66 | 9.61 | 93.4 | 8
2009-08-31 | 4:16 | 13.98 | 43 | 66 | 9.61 | 93.2 | 8
2009-08-31 | 4:31 | 13.88 | 43 | 67 | 9.67 | 93.6 | 8.01
2009-08-31 | 4:46 | 13.78 | 43 | 67 | 9.72 | 93.8 | 8.01
2009-08-31 | 5:01 | 13.67 | 44 | 67 | 9.75 | 93.9 | 8.01
2009-08-31 | 5:16 | 13.58 | 44 | 67 | 9.74 | 93.7 | 8.02
2009-08-31 | 5:31 | 13.49 | 44 | 68 | 9.75 | 93.5 | 8.02
2009-08-31 | 5:46 | 13.4 | 44 | 68 | 9.77 | 93.6 | 8.03
2009-08-31 | 6:01 | 13.32 | 44 | 69 | 9.82 | 93.8 | 8.03
2009-08-31 | 6:16 | 13.25 | 45 | 69 | 9.82 | 93.7 | 8.03
2009-08-31 | 6:31 | 13.18 | 45 | 69 | 9.83 | 93.7 | 8.03
2009-08-31 | 6:46 | 13.1 | 45 | 70 | 9.89 | 94.1 | 8.04
2009-08-31 | 7:01 | 13.03 | 46 | 70 | 9.94 | 94.4 | 8.05
2009-08-31 | 7:16 | 12.95 | 46 | 71 | 9.98 | 94.6 | 8.06
2009-08-31 | 7:31 | 12.86 | 46 | 72 | 10.02 | 94.8 | 8.07
2009-08-31 | 7:46 | 12.78 | 47 | 72 | 10.06 | 95 | 8.08
2009-08-31 | 8:01 | 12.7 | 47 | 73 | 10.09 | 95.2 | 8.09
2009-08-31 | 8:16 | 12.63 | 47 | 73 | 10.1 | 95 | 8.1
2009-08-31 | 8:31 | 12.58 | 47 | 73 | 10.11 | 95 | 8.11
2009-08-31 | 8:46 | 12.56 | 48 | 74 | 10.16 | 95.4 | 8.11
2009-08-31 | 9:01 | 12.57 | 48 | 74 | 10.19 | 95.8 | 8.13
2009-08-31 | 9:16 | 12.58 | 48 | 74 | 10.21 | 96 | 8.14
2009-08-31 | 9:31 | 12.61 | 48 | 74 | 10.24 | 96.3 | 8.15
2009-08-31 | 9:46 | 12.67 | 48 | 73 | 10.21 | 96.2 | 8.16
2009-08-31 | 10:01 | 12.75 | 48 | 73 | 10.25 | 96.7 | 8.17
2009-08-31 | 10:16 | 12.81 | 47 | 71 | 10.26 | 97 | 8.17
2009-08-31 | 10:31 | 12.9 | 46 | 70 | 10.26 | 97.2 | 8.17
2009-08-31 | 10:46 | 13.02 | 46 | 69 | 10.22 | 97.1 | 8.18
2009-08-31 | 11:01 | 13.15 | 45 | 68 | 10.2 | 97.2 | 8.18
2009-08-31 | 11:16 | 13.32 | 44 | 66 | 10.21 | 97.6 | 8.18
2009-08-31 | 11:31 | 13.54 | 43 | 65 | 10.14 | 97.4 | 8.19
2009-08-31 | 11:46 | 13.83 | 42 | 65 | 10.12 | 97.8 | 8.19
2009-08-31 | 12:01 | 14.18 | 42 | 64 | 10.05 | 97.9 | 8.2
2009-08-31 | 12:16 | 14.61 | 42 | 63 | 9.92 | 97.6 | 8.21
2009-08-31 | 12:31 | 15.07 | 41 | 63 | 9.81 | 97.5 | 8.22
2009-08-31 | 12:46 | 15.56 | 41 | 62 | 9.7 | 97.4 | 8.24
2009-08-31 | 13:01 | 16.07 | 41 | 62 | 9.59 | 97.3 | 8.26
2009-08-31 | 13:16 | 16.59 | 40 | 62 | 9.51 | 97.5 | 8.27
2009-08-31 | 13:31 | 17.12 | 40 | 62 | 9.38 | 97.3 | 8.29
2009-08-31 | 13:46 | 17.65 | 40 | 61 | 9.3 | 97.5 | 8.3
2009-08-31 | 14:01 | 18.17 | 40 | 61 | 9.17 | 97.2 | 8.32
2009-08-31 | 14:16 | 18.69 | 40 | 61 | 9.05 | 97 | 8.34
2009-08-31 | 14:31 | 19.21 | 40 | 61 | 8.97 | 97.2 | 8.35
2009-08-31 | 14:46 | 19.63 | 40 | 61 | 8.89 | 97.1 | 8.37
2009-08-31 | 15:01 | 19.98 | 40 | 61 | 8.79 | 96.7 | 8.37
2009-08-31 | 15:16 | 20.29 | 40 | 61 | 8.71 | 96.4 | 8.38
2009-08-31 | 15:31 | 20.55 | 40 | 61 | 8.67 | 96.4 | 8.38
2009-08-31 | 15:46 | 20.72 | 40 | 61 | 8.61 | 96 | 8.37
2009-08-31 | 16:01 | 20.9 | 40 | 61 | 8.56 | 95.8 | 8.37
2009-08-31 | 16:16 | 20.97 | 40 | 61 | 8.52 | 95.6 | 8.37
2009-08-31 | 16:31 | 20.95 | 40 | 61 | 8.5 | 95.2 | 8.35
2009-08-31 | 16:46 | 20.88 | 40 | 61 | 8.5 | 95.2 | 8.32
2009-08-31 | 17:01 | 20.8 | 40 | 61 | 8.49 | 94.9 | 8.29
2009-08-31 | 17:16 | 20.64 | 40 | 61 | 8.46 | 94.3 | 8.27
2009-08-31 | 17:31 | 20.49 | 40 | 62 | 8.47 | 94.1 | 8.24
2009-08-31 | 17:46 | 20.36 | 40 | 62 | 8.46 | 93.8 | 8.21
2009-08-31 | 18:01 | 20.23 | 40 | 62 | 8.47 | 93.6 | 8.18
2009-08-31 | 18:16 | 20.12 | 40 | 62 | 8.45 | 93.2 | 8.16
2009-08-31 | 18:31 | 20.04 | 40 | 62 | 8.47 | 93.2 | 8.14
2009-08-31 | 18:46 | 19.99 | 40 | 62 | 8.47 | 93.1 | 8.11
2009-08-31 | 19:01 | 19.95 | 40 | 62 | 8.46 | 93 | 8.1
2009-08-31 | 19:16 | 19.9 | 40 | 62 | 8.46 | 92.8 | 8.08
2009-08-31 | 19:31 | 19.82 | 40 | 62 | 8.43 | 92.4 | 8.06
2009-08-31 | 19:46 | 19.75 | 40 | 62 | 8.41 | 92.1 | 8.05
2009-08-31 | 20:01 | 19.68 | 40 | 62 | 8.4 | 91.8 | 8.04
2009-08-31 | 20:16 | 19.59 | 40 | 62 | 8.44 | 92 | 8.02
2009-08-31 | 20:31 | 19.49 | 41 | 62 | 8.45 | 92 | 8.02
2009-08-31 | 20:46 | 19.4 | 41 | 62 | 8.47 | 92.1 | 8.01
2009-08-31 | 21:01 | 19.31 | 41 | 62 | 8.49 | 92.1 | 8.01
2009-08-31 | 21:16 | 19.2 | 41 | 63 | 8.49 | 91.9 | 8
2009-08-31 | 21:31 | 19.07 | 41 | 63 | 8.53 | 92.1 | 8
2009-08-31 | 21:46 | 18.93 | 41 | 63 | 8.54 | 92 | 8
2009-08-31 | 22:01 | 18.78 | 41 | 63 | 8.57 | 91.9 | 8
2009-08-31 | 22:16 | 18.62 | 41 | 63 | 8.63 | 92.3 | 8
2009-08-31 | 22:31 | 18.46 | 41 | 63 | 8.64 | 92.1 | 8
2009-08-31 | 22:46 | 18.3 | 41 | 63 | 8.66 | 92.1 | 8
2009-08-31 | 23:01 | 18.16 | 41 | 64 | 8.72 | 92.4 | 8
2009-08-31 | 23:16 | 18 | 41 | 64 | 8.76 | 92.5 | 8
2009-08-31 | 23:31 | 17.85 | 41 | 64 | 8.79 | 92.6 | 8
2009-08-31 | 23:46 | 17.7 | 41 | 64 | 8.83 | 92.7 | 8
2009-09-01 | 0:01 | 17.55 | 41 | 64 | 8.86 | 92.7 | 8
2009-09-01 | 0:16 | 17.42 | 42 | 64 | 8.87 | 92.5 | 8
2009-09-01 | 0:31 | 17.29 | 42 | 64 | 8.88 | 92.5 | 8
2009-09-01 | 0:46 | 17.15 | 42 | 64 | 8.91 | 92.5 | 8
2009-09-01 | 1:01 | 17.02 | 42 | 64 | 8.93 | 92.5 | 8
2009-09-01 | 1:16 | 16.9 | 42 | 64 | 8.96 | 92.5 | 8
2009-09-01 | 1:31 | 16.79 | 42 | 64 | 8.99 | 92.6 | 8
2009-09-01 | 1:46 | 16.68 | 42 | 64 | 9.01 | 92.6 | 8
2009-09-01 | 2:01 | 16.57 | 42 | 64 | 9.06 | 92.9 | 8
2009-09-01 | 2:16 | 16.47 | 42 | 64 | 9.06 | 92.7 | 8
2009-09-01 | 2:31 | 16.39 | 42 | 64 | 9.08 | 92.7 | 8
2009-09-01 | 2:46 | 16.31 | 42 | 65 | 9.1 | 92.8 | 8
2009-09-01 | 3:01 | 16.23 | 42 | 65 | 9.12 | 92.9 | 8
2009-09-01 | 3:16 | 16.16 | 42 | 65 | 9.15 | 93 | 8.01
2009-09-01 | 3:31 | 16.08 | 42 | 65 | 9.18 | 93.1 | 8.01
2009-09-01 | 3:46 | 16 | 42 | 65 | 9.2 | 93.2 | 8.02
2009-09-01 | 4:01 | 15.94 | 42 | 66 | 9.25 | 93.6 | 8.02
2009-09-01 | 4:16 | 15.86 | 43 | 66 | 9.25 | 93.5 | 8.02
2009-09-01 | 4:31 | 15.79 | 43 | 66 | 9.27 | 93.5 | 8.02
2009-09-01 | 4:46 | 15.72 | 43 | 67 | 9.28 | 93.5 | 8.02
2009-09-01 | 5:01 | 15.67 | 43 | 67 | 9.32 | 93.8 | 8.03
2009-09-01 | 5:16 | 15.61 | 44 | 68 | 9.34 | 93.9 | 8.03
2009-09-01 | 5:31 | 15.56 | 44 | 69 | 9.36 | 93.9 | 8.04
2009-09-01 | 5:46 | 15.5 | 45 | 69 | 9.34 | 93.7 | 8.04
2009-09-01 | 6:01 | 15.43 | 45 | 70 | 9.35 | 93.6 | 8.05
2009-09-01 | 6:16 | 15.37 | 46 | 71 | 9.37 | 93.6 | 8.06
2009-09-01 | 6:31 | 15.29 | 46 | 72 | 9.41 | 93.9 | 8.07
2009-09-01 | 6:46 | 15.18 | 47 | 73 | 9.43 | 93.9 | 8.07
2009-09-01 | 7:01 | 15.06 | 48 | 74 | 9.46 | 94 | 8.08
2009-09-01 | 7:16 | 14.93 | 48 | 74 | 9.54 | 94.5 | 8.09
2009-09-01 | 7:31 | 14.81 | 48 | 74 | 9.57 | 94.5 | 8.1
2009-09-01 | 7:46 | 14.68 | 48 | 73 | 9.64 | 95 | 8.1
2009-09-01 | 8:01 | 14.56 | 48 | 72 | 9.67 | 95 | 8.11
2009-09-01 | 8:16 | 14.46 | 47 | 71 | 9.72 | 95.3 | 8.1
2009-09-01 | 8:31 | 14.38 | 46 | 69 | 9.74 | 95.3 | 8.11
2009-09-01 | 8:46 | 14.33 | 45 | 68 | 9.76 | 95.4 | 8.1
2009-09-01 | 9:01 | 14.32 | 44 | 66 | 9.81 | 95.9 | 8.1
2009-09-01 | 9:16 | 14.32 | 43 | 65 | 9.84 | 96.2 | 8.1
2009-09-01 | 9:31 | 14.36 | 42 | 63 | 9.83 | 96.1 | 8.1
2009-09-01 | 9:46 | 14.42 | 41 | 62 | 9.87 | 96.6 | 8.12
2009-09-01 | 10:01 | 14.49 | 41 | 62 | 9.89 | 97 | 8.12
2009-09-01 | 10:16 | 14.56 | 40 | 61 | 9.9 | 97.2 | 8.13
2009-09-01 | 10:31 | 14.62 | 40 | 60 | 9.87 | 97.1 | 8.13
2009-09-01 | 10:46 | 14.73 | 39 | 60 | 9.86 | 97.3 | 8.15
2009-09-01 | 11:01 | 14.85 | 39 | 59 | 9.85 | 97.4 | 8.16
2009-09-01 | 11:16 | 15.01 | 38 | 59 | 9.87 | 97.9 | 8.17
2009-09-01 | 11:31 | 15.19 | 38 | 59 | 9.83 | 97.9 | 8.18
2009-09-01 | 11:46 | 15.45 | 38 | 58 | 9.8 | 98.2 | 8.2
2009-09-01 | 12:01 | 15.77 | 38 | 58 | 9.76 | 98.4 | 8.21
2009-09-01 | 12:16 | 16.13 | 38 | 58 | 9.68 | 98.4 | 8.22
2009-09-01 | 12:31 | 16.53 | 38 | 58 | 9.61 | 98.4 | 8.25
2009-09-01 | 12:46 | 16.94 | 37 | 58 | 9.52 | 98.4 | 8.25
2009-09-01 | 13:01 | 17.34 | 37 | 57 | 9.4 | 98 | 8.26
2009-09-01 | 13:16 | 17.74 | 37 | 57 | 9.34 | 98.2 | 8.28
2009-09-01 | 13:31 | 18.19 | 37 | 57 | 9.23 | 97.9 | 8.31
2009-09-01 | 13:46 | 18.48 | 37 | 57 | 9.15 | 97.6 | 8.33
2009-09-01 | 14:01 | 18.79 | 37 | 57 | 9.1 | 97.7 | 8.34
2009-09-01 | 14:16 | 19.29 | 37 | 57 | 9.02 | 97.8 | 8.36
2009-09-01 | 14:31 | 19.62 | 37 | 57 | 9 | 98.3 | 8.38
2009-09-01 | 14:46 | 19.92 | 37 | 57 | 8.91 | 97.9 | 8.4
2009-09-01 | 15:01 | 20.1 | 37 | 58 | 8.88 | 97.9 | 8.4
2009-09-01 | 15:16 | 20.28 | 37 | 57 | 8.82 | 97.5 | 8.39
2009-09-01 | 15:31 | 20.48 | 37 | 57 | 8.75 | 97.2 | 8.38
2009-09-01 | 15:46 | 20.63 | 37 | 58 | 8.72 | 97.1 | 8.39
2009-09-01 | 16:01 | 20.75 | 37 | 58 | 8.67 | 96.8 | 8.38
2009-09-01 | 16:16 | 20.75 | 37 | 58 | 8.67 | 96.8 | 8.37
2009-09-01 | 16:31 | 20.69 | 38 | 58 | 8.64 | 96.3 | 8.35
2009-09-01 | 16:46 | 20.6 | 38 | 58 | 8.64 | 96.2 | 8.32
2009-09-01 | 17:01 | 20.49 | 38 | 58 | 8.63 | 95.9 | 8.3
2009-09-01 | 17:16 | 20.3 | 38 | 58 | 8.61 | 95.3 | 8.26
2009-09-01 | 17:31 | 20.16 | 38 | 58 | 8.6 | 94.9 | 8.23
2009-09-01 | 17:46 | 20.01 | 38 | 58 | 8.59 | 94.5 | 8.19
2009-09-01 | 18:01 | 19.88 | 38 | 58 | 8.61 | 94.4 | 8.16
2009-09-01 | 18:16 | 19.77 | 38 | 58 | 8.62 | 94.3 | 8.14
2009-09-01 | 18:31 | 19.7 | 38 | 58 | 8.59 | 93.9 | 8.11
2009-09-01 | 18:46 | 19.62 | 38 | 58 | 8.6 | 93.9 | 8.1
2009-09-01 | 19:01 | 19.56 | 38 | 59 | 8.59 | 93.6 | 8.08
2009-09-01 | 19:16 | 19.53 | 38 | 59 | 8.59 | 93.6 | 8.07
2009-09-01 | 19:31 | 19.47 | 38 | 59 | 8.59 | 93.5 | 8.06
2009-09-01 | 19:46 | 19.43 | 38 | 59 | 8.57 | 93.1 | 8.04
2009-09-01 | 20:01 | 19.36 | 38 | 59 | 8.55 | 92.8 | 8.03
2009-09-01 | 20:16 | 19.27 | 38 | 59 | 8.56 | 92.7 | 8.02
2009-09-01 | 20:31 | 19.17 | 38 | 59 | 8.6 | 93 | 8.01
2009-09-01 | 20:46 | 19.06 | 38 | 59 | 8.6 | 92.8 | 8
2009-09-01 | 21:01 | 18.94 | 39 | 59 | 8.62 | 92.8 | 8
2009-09-01 | 21:16 | 18.82 | 39 | 59 | 8.65 | 92.9 | 8
2009-09-01 | 21:31 | 18.7 | 39 | 59 | 8.67 | 92.9 | 7.99
2009-09-01 | 21:46 | 18.58 | 39 | 60 | 8.69 | 92.9 | 7.99
2009-09-01 | 22:01 | 18.44 | 39 | 60 | 8.72 | 93 | 7.99
2009-09-01 | 22:16 | 18.31 | 39 | 60 | 8.78 | 93.3 | 7.99
2009-09-01 | 22:31 | 18.16 | 39 | 60 | 8.79 | 93.2 | 7.98
2009-09-01 | 22:46 | 18 | 39 | 60 | 8.85 | 93.5 | 7.99
2009-09-01 | 23:01 | 17.83 | 39 | 60 | 8.89 | 93.6 | 7.98
2009-09-01 | 23:16 | 17.66 | 39 | 61 | 8.91 | 93.4 | 7.98
2009-09-01 | 23:31 | 17.49 | 40 | 61 | 8.96 | 93.7 | 7.98
2009-09-01 | 23:46 | 17.33 | 39 | 61 | 8.98 | 93.5 | 7.98
2009-09-02 | 0:01 | 17.17 | 39 | 61 | 9.01 | 93.5 | 7.98
2009-09-02 | 0:16 | 17.01 | 39 | 61 | 9.04 | 93.6 | 7.98
2009-09-02 | 0:31 | 16.86 | 39 | 61 | 9.1 | 93.8 | 7.99
2009-09-02 | 0:46 | 16.7 | 39 | 61 | 9.11 | 93.7 | 7.98
2009-09-02 | 1:01 | 16.55 | 39 | 61 | 9.14 | 93.7 | 7.98
2009-09-02 | 1:16 | 16.4 | 39 | 61 | 9.17 | 93.7 | 7.99
2009-09-02 | 1:31 | 16.26 | 39 | 61 | 9.2 | 93.7 | 7.99
2009-09-02 | 1:46 | 16.12 | 40 | 61 | 9.26 | 94.1 | 7.98
2009-09-02 | 2:01 | 15.98 | 40 | 61 | 9.3 | 94.2 | 7.99
2009-09-02 | 2:16 | 15.84 | 40 | 61 | 9.32 | 94.1 | 7.99
2009-09-02 | 2:31 | 15.7 | 40 | 61 | 9.35 | 94.1 | 7.99
2009-09-02 | 2:46 | 15.56 | 40 | 61 | 9.41 | 94.5 | 7.99
2009-09-02 | 3:01 | 15.42 | 40 | 62 | 9.43 | 94.4 | 7.99
2009-09-02 | 3:16 | 15.3 | 40 | 62 | 9.45 | 94.3 | 7.99
2009-09-02 | 3:31 | 15.17 | 40 | 62 | 9.51 | 94.7 | 8
2009-09-02 | 3:46 | 15.04 | 40 | 62 | 9.56 | 94.9 | 8
2009-09-02 | 4:01 | 14.92 | 41 | 63 | 9.56 | 94.7 | 8
2009-09-02 | 4:16 | 14.8 | 41 | 63 | 9.61 | 94.9 | 8
2009-09-02 | 4:31 | 14.7 | 41 | 63 | 9.64 | 95 | 8
2009-09-02 | 4:46 | 14.6 | 41 | 64 | 9.64 | 94.8 | 8.01
2009-09-02 | 5:01 | 14.51 | 41 | 64 | 9.66 | 94.8 | 8.01
2009-09-02 | 5:16 | 14.4 | 42 | 65 | 9.71 | 95.1 | 8.01
2009-09-02 | 5:31 | 14.29 | 42 | 65 | 9.75 | 95.2 | 8.02
2009-09-02 | 5:46 | 14.17 | 42 | 66 | 9.75 | 95 | 8.02
2009-09-02 | 6:01 | 14.05 | 43 | 66 | 9.8 | 95.2 | 8.03
2009-09-02 | 6:16 | 13.93 | 43 | 66 | 9.84 | 95.4 | 8.03
2009-09-02 | 6:31 | 13.8 | 43 | 67 | 9.89 | 95.5 | 8.03
2009-09-02 | 6:46 | 13.66 | 43 | 67 | 9.9 | 95.3 | 8.03
2009-09-02 | 7:01 | 13.53 | 44 | 67 | 9.94 | 95.4 | 8.04
2009-09-02 | 7:16 | 13.4 | 44 | 67 | 9.98 | 95.5 | 8.04
2009-09-02 | 7:31 | 13.28 | 44 | 67 | 10.05 | 96 | 8.05
2009-09-02 | 7:46 | 13.18 | 43 | 66 | 10.11 | 96.4 | 8.05
2009-09-02 | 8:01 | 13.1 | 43 | 65 | 10.14 | 96.4 | 8.06
2009-09-02 | 8:16 | 13.05 | 42 | 64 | 10.2 | 96.9 | 8.06
2009-09-02 | 8:31 | 13 | 42 | 63 | 10.23 | 97.1 | 8.07
2009-09-02 | 8:46 | 12.98 | 41 | 62 | 10.25 | 97.3 | 8.07
2009-09-02 | 9:01 | 12.99 | 41 | 62 | 10.23 | 97.1 | 8.07
2009-09-02 | 9:16 | 13.01 | 40 | 61 | 10.27 | 97.5 | 8.07
2009-09-02 | 9:31 | 13.06 | 40 | 60 | 10.26 | 97.5 | 8.09
2009-09-02 | 9:46 | 13.14 | 39 | 60 | 10.26 | 97.7 | 8.1
2009-09-02 | 10:01 | 13.22 | 39 | 59 | 10.3 | 98.3 | 8.11
2009-09-02 | 10:16 | 13.32 | 38 | 59 | 10.35 | 98.9 | 8.12
2009-09-02 | 10:31 | 13.39 | 38 | 58 | 10.29 | 98.5 | 8.13
2009-09-02 | 10:46 | 13.53 | 38 | 58 | 10.28 | 98.8 | 8.14
2009-09-02 | 11:01 | 13.68 | 38 | 57 | 10.3 | 99.3 | 8.16
2009-09-02 | 11:16 | 13.84 | 37 | 57 | 10.28 | 99.4 | 8.17
2009-09-02 | 11:31 | 14.06 | 37 | 57 | 10.21 | 99.2 | 8.17
2009-09-02 | 11:46 | 14.35 | 37 | 57 | 10.18 | 99.5 | 8.2
2009-09-02 | 12:01 | 14.69 | 37 | 56 | 10.09 | 99.4 | 8.2
2009-09-02 | 12:16 | 15.13 | 37 | 56 | 10.02 | 99.7 | 8.22
2009-09-02 | 12:31 | 15.53 | 37 | 56 | 9.94 | 99.7 | 8.23
2009-09-02 | 12:46 | 15.95 | 36 | 56 | 9.79 | 99.1 | 8.25
2009-09-02 | 13:01 | 16.44 | 37 | 56 | 9.66 | 98.7 | 8.26
2009-09-02 | 13:16 | 16.9 | 36 | 56 | 9.55 | 98.6 | 8.26
2009-09-02 | 13:31 | 17.37 | 37 | 56 | 9.44 | 98.5 | 8.28
2009-09-02 | 13:46 | 17.86 | 36 | 56 | 9.34 | 98.4 | 8.3
2009-09-02 | 14:01 | 18.33 | 36 | 56 | 9.27 | 98.6 | 8.31
2009-09-02 | 14:16 | 18.83 | 36 | 56 | 9.15 | 98.3 | 8.32
2009-09-02 | 14:31 | 19.31 | 37 | 56 | 9.04 | 98.1 | 8.34
2009-09-02 | 14:46 | 19.72 | 37 | 56 | 8.98 | 98.2 | 8.36
2009-09-02 | 15:01 | 20.04 | 37 | 56 | 8.91 | 98.1 | 8.36
2009-09-02 | 15:16 | 20.27 | 37 | 57 | 8.81 | 97.5 | 8.37
2009-09-02 | 15:31 | 20.47 | 37 | 57 | 8.77 | 97.4 | 8.36
2009-09-02 | 15:46 | 20.63 | 37 | 57 | 8.75 | 97.4 | 8.36
2009-09-02 | 16:01 | 20.61 | 37 | 57 | 8.73 | 97.2 | 8.35
2009-09-02 | 16:16 | 20.61 | 37 | 57 | 8.71 | 96.9 | 8.34
2009-09-02 | 16:31 | 20.56 | 37 | 57 | 8.59 | 95.5 | 8.31
2009-09-02 | 16:46 | 20.48 | 37 | 57 | 8.68 | 96.4 | 8.3
2009-09-02 | 17:01 | 20.44 | 37 | 57 | 8.68 | 96.3 | 8.27
2009-09-02 | 17:16 | 20.31 | 37 | 57 | 8.69 | 96.2 | 8.25
2009-09-02 | 17:31 | 20.2 | 37 | 57 | 8.67 | 95.7 | 8.23
2009-09-02 | 17:46 | 20.12 | 37 | 58 | 8.67 | 95.6 | 8.21
2009-09-02 | 18:01 | 20.01 | 37 | 58 | 8.64 | 95 | 8.18
2009-09-02 | 18:16 | 19.91 | 37 | 58 | 8.61 | 94.5 | 8.15
2009-09-02 | 18:31 | 19.83 | 37 | 58 | 8.62 | 94.4 | 8.11
2009-09-02 | 18:46 | 19.74 | 38 | 58 | 8.59 | 94 | 8.09
2009-09-02 | 19:01 | 19.66 | 38 | 58 | 8.58 | 93.7 | 8.07
2009-09-02 | 19:16 | 19.58 | 38 | 58 | 8.58 | 93.5 | 8.05
2009-09-02 | 19:31 | 19.49 | 38 | 58 | 8.57 | 93.3 | 8.04
2009-09-02 | 19:46 | 19.41 | 38 | 58 | 8.56 | 93.1 | 8.02
2009-09-02 | 20:01 | 19.33 | 38 | 58 | 8.57 | 93 | 8
2009-09-02 | 20:16 | 19.25 | 38 | 59 | 8.57 | 92.8 | 7.99
2009-09-02 | 20:31 | 19.16 | 38 | 59 | 8.58 | 92.8 | 7.98
2009-09-02 | 20:46 | 19.07 | 38 | 59 | 8.59 | 92.8 | 7.98
2009-09-02 | 21:01 | 18.99 | 38 | 59 | 8.64 | 93.1 | 7.98
2009-09-02 | 21:16 | 18.92 | 38 | 59 | 8.66 | 93.2 | 7.97
2009-09-02 | 21:31 | 18.85 | 38 | 59 | 8.66 | 93 | 7.97
2009-09-02 | 21:46 | 18.78 | 38 | 59 | 8.67 | 93.1 | 7.97
2009-09-02 | 22:01 | 18.68 | 38 | 59 | 8.72 | 93.3 | 7.97
2009-09-02 | 22:16 | 18.58 | 39 | 59 | 8.75 | 93.5 | 7.96
2009-09-02 | 22:31 | 18.46 | 39 | 60 | 8.78 | 93.6 | 7.96
2009-09-02 | 22:46 | 18.33 | 39 | 60 | 8.81 | 93.7 | 7.96
2009-09-02 | 23:01 | 18.18 | 39 | 60 | 8.81 | 93.4 | 7.96
2009-09-02 | 23:16 | 18.03 | 39 | 60 | 8.83 | 93.4 | 7.96
2009-09-02 | 23:31 | 17.88 | 39 | 60 | 8.88 | 93.6 | 7.96
2009-09-02 | 23:46 | 17.74 | 39 | 60 | 8.89 | 93.4 | 7.97
2009-09-03 | 0:01 | 17.6 | 39 | 60 | 8.95 | 93.7 | 7.97
2009-09-03 | 0:16 | 17.45 | 39 | 60 | 8.99 | 93.9 | 7.97
2009-09-03 | 0:31 | 17.31 | 39 | 60 | 8.99 | 93.7 | 7.97
2009-09-03 | 0:46 | 17.18 | 39 | 60 | 9.04 | 93.9 | 7.97
2009-09-03 | 1:01 | 17.05 | 39 | 60 | 9.08 | 94.1 | 7.97
2009-09-03 | 1:16 | 16.93 | 39 | 60 | 9.11 | 94.2 | 7.97
2009-09-03 | 1:31 | 16.81 | 39 | 60 | 9.11 | 93.9 | 7.97
2009-09-03 | 1:46 | 16.72 | 39 | 60 | 9.16 | 94.2 | 7.97
2009-09-03 | 2:01 | 16.62 | 39 | 61 | 9.18 | 94.3 | 7.98
2009-09-03 | 2:16 | 16.55 | 39 | 61 | 9.21 | 94.4 | 7.98
2009-09-03 | 2:31 | 16.45 | 39 | 61 | 9.24 | 94.5 | 7.99
2009-09-03 | 2:46 | 16.31 | 40 | 61 | 9.28 | 94.7 | 7.99
2009-09-03 | 3:01 | 16.21 | 40 | 61 | 9.31 | 94.8 | 7.99
2009-09-03 | 3:16 | 16.1 | 40 | 62 | 9.34 | 94.9 | 8
2009-09-03 | 3:31 | 15.99 | 40 | 62 | 9.38 | 95 | 8
2009-09-03 | 3:46 | 15.89 | 40 | 62 | 9.41 | 95.1 | 8.01
2009-09-03 | 4:01 | 15.78 | 40 | 63 | 9.43 | 95.1 | 8.01
2009-09-03 | 4:16 | 15.71 | 41 | 63 | 9.45 | 95.2 | 8.01
2009-09-03 | 4:31 | 15.62 | 41 | 64 | 9.48 | 95.2 | 8.02
2009-09-03 | 4:46 | 15.51 | 41 | 64 | 9.48 | 95 | 8.02
2009-09-03 | 5:01 | 15.4 | 42 | 65 | 9.49 | 95 | 8.03
2009-09-03 | 5:16 | 15.31 | 42 | 65 | 9.55 | 95.3 | 8.03
2009-09-03 | 5:31 | 15.2 | 42 | 66 | 9.58 | 95.4 | 8.04
2009-09-03 | 5:46 | 15.1 | 43 | 66 | 9.61 | 95.5 | 8.05
2009-09-03 | 6:01 | 14.98 | 43 | 67 | 9.62 | 95.3 | 8.05
2009-09-03 | 6:16 | 14.86 | 44 | 68 | 9.64 | 95.3 | 8.05
2009-09-03 | 6:31 | 14.73 | 44 | 68 | 9.7 | 95.6 | 8.06
2009-09-03 | 6:46 | 14.6 | 44 | 68 | 9.75 | 95.8 | 8.07
2009-09-03 | 7:01 | 14.46 | 44 | 68 | 9.75 | 95.5 | 8.06
2009-09-03 | 7:16 | 14.33 | 44 | 68 | 9.81 | 95.9 | 8.07
2009-09-03 | 7:31 | 14.21 | 44 | 68 | 9.88 | 96.4 | 8.09
2009-09-03 | 7:46 | 14.13 | 44 | 67 | 9.97 | 97 | 8.1
2009-09-03 | 8:01 | 14.05 | 44 | 66 | 9.97 | 96.8 | 8.1
2009-09-03 | 8:16 | 13.99 | 43 | 64 | 9.98 | 96.9 | 8.1
2009-09-03 | 8:31 | 13.95 | 42 | 63 | 10.04 | 97.3 | 8.1
2009-09-03 | 8:46 | 13.93 | 41 | 61 | 10.04 | 97.3 | 8.1
2009-09-03 | 9:01 | 13.93 | 40 | 60 | 10.09 | 97.8 | 8.11
2009-09-03 | 9:16 | 13.92 | 39 | 59 | 10.13 | 98.1 | 8.11
2009-09-03 | 9:31 | 13.93 | 39 | 59 | 10.16 | 98.4 | 8.12
2009-09-03 | 9:46 | 13.95 | 38 | 58 | 10.16 | 98.4 | 8.13
2009-09-03 | 10:01 | 14 | 38 | 57 | 10.19 | 98.9 | 8.15
2009-09-03 | 10:16 | 14.06 | 37 | 57 | 10.17 | 98.8 | 8.16
2009-09-03 | 10:31 | 14.1 | 37 | 57 | 10.22 | 99.4 | 8.18
2009-09-03 | 10:46 | 14.16 | 37 | 56 | 10.25 | 99.8 | 8.2
2009-09-03 | 11:01 | 14.29 | 37 | 56 | 10.26 | 100.2 | 8.21
2009-09-03 | 11:16 | 14.43 | 36 | 56 | 10.25 | 100.4 | 8.23
2009-09-03 | 11:31 | 14.59 | 36 | 56 | 10.19 | 100.2 | 8.25
2009-09-03 | 11:46 | 14.81 | 36 | 55 | 10.16 | 100.4 | 8.24
2009-09-03 | 12:01 | 15.04 | 36 | 55 | 10.12 | 100.4 | 8.25
2009-09-03 | 12:16 | 15.31 | 36 | 55 | 10.05 | 100.3 | 8.25
2009-09-03 | 12:31 | 15.64 | 36 | 55 | 9.98 | 100.3 | 8.25
2009-09-03 | 12:46 | 15.97 | 36 | 55 | 9.86 | 99.9 | 8.26
2009-09-03 | 13:01 | 16.24 | 36 | 55 | 9.82 | 100 | 8.27
2009-09-03 | 13:16 | 16.64 | 36 | 55 | 9.75 | 100.2 | 8.29
2009-09-03 | 13:31 | 16.95 | 36 | 55 | 9.69 | 100.1 | 8.31
2009-09-03 | 13:46 | 17.28 | 36 | 55 | 9.58 | 99.7 | 8.32
2009-09-03 | 14:01 | 17.71 | 36 | 55 | 9.51 | 99.9 | 8.33
2009-09-03 | 14:16 | 18.14 | 36 | 55 | 9.41 | 99.7 | 8.34
2009-09-03 | 14:31 | 18.5 | 36 | 55 | 9.3 | 99.3 | 8.34
2009-09-03 | 14:46 | 18.88 | 36 | 55 | 9.24 | 99.3 | 8.34
2009-09-03 | 15:01 | 19.14 | 36 | 56 | 9.18 | 99.2 | 8.35
2009-09-03 | 15:16 | 19.33 | 36 | 56 | 9.11 | 98.8 | 8.35
2009-09-03 | 15:31 | 19.58 | 36 | 56 | 9.06 | 98.8 | 8.36
2009-09-03 | 15:46 | 19.74 | 36 | 56 | 9.01 | 98.6 | 8.35
2009-09-03 | 16:01 | 19.8 | 36 | 56 | 8.99 | 98.5 | 8.35
2009-09-03 | 16:16 | 19.81 | 36 | 56 | 8.98 | 98.4 | 8.34
2009-09-03 | 16:31 | 19.72 | 36 | 56 | 9 | 98.4 | 8.33
2009-09-03 | 16:46 | 19.65 | 36 | 56 | 9.01 | 98.4 | 8.33
2009-09-03 | 17:01 | 19.54 | 37 | 56 | 9 | 98.1 | 8.3
2009-09-03 | 17:16 | 19.42 | 37 | 56 | 9 | 97.8 | 8.29
2009-09-03 | 17:31 | 19.27 | 37 | 56 | 8.95 | 97 | 8.25
2009-09-03 | 17:46 | 19.12 | 37 | 57 | 8.96 | 96.8 | 8.21
2009-09-03 | 18:01 | 18.98 | 37 | 57 | 8.94 | 96.3 | 8.18
2009-09-03 | 18:16 | 18.86 | 37 | 57 | 8.94 | 96.1 | 8.16
2009-09-03 | 18:31 | 18.78 | 37 | 57 | 8.93 | 95.9 | 8.14
2009-09-03 | 18:46 | 18.71 | 37 | 57 | 8.95 | 95.9 | 8.12
2009-09-03 | 19:01 | 18.61 | 37 | 57 | 8.93 | 95.5 | 8.1
2009-09-03 | 19:16 | 18.54 | 37 | 57 | 8.93 | 95.4 | 8.08
2009-09-03 | 19:31 | 18.45 | 37 | 57 | 8.92 | 95.1 | 8.06
2009-09-03 | 19:46 | 18.38 | 37 | 57 | 8.91 | 94.9 | 8.04
2009-09-03 | 20:01 | 18.31 | 37 | 57 | 8.91 | 94.8 | 8.02
2009-09-03 | 20:16 | 18.2 | 37 | 58 | 8.9 | 94.4 | 8.01
2009-09-03 | 20:31 | 18.07 | 37 | 58 | 8.91 | 94.2 | 7.99
2009-09-03 | 20:46 | 17.99 | 37 | 58 | 8.92 | 94.3 | 7.99
2009-09-03 | 21:01 | 17.86 | 38 | 58 | 8.98 | 94.5 | 7.98
2009-09-03 | 21:16 | 17.72 | 38 | 58 | 9.02 | 94.7 | 7.98
2009-09-03 | 21:31 | 17.56 | 38 | 58 | 9.06 | 94.9 | 7.98
2009-09-03 | 21:46 | 17.41 | 38 | 58 | 9.07 | 94.6 | 7.97
2009-09-03 | 22:01 | 17.24 | 38 | 58 | 9.1 | 94.6 | 7.97
2009-09-03 | 22:16 | 17.06 | 38 | 58 | 9.16 | 94.9 | 7.97
2009-09-03 | 22:31 | 16.88 | 38 | 58 | 9.21 | 95 | 7.96
2009-09-03 | 22:46 | 16.7 | 38 | 58 | 9.25 | 95.2 | 7.96
2009-09-03 | 23:01 | 16.54 | 38 | 58 | 9.29 | 95.2 | 7.96
2009-09-03 | 23:16 | 16.37 | 38 | 58 | 9.33 | 95.3 | 7.96
2009-09-03 | 23:31 | 16.21 | 38 | 58 | 9.37 | 95.4 | 7.96
2009-09-03 | 23:46 | 16.05 | 38 | 58 | 9.41 | 95.4 | 7.96
2009-09-04 | 0:01 | 15.89 | 38 | 58 | 9.45 | 95.5 | 7.96
2009-09-04 | 0:16 | 15.71 | 38 | 58 | 9.49 | 95.5 | 7.96
2009-09-04 | 0:31 | 15.53 | 38 | 58 | 9.5 | 95.3 | 7.96
2009-09-04 | 0:46 | 15.35 | 38 | 58 | 9.52 | 95.2 | 7.96
2009-09-04 | 1:01 | 15.17 | 38 | 58 | 9.56 | 95.2 | 7.95
2009-09-04 | 1:16 | 14.99 | 38 | 58 | 9.6 | 95.2 | 7.95
2009-09-04 | 1:31 | 14.81 | 38 | 58 | 9.64 | 95.2 | 7.95
2009-09-04 | 1:46 | 14.64 | 38 | 58 | 9.68 | 95.3 | 7.96
2009-09-04 | 2:01 | 14.48 | 38 | 58 | 9.75 | 95.6 | 7.96
2009-09-04 | 2:16 | 14.31 | 38 | 59 | 9.8 | 95.8 | 7.96
2009-09-04 | 2:31 | 14.17 | 38 | 59 | 9.81 | 95.6 | 7.95
2009-09-04 | 2:46 | 14.02 | 38 | 59 | 9.84 | 95.5 | 7.96
2009-09-04 | 3:01 | 13.88 | 39 | 60 | 9.9 | 95.8 | 7.97
2009-09-04 | 3:16 | 13.74 | 39 | 60 | 9.91 | 95.6 | 7.97
2009-09-04 | 3:31 | 13.59 | 39 | 60 | 9.94 | 95.6 | 7.97
2009-09-04 | 3:46 | 13.46 | 39 | 60 | 9.96 | 95.6 | 7.97
2009-09-04 | 4:01 | 13.32 | 39 | 61 | 10 | 95.6 | 7.97
2009-09-04 | 4:16 | 13.19 | 39 | 61 | 10.06 | 95.9 | 7.97
2009-09-04 | 4:31 | 13.06 | 40 | 61 | 10.11 | 96.1 | 7.98
2009-09-04 | 4:46 | 12.94 | 40 | 61 | 10.12 | 95.9 | 7.98
2009-09-04 | 5:01 | 12.83 | 40 | 62 | 10.18 | 96.2 | 7.97
2009-09-04 | 5:16 | 12.71 | 40 | 62 | 10.19 | 96.1 | 7.98
2009-09-04 | 5:31 | 12.62 | 40 | 62 | 10.21 | 96.1 | 7.98
2009-09-04 | 5:46 | 12.51 | 40 | 62 | 10.23 | 96.1 | 7.99
2009-09-04 | 6:01 | 12.41 | 41 | 63 | 10.3 | 96.4 | 7.99
2009-09-04 | 6:16 | 12.31 | 41 | 63 | 10.34 | 96.6 | 7.99
2009-09-04 | 6:31 | 12.21 | 41 | 63 | 10.38 | 96.8 | 7.99
2009-09-04 | 6:46 | 12.11 | 41 | 64 | 10.42 | 96.9 | 8
2009-09-04 | 7:01 | 12.03 | 41 | 64 | 10.47 | 97.2 | 8.01
2009-09-04 | 7:16 | 11.97 | 42 | 64 | 10.52 | 97.5 | 8.03
2009-09-04 | 7:31 | 11.9 | 42 | 65 | 10.52 | 97.4 | 8.04
2009-09-04 | 7:46 | 11.84 | 42 | 65 | 10.54 | 97.4 | 8.06
2009-09-04 | 8:01 | 11.8 | 42 | 66 | 10.61 | 98 | 8.07
2009-09-04 | 8:16 | 11.76 | 43 | 66 | 10.67 | 98.5 | 8.09
2009-09-04 | 8:31 | 11.74 | 43 | 66 | 10.71 | 98.8 | 8.11
2009-09-04 | 8:46 | 11.74 | 43 | 67 | 10.74 | 99.1 | 8.13
2009-09-04 | 9:01 | 11.78 | 43 | 67 | 10.78 | 99.5 | 8.15
2009-09-04 | 9:16 | 11.84 | 43 | 67 | 10.8 | 99.8 | 8.17
2009-09-04 | 9:31 | 11.93 | 44 | 68 | 10.8 | 100 | 8.2
2009-09-04 | 9:46 | 12.06 | 44 | 68 | 10.76 | 100 | 8.22
2009-09-04 | 10:01 | 12.13 | 44 | 68 | 10.74 | 100 | 8.24
2009-09-04 | 10:16 | 12.23 | 44 | 69 | 10.73 | 100.1 | 8.26
2009-09-04 | 10:31 | 12.35 | 45 | 69 | 10.72 | 100.2 | 8.27
2009-09-04 | 10:46 | 12.42 | 45 | 69 | 10.75 | 100.7 | 8.3
2009-09-04 | 11:01 | 12.49 | 45 | 70 | 10.72 | 100.6 | 8.31
2009-09-04 | 11:16 | 12.63 | 45 | 70 | 10.74 | 101.1 | 8.32
2009-09-04 | 11:31 | 12.82 | 45 | 70 | 10.73 | 101.4 | 8.35
2009-09-04 | 11:46 | 13 | 45 | 70 | 10.71 | 101.6 | 8.36
2009-09-04 | 12:01 | 13.2 | 45 | 69 | 10.63 | 101.3 | 8.39
2009-09-04 | 12:16 | 13.47 | 45 | 69 | 10.55 | 101.2 | 8.4
2009-09-04 | 12:31 | 13.73 | 45 | 68 | 10.48 | 101.1 | 8.41
2009-09-04 | 12:46 | 14.06 | 44 | 68 | 10.45 | 101.6 | 8.43
2009-09-04 | 13:01 | 14.31 | 44 | 67 | 10.39 | 101.5 | 8.44
2009-09-04 | 13:16 | 14.53 | 44 | 66 | 10.33 | 101.4 | 8.44
2009-09-04 | 13:31 | 14.8 | 43 | 66 | 10.27 | 101.4 | 8.44
2009-09-04 | 13:46 | 15.13 | 43 | 65 | 10.21 | 101.5 | 8.47
2009-09-04 | 14:01 | 15.31 | 43 | 65 | 10.11 | 100.9 | 8.47
2009-09-04 | 14:16 | 15.65 | 42 | 65 | 10.02 | 100.8 | 8.47
2009-09-04 | 14:31 | 16.13 | 42 | 64 | 9.94 | 100.9 | 8.47
2009-09-04 | 14:46 | 16.48 | 42 | 64 | 9.81 | 100.4 | 8.47
2009-09-04 | 15:01 | 16.75 | 42 | 64 | 9.75 | 100.4 | 8.46
2009-09-04 | 15:16 | 17.1 | 41 | 64 | 9.68 | 100.3 | 8.46
2009-09-04 | 15:31 | 17.28 | 41 | 64 | 9.58 | 99.7 | 8.45
2009-09-04 | 15:46 | 17.36 | 41 | 63 | 9.55 | 99.5 | 8.45
2009-09-04 | 16:01 | 17.45 | 41 | 63 | 9.47 | 99 | 8.44
2009-09-04 | 16:16 | 17.51 | 41 | 63 | 9.47 | 99 | 8.43
2009-09-04 | 16:31 | 17.53 | 41 | 63 | 9.46 | 99 | 8.42
2009-09-04 | 16:46 | 17.52 | 41 | 63 | 9.43 | 98.6 | 8.4
2009-09-04 | 17:01 | 17.56 | 41 | 63 | 9.38 | 98.1 | 8.38
2009-09-04 | 17:16 | 17.51 | 41 | 63 | 9.34 | 97.7 | 8.34
2009-09-04 | 17:31 | 17.45 | 41 | 63 | 9.35 | 97.7 | 8.31
2009-09-04 | 17:46 | 17.42 | 41 | 63 | 9.33 | 97.4 | 8.28
2009-09-04 | 18:01 | 17.42 | 41 | 63 | 9.29 | 96.9 | 8.26
2009-09-04 | 18:16 | 17.38 | 41 | 63 | 9.26 | 96.5 | 8.23
2009-09-04 | 18:31 | 17.36 | 41 | 63 | 9.2 | 96 | 8.19
2009-09-04 | 18:46 | 17.33 | 41 | 64 | 9.17 | 95.5 | 8.17
2009-09-04 | 19:01 | 17.27 | 41 | 64 | 9.18 | 95.5 | 8.14
2009-09-04 | 19:16 | 17.22 | 41 | 64 | 9.17 | 95.3 | 8.12
2009-09-04 | 19:31 | 17.17 | 42 | 64 | 9.13 | 94.8 | 8.09
2009-09-04 | 19:46 | 17.1 | 42 | 64 | 9.14 | 94.7 | 8.07
2009-09-04 | 20:01 | 17.01 | 42 | 64 | 9.14 | 94.6 | 8.06
2009-09-04 | 20:16 | 16.92 | 42 | 64 | 9.12 | 94.2 | 8.05
2009-09-04 | 20:31 | 16.84 | 42 | 64 | 9.13 | 94.1 | 8.04
2009-09-04 | 20:46 | 16.77 | 42 | 65 | 9.13 | 94.1 | 8.02
2009-09-04 | 21:01 | 16.7 | 42 | 65 | 9.17 | 94.3 | 8.02
2009-09-04 | 21:16 | 16.61 | 42 | 65 | 9.17 | 94.1 | 8.02
2009-09-04 | 21:31 | 16.53 | 42 | 65 | 9.18 | 94.1 | 8.01
2009-09-04 | 21:46 | 16.43 | 42 | 65 | 9.2 | 94.1 | 8.01
2009-09-04 | 22:01 | 16.34 | 42 | 65 | 9.22 | 94.1 | 8.01
2009-09-04 | 22:16 | 16.24 | 43 | 66 | 9.24 | 94.1 | 8.01
2009-09-04 | 22:31 | 16.15 | 43 | 66 | 9.3 | 94.5 | 8
2009-09-04 | 22:46 | 16.06 | 43 | 66 | 9.33 | 94.6 | 8.01
2009-09-04 | 23:01 | 15.98 | 43 | 66 | 9.32 | 94.4 | 8.01
2009-09-04 | 23:16 | 15.88 | 43 | 66 | 9.37 | 94.7 | 8.01
2009-09-04 | 23:31 | 15.78 | 43 | 66 | 9.4 | 94.8 | 8.01
2009-09-04 | 23:46 | 15.69 | 43 | 66 | 9.39 | 94.5 | 8.01
2009-09-05 | 0:01 | 15.59 | 43 | 67 | 9.4 | 94.4 | 8.01
2009-09-05 | 0:16 | 15.49 | 43 | 67 | 9.42 | 94.4 | 8.01
2009-09-05 | 0:31 | 15.38 | 43 | 67 | 9.44 | 94.4 | 8.01
2009-09-05 | 0:46 | 15.27 | 43 | 67 | 9.49 | 94.7 | 8.01
2009-09-05 | 1:01 | 15.14 | 43 | 67 | 9.5 | 94.5 | 8.01
2009-09-05 | 1:16 | 15.01 | 43 | 67 | 9.55 | 94.8 | 8.01
2009-09-05 | 1:31 | 14.88 | 43 | 67 | 9.6 | 94.9 | 8.01
2009-09-05 | 1:46 | 14.75 | 43 | 67 | 9.63 | 95 | 8.01
2009-09-05 | 2:01 | 14.63 | 44 | 67 | 9.66 | 95.1 | 8.01
2009-09-05 | 2:16 | 14.51 | 44 | 67 | 9.66 | 94.8 | 8.01
2009-09-05 | 2:31 | 14.37 | 44 | 67 | 9.68 | 94.7 | 8.01
2009-09-05 | 2:46 | 14.25 | 44 | 67 | 9.72 | 94.9 | 8.01
2009-09-05 | 3:01 | 14.14 | 44 | 67 | 9.76 | 95 | 8.01
2009-09-05 | 3:16 | 14.01 | 44 | 67 | 9.81 | 95.2 | 8.01
2009-09-05 | 3:31 | 13.91 | 44 | 67 | 9.81 | 95 | 8.02
2009-09-05 | 3:46 | 13.81 | 44 | 68 | 9.86 | 95.3 | 8.02
2009-09-05 | 4:01 | 13.72 | 44 | 68 | 9.86 | 95.1 | 8.02
2009-09-05 | 4:16 | 13.63 | 44 | 68 | 9.91 | 95.4 | 8.02
2009-09-05 | 4:31 | 13.55 | 44 | 68 | 9.95 | 95.6 | 8.02
2009-09-05 | 4:46 | 13.45 | 44 | 68 | 9.95 | 95.4 | 8.03
2009-09-05 | 5:01 | 13.37 | 44 | 68 | 10.01 | 95.7 | 8.03
2009-09-05 | 5:16 | 13.29 | 44 | 69 | 10.04 | 95.9 | 8.03
2009-09-05 | 5:31 | 13.22 | 45 | 69 | 10.06 | 96 | 8.04
2009-09-05 | 5:46 | 13.17 | 45 | 70 | 10.08 | 96 | 8.05
2009-09-05 | 6:01 | 13.14 | 46 | 71 | 10.06 | 95.8 | 8.06
2009-09-05 | 6:16 | 13.11 | 46 | 71 | 10.09 | 96 | 8.06
2009-09-05 | 6:31 | 13.09 | 46 | 72 | 10.11 | 96.1 | 8.07
2009-09-05 | 6:46 | 13.06 | 47 | 73 | 10.09 | 95.9 | 8.07
2009-09-05 | 7:01 | 13.04 | 48 | 74 | 10.09 | 95.8 | 8.09
2009-09-05 | 7:16 | 13.01 | 48 | 75 | 10.12 | 96.1 | 8.1
2009-09-05 | 7:31 | 12.97 | 49 | 76 | 10.15 | 96.2 | 8.12
2009-09-05 | 7:46 | 12.92 | 50 | 77 | 10.18 | 96.4 | 8.13
2009-09-05 | 8:01 | 12.86 | 50 | 78 | 10.26 | 97.1 | 8.14
2009-09-05 | 8:16 | 12.81 | 50 | 78 | 10.32 | 97.6 | 8.16
2009-09-05 | 8:31 | 12.76 | 50 | 77 | 10.33 | 97.5 | 8.17
2009-09-05 | 8:46 | 12.73 | 50 | 76 | 10.4 | 98.1 | 8.18
2009-09-05 | 9:01 | 12.72 | 49 | 75 | 10.46 | 98.6 | 8.2
2009-09-05 | 9:16 | 12.73 | 49 | 74 | 10.46 | 98.7 | 8.21
2009-09-05 | 9:31 | 12.78 | 48 | 72 | 10.51 | 99.3 | 8.23
2009-09-05 | 9:46 | 12.84 | 47 | 71 | 10.53 | 99.6 | 8.25
2009-09-05 | 10:01 | 12.91 | 46 | 70 | 10.5 | 99.4 | 8.26
2009-09-05 | 10:16 | 12.97 | 45 | 68 | 10.47 | 99.3 | 8.27
2009-09-05 | 10:31 | 13.06 | 44 | 67 | 10.5 | 99.8 | 8.28
2009-09-05 | 10:46 | 13.2 | 43 | 65 | 10.51 | 100.2 | 8.29
2009-09-05 | 11:01 | 13.31 | 42 | 64 | 10.48 | 100.2 | 8.29
2009-09-05 | 11:16 | 13.47 | 42 | 63 | 10.46 | 100.3 | 8.3
2009-09-05 | 11:31 | 13.71 | 41 | 62 | 10.43 | 100.5 | 8.3
2009-09-05 | 11:46 | 13.89 | 40 | 61 | 10.36 | 100.3 | 8.31
2009-09-05 | 12:01 | 14.06 | 40 | 61 | 10.28 | 99.9 | 8.32
2009-09-05 | 12:16 | 14.26 | 39 | 60 | 10.27 | 100.2 | 8.32
2009-09-05 | 12:31 | 14.37 | 39 | 60 | 10.2 | 99.8 | 8.32
2009-09-05 | 12:46 | 14.49 | 39 | 59 | 10.17 | 99.7 | 8.3
2009-09-05 | 13:01 | 14.68 | 39 | 59 | 10.14 | 99.9 | 8.31
2009-09-05 | 13:16 | 14.8 | 39 | 59 | 10.11 | 99.8 | 8.3
2009-09-05 | 13:31 | 15 | 38 | 59 | 10.06 | 99.7 | 8.31
2009-09-05 | 13:46 | 15.17 | 38 | 59 | 9.97 | 99.3 | 8.31
2009-09-05 | 14:01 | 15.27 | 38 | 59 | 9.93 | 99 | 8.29
2009-09-05 | 14:16 | 15.51 | 38 | 59 | 9.86 | 98.9 | 8.29
2009-09-05 | 14:31 | 15.59 | 38 | 59 | 9.84 | 98.8 | 8.3
2009-09-05 | 14:46 | 15.64 | 38 | 59 | 9.82 | 98.7 | 8.28
2009-09-05 | 15:01 | 15.73 | 38 | 59 | 9.73 | 98.1 | 8.27
2009-09-05 | 15:16 | 15.71 | 38 | 59 | 9.74 | 98.1 | 8.25
2009-09-05 | 15:31 | 15.85 | 38 | 59 | 9.79 | 98.9 | 8.26
2009-09-05 | 15:46 | 16.04 | 38 | 59 | 9.79 | 99.3 | 8.28
2009-09-05 | 16:01 | 16.23 | 38 | 59 | 9.72 | 98.9 | 8.29
2009-09-05 | 16:16 | 16.34 | 38 | 59 | 9.66 | 98.6 | 8.28
2009-09-05 | 16:31 | 16.43 | 38 | 59 | 9.67 | 98.9 | 8.28
2009-09-05 | 16:46 | 16.52 | 38 | 59 | 9.63 | 98.7 | 8.29
2009-09-05 | 17:01 | 16.48 | 38 | 59 | 9.61 | 98.3 | 8.27
2009-09-05 | 17:16 | 16.4 | 38 | 59 | 9.57 | 97.8 | 8.24
2009-09-05 | 17:31 | 16.33 | 38 | 59 | 9.51 | 97.1 | 8.2
2009-09-05 | 17:46 | 16.19 | 38 | 59 | 9.53 | 97 | 8.17
2009-09-05 | 18:01 | 16.03 | 39 | 59 | 9.55 | 96.8 | 8.15
2009-09-05 | 18:16 | 15.89 | 39 | 59 | 9.59 | 96.9 | 8.14
2009-09-05 | 18:31 | 15.78 | 39 | 60 | 9.6 | 96.8 | 8.13
2009-09-05 | 18:46 | 15.69 | 39 | 59 | 9.61 | 96.7 | 8.12
2009-09-05 | 19:01 | 15.59 | 39 | 60 | 9.63 | 96.7 | 8.11
2009-09-05 | 19:16 | 15.51 | 39 | 60 | 9.59 | 96.2 | 8.09
2009-09-05 | 19:31 | 15.45 | 39 | 60 | 9.59 | 96.1 | 8.07
2009-09-05 | 19:46 | 15.39 | 39 | 60 | 9.56 | 95.6 | 8.06
2009-09-05 | 20:01 | 15.35 | 39 | 60 | 9.53 | 95.3 | 8.04
2009-09-05 | 20:16 | 15.27 | 39 | 60 | 9.54 | 95.1 | 8.02
2009-09-05 | 20:31 | 15.21 | 39 | 60 | 9.58 | 95.4 | 8.01
2009-09-05 | 20:46 | 15.12 | 39 | 60 | 9.58 | 95.2 | 8.02
2009-09-05 | 21:01 | 15.05 | 39 | 60 | 9.62 | 95.5 | 8.01
2009-09-05 | 21:16 | 14.96 | 39 | 60 | 9.62 | 95.3 | 8
2009-09-05 | 21:31 | 14.84 | 39 | 60 | 9.64 | 95.3 | 8
2009-09-05 | 21:46 | 14.72 | 39 | 60 | 9.69 | 95.6 | 8
2009-09-05 | 22:01 | 14.6 | 39 | 60 | 9.71 | 95.5 | 8
2009-09-05 | 22:16 | 14.48 | 39 | 60 | 9.76 | 95.7 | 8
2009-09-05 | 22:31 | 14.36 | 39 | 60 | 9.77 | 95.5 | 7.99
2009-09-05 | 22:46 | 14.25 | 39 | 60 | 9.82 | 95.8 | 7.99
2009-09-05 | 23:01 | 14.14 | 39 | 60 | 9.86 | 95.9 | 8
2009-09-05 | 23:16 | 14.02 | 39 | 61 | 9.89 | 96 | 7.99
2009-09-05 | 23:31 | 13.91 | 39 | 61 | 9.92 | 96.1 | 7.99
2009-09-05 | 23:46 | 13.79 | 39 | 61 | 9.93 | 95.9 | 7.99
2009-09-06 | 0:01 | 13.68 | 39 | 61 | 9.94 | 95.8 | 7.99
2009-09-06 | 0:16 | 13.56 | 39 | 61 | 10 | 96.1 | 7.99
2009-09-06 | 0:31 | 13.45 | 40 | 61 | 10.01 | 96 | 7.99
2009-09-06 | 0:46 | 13.34 | 40 | 61 | 10.06 | 96.2 | 7.99
2009-09-06 | 1:01 | 13.24 | 40 | 61 | 10.07 | 96.1 | 7.98
2009-09-06 | 1:16 | 13.15 | 40 | 61 | 10.11 | 96.3 | 7.99
2009-09-06 | 1:31 | 13.06 | 40 | 61 | 10.12 | 96.2 | 7.99
2009-09-06 | 1:46 | 12.97 | 40 | 61 | 10.17 | 96.4 | 7.99
2009-09-06 | 2:01 | 12.87 | 40 | 62 | 10.17 | 96.2 | 8
2009-09-06 | 2:16 | 12.77 | 40 | 62 | 10.22 | 96.5 | 7.99
2009-09-06 | 2:31 | 12.67 | 40 | 62 | 10.26 | 96.6 | 7.99
2009-09-06 | 2:46 | 12.58 | 40 | 62 | 10.28 | 96.7 | 8
2009-09-06 | 3:01 | 12.5 | 40 | 62 | 10.31 | 96.8 | 8
2009-09-06 | 3:16 | 12.4 | 40 | 62 | 10.34 | 96.8 | 8
2009-09-06 | 3:31 | 12.29 | 41 | 63 | 10.34 | 96.6 | 8
2009-09-06 | 3:46 | 12.18 | 41 | 63 | 10.36 | 96.5 | 8
2009-09-06 | 4:01 | 12.05 | 41 | 63 | 10.39 | 96.6 | 8
2009-09-06 | 4:16 | 11.95 | 41 | 63 | 10.46 | 96.9 | 8
2009-09-06 | 4:31 | 11.86 | 41 | 64 | 10.5 | 97.1 | 8.01
2009-09-06 | 4:46 | 11.78 | 41 | 64 | 10.5 | 96.9 | 8.01
2009-09-06 | 5:01 | 11.72 | 42 | 64 | 10.54 | 97.2 | 8.01
2009-09-06 | 5:16 | 11.65 | 42 | 65 | 10.54 | 97 | 8.02
2009-09-06 | 5:31 | 11.58 | 42 | 65 | 10.58 | 97.2 | 8.02
2009-09-06 | 5:46 | 11.5 | 42 | 65 | 10.61 | 97.4 | 8.02
2009-09-06 | 6:01 | 11.41 | 43 | 66 | 10.64 | 97.4 | 8.03
2009-09-06 | 6:16 | 11.34 | 43 | 66 | 10.66 | 97.5 | 8.04
2009-09-06 | 6:31 | 11.29 | 43 | 67 | 10.68 | 97.5 | 8.04
2009-09-06 | 6:46 | 11.21 | 43 | 67 | 10.71 | 97.6 | 8.04
2009-09-06 | 7:01 | 11.15 | 44 | 67 | 10.75 | 97.8 | 8.05
2009-09-06 | 7:16 | 11.07 | 44 | 67 | 10.76 | 97.7 | 8.06
2009-09-06 | 7:31 | 11 | 44 | 67 | 10.8 | 97.9 | 8.07
2009-09-06 | 7:46 | 10.94 | 44 | 68 | 10.82 | 98 | 8.08
2009-09-06 | 8:01 | 10.89 | 44 | 67 | 10.89 | 98.5 | 8.08
2009-09-06 | 8:16 | 10.88 | 44 | 67 | 10.93 | 98.9 | 8.09
2009-09-06 | 8:31 | 10.84 | 44 | 67 | 10.92 | 98.7 | 8.1
2009-09-06 | 8:46 | 10.84 | 44 | 67 | 10.98 | 99.2 | 8.11
2009-09-06 | 9:01 | 10.85 | 44 | 67 | 11.01 | 99.5 | 8.13
2009-09-06 | 9:16 | 10.87 | 43 | 66 | 11.03 | 99.7 | 8.13
2009-09-06 | 9:31 | 10.89 | 43 | 66 | 10.99 | 99.5 | 8.13
2009-09-06 | 9:46 | 10.91 | 43 | 65 | 10.99 | 99.5 | 8.13
2009-09-06 | 10:01 | 10.97 | 42 | 65 | 10.99 | 99.6 | 8.15
2009-09-06 | 10:16 | 11.05 | 42 | 64 | 11.04 | 100.2 | 8.15
2009-09-06 | 10:31 | 11.03 | 42 | 64 | 11.06 | 100.4 | 8.16
2009-09-06 | 10:46 | 11.15 | 41 | 63 | 11.05 | 100.6 | 8.17
2009-09-06 | 11:01 | 11.25 | 41 | 63 | 11.06 | 100.9 | 8.18
2009-09-06 | 11:16 | 11.42 | 41 | 62 | 11.03 | 101 | 8.19
2009-09-06 | 11:31 | 11.46 | 41 | 62 | 11 | 100.8 | 8.19
2009-09-06 | 11:46 | 11.53 | 40 | 61 | 10.96 | 100.6 | 8.2
2009-09-06 | 12:01 | 11.65 | 40 | 61 | 10.91 | 100.4 | 8.21
2009-09-06 | 12:16 | 11.92 | 40 | 61 | 10.87 | 100.6 | 8.21
2009-09-06 | 12:31 | 12.22 | 39 | 60 | 10.8 | 100.7 | 8.23
2009-09-06 | 12:46 | 12.37 | 39 | 60 | 10.77 | 100.8 | 8.23
2009-09-06 | 13:01 | 12.52 | 39 | 60 | 10.74 | 100.8 | 8.23
2009-09-06 | 13:16 | 12.67 | 39 | 60 | 10.7 | 100.8 | 8.23
2009-09-06 | 13:31 | 12.87 | 39 | 59 | 10.69 | 101.2 | 8.25
2009-09-06 | 13:46 | 13.28 | 39 | 59 | 10.62 | 101.4 | 8.26
2009-09-06 | 14:01 | 13.66 | 39 | 59 | 10.53 | 101.4 | 8.26
2009-09-06 | 14:16 | 13.98 | 39 | 59 | 10.13 | 98.3 | 8.26
2009-09-06 | 14:31 | 14.32 | 39 | 59 | 10.34 | 101 | 8.28
2009-09-06 | 14:46 | 14.38 | 39 | 59 | 10.28 | 100.6 | 8.27
2009-09-06 | 15:01 | 14.43 | 39 | 59 | 10.21 | 100 | 8.25
2009-09-06 | 15:16 | 14.74 | 39 | 59 | 10.19 | 100.5 | 8.27
2009-09-06 | 15:31 | 14.77 | 39 | 59 | 10.22 | 100.8 | 8.27
2009-09-06 | 15:46 | 14.88 | 39 | 59 | 10.17 | 100.6 | 8.28
2009-09-06 | 16:01 | 15.02 | 39 | 59 | 10.1 | 100.2 | 8.29
2009-09-06 | 16:16 | 14.97 | 38 | 59 | 10.14 | 100.5 | 8.28
2009-09-06 | 16:31 | 14.96 | 38 | 59 | 10.09 | 100 | 8.28
2009-09-06 | 16:46 | 14.98 | 38 | 59 | 10.11 | 100.2 | 8.28
2009-09-06 | 17:01 | 14.91 | 38 | 58 | 10.07 | 99.7 | 8.25
2009-09-06 | 17:16 | 14.81 | 38 | 58 | 10.05 | 99.3 | 8.22
2009-09-06 | 17:31 | 14.74 | 38 | 59 | 10.03 | 98.9 | 8.19
2009-09-06 | 17:46 | 14.67 | 38 | 58 | 9.98 | 98.2 | 8.16
2009-09-06 | 18:01 | 14.58 | 38 | 58 | 9.93 | 97.6 | 8.15
2009-09-06 | 18:16 | 14.53 | 38 | 58 | 9.99 | 98.1 | 8.12
2009-09-06 | 18:31 | 14.43 | 38 | 58 | 9.97 | 97.6 | 8.11
2009-09-06 | 18:46 | 14.34 | 38 | 58 | 9.99 | 97.7 | 8.09
2009-09-06 | 19:01 | 14.28 | 38 | 59 | 9.96 | 97.3 | 8.08
2009-09-06 | 19:16 | 14.19 | 38 | 59 | 9.99 | 97.4 | 8.07
2009-09-06 | 19:31 | 14.08 | 38 | 59 | 10.01 | 97.3 | 8.05
2009-09-06 | 19:46 | 13.99 | 38 | 59 | 9.99 | 96.9 | 8.03
2009-09-06 | 20:01 | 13.91 | 38 | 59 | 10.02 | 97.1 | 8.02
2009-09-06 | 20:16 | 13.83 | 38 | 59 | 10.01 | 96.8 | 8.01
2009-09-06 | 20:31 | 13.78 | 38 | 59 | 10.05 | 97 | 8.01
2009-09-06 | 20:46 | 13.72 | 38 | 59 | 10.04 | 96.8 | 8.01
2009-09-06 | 21:01 | 13.63 | 39 | 59 | 10.08 | 97.1 | 8
2009-09-06 | 21:16 | 13.52 | 39 | 59 | 10.12 | 97.2 | 7.99
2009-09-06 | 21:31 | 13.42 | 39 | 60 | 10.12 | 96.9 | 7.99
2009-09-06 | 21:46 | 13.31 | 39 | 60 | 10.17 | 97.2 | 7.99
2009-09-06 | 22:01 | 13.19 | 39 | 60 | 10.21 | 97.3 | 7.99
2009-09-06 | 22:16 | 13.08 | 39 | 60 | 10.24 | 97.4 | 7.99
2009-09-06 | 22:31 | 12.98 | 39 | 60 | 10.27 | 97.4 | 7.99
2009-09-06 | 22:46 | 12.86 | 39 | 60 | 10.27 | 97.2 | 7.99
2009-09-06 | 23:01 | 12.75 | 39 | 60 | 10.29 | 97.1 | 7.99
2009-09-06 | 23:16 | 12.66 | 39 | 60 | 10.31 | 97.1 | 7.99
2009-09-06 | 23:31 | 12.57 | 39 | 60 | 10.37 | 97.4 | 7.99
2009-09-06 | 23:46 | 12.47 | 39 | 60 | 10.4 | 97.6 | 7.99
2009-09-07 | 0:01 | 12.37 | 39 | 60 | 10.4 | 97.3 | 7.99
2009-09-07 | 0:16 | 12.27 | 39 | 60 | 10.42 | 97.3 | 7.99
2009-09-07 | 0:31 | 12.18 | 39 | 60 | 10.44 | 97.3 | 7.99
2009-09-07 | 0:46 | 12.09 | 39 | 61 | 10.5 | 97.6 | 7.99
2009-09-07 | 1:01 | 11.99 | 39 | 61 | 10.5 | 97.5 | 7.99
2009-09-07 | 1:16 | 11.87 | 39 | 61 | 10.53 | 97.4 | 7.99
2009-09-07 | 1:31 | 11.76 | 39 | 61 | 10.59 | 97.8 | 7.98
2009-09-07 | 1:46 | 11.66 | 39 | 61 | 10.63 | 97.9 | 7.99
2009-09-07 | 2:01 | 11.57 | 40 | 61 | 10.66 | 97.9 | 7.99
2009-09-07 | 2:16 | 11.48 | 40 | 61 | 10.68 | 98 | 7.99
2009-09-07 | 2:31 | 11.41 | 40 | 61 | 10.67 | 97.7 | 7.99
2009-09-07 | 2:46 | 11.37 | 40 | 62 | 10.71 | 98 | 7.99
2009-09-07 | 3:01 | 11.31 | 40 | 62 | 10.74 | 98.1 | 8
2009-09-07 | 3:16 | 11.23 | 40 | 62 | 10.73 | 97.8 | 8
2009-09-07 | 3:31 | 11.16 | 40 | 62 | 10.78 | 98.1 | 8
2009-09-07 | 3:46 | 11.08 | 40 | 62 | 10.77 | 97.9 | 8
2009-09-07 | 4:01 | 10.99 | 41 | 63 | 10.79 | 97.8 | 8
2009-09-07 | 4:16 | 10.89 | 41 | 63 | 10.81 | 97.8 | 8
2009-09-07 | 4:31 | 10.79 | 41 | 63 | 10.84 | 97.9 | 7.99
2009-09-07 | 4:46 | 10.7 | 41 | 63 | 10.87 | 97.9 | 8.01
2009-09-07 | 5:01 | 10.6 | 41 | 64 | 10.9 | 97.9 | 8.01
2009-09-07 | 5:16 | 10.51 | 41 | 64 | 10.97 | 98.3 | 8.01
2009-09-07 | 5:31 | 10.43 | 42 | 64 | 11 | 98.4 | 8.01
2009-09-07 | 5:46 | 10.35 | 42 | 64 | 10.99 | 98.2 | 8.01
2009-09-07 | 6:01 | 10.28 | 42 | 65 | 11.04 | 98.5 | 8.02
2009-09-07 | 6:16 | 10.18 | 42 | 65 | 11.05 | 98.3 | 8.02
2009-09-07 | 6:31 | 10.07 | 42 | 65 | 11.07 | 98.3 | 8.02
2009-09-07 | 6:46 | 9.96 | 42 | 65 | 11.14 | 98.6 | 8.02
2009-09-07 | 7:01 | 9.84 | 42 | 65 | 11.17 | 98.6 | 8.04
2009-09-07 | 7:16 | 9.73 | 43 | 66 | 11.26 | 99.1 | 8.03
2009-09-07 | 7:31 | 9.62 | 43 | 66 | 11.32 | 99.4 | 8.05
2009-09-07 | 7:46 | 9.52 | 43 | 66 | 11.37 | 99.6 | 8.06
2009-09-07 | 8:01 | 9.44 | 43 | 66 | 11.41 | 99.8 | 8.06
2009-09-07 | 8:16 | 9.39 | 43 | 66 | 11.42 | 99.7 | 8.07
2009-09-07 | 8:31 | 9.34 | 43 | 66 | 11.48 | 100.1 | 8.08
2009-09-07 | 8:46 | 9.33 | 43 | 66 | 11.5 | 100.3 | 8.09
2009-09-07 | 9:01 | 9.35 | 43 | 66 | 11.49 | 100.2 | 8.1
2009-09-07 | 9:16 | 9.37 | 43 | 66 | 11.54 | 100.7 | 8.11
2009-09-07 | 9:31 | 9.42 | 43 | 66 | 11.52 | 100.7 | 8.12
2009-09-07 | 9:46 | 9.5 | 43 | 66 | 11.53 | 100.9 | 8.13
2009-09-07 | 10:01 | 9.57 | 43 | 66 | 11.58 | 101.6 | 8.15
2009-09-07 | 10:16 | 9.64 | 43 | 66 | 11.6 | 101.9 | 8.16
2009-09-07 | 10:31 | 9.71 | 43 | 66 | 11.58 | 101.9 | 8.18
2009-09-07 | 10:46 | 9.81 | 43 | 66 | 11.6 | 102.4 | 8.19
2009-09-07 | 11:01 | 9.92 | 43 | 66 | 11.56 | 102.3 | 8.2
2009-09-07 | 11:16 | 10.04 | 43 | 66 | 11.53 | 102.3 | 8.2
2009-09-07 | 11:31 | 10.2 | 43 | 66 | 11.53 | 102.7 | 8.21
2009-09-07 | 11:46 | 10.39 | 43 | 67 | 11.46 | 102.4 | 8.22
2009-09-07 | 12:01 | 10.67 | 43 | 67 | 11.37 | 102.3 | 8.23
2009-09-07 | 12:16 | 11.02 | 43 | 67 | 11.31 | 102.6 | 8.25
2009-09-07 | 12:31 | 11.4 | 43 | 67 | 11.21 | 102.6 | 8.26
2009-09-07 | 12:46 | 11.83 | 43 | 67 | 11.1 | 102.6 | 8.27
2009-09-07 | 13:01 | 12.3 | 43 | 67 | 10.95 | 102.4 | 8.28
2009-09-07 | 13:16 | 12.84 | 43 | 67 | 10.81 | 102.3 | 8.31
2009-09-07 | 13:31 | 13.26 | 43 | 67 | 10.67 | 101.9 | 8.32
2009-09-07 | 13:46 | 13.7 | 43 | 67 | 10.55 | 101.7 | 8.34
2009-09-07 | 14:01 | 13.94 | 43 | 67 | 10.43 | 101.1 | 8.35
2009-09-07 | 14:16 | 14.34 | 43 | 67 | 10.37 | 101.3 | 8.36
2009-09-07 | 14:31 | 14.62 | 43 | 67 | 10.3 | 101.3 | 8.36
2009-09-07 | 14:46 | 15.06 | 43 | 67 | 10.17 | 101 | 8.36
2009-09-07 | 15:01 | 15.56 | 43 | 67 | 10.05 | 100.9 | 8.39
2009-09-07 | 15:16 | 15.92 | 43 | 67 | 9.97 | 100.8 | 8.39
2009-09-07 | 15:31 | 16.26 | 43 | 67 | 9.89 | 100.8 | 8.4
2009-09-07 | 15:46 | 16.56 | 43 | 67 | 9.8 | 100.5 | 8.39
2009-09-07 | 16:01 | 16.76 | 43 | 67 | 9.72 | 100.1 | 8.4
2009-09-07 | 16:16 | 16.84 | 43 | 67 | 9.69 | 100 | 8.39
2009-09-07 | 16:31 | 16.89 | 43 | 67 | 9.65 | 99.6 | 8.39
2009-09-07 | 16:46 | 16.82 | 43 | 67 | 9.65 | 99.5 | 8.36
2009-09-07 | 17:01 | 16.6 | 43 | 67 | 9.63 | 98.8 | 8.33
2009-09-07 | 17:16 | 16.35 | 43 | 66 | 9.63 | 98.3 | 8.3
2009-09-07 | 17:31 | 16.11 | 43 | 66 | 9.68 | 98.3 | 8.26
2009-09-07 | 17:46 | 15.91 | 43 | 66 | 9.7 | 98.1 | 8.23
2009-09-07 | 18:01 | 15.75 | 43 | 66 | 9.71 | 97.9 | 8.21
2009-09-07 | 18:16 | 15.59 | 43 | 66 | 9.7 | 97.4 | 8.18
2009-09-07 | 18:31 | 15.49 | 43 | 66 | 9.73 | 97.5 | 8.16
2009-09-07 | 18:46 | 15.4 | 43 | 66 | 9.74 | 97.4 | 8.15
2009-09-07 | 19:01 | 15.33 | 43 | 66 | 9.71 | 97 | 8.13
2009-09-07 | 19:16 | 15.26 | 43 | 65 | 9.73 | 97.1 | 8.11
2009-09-07 | 19:31 | 15.18 | 43 | 65 | 9.74 | 97 | 8.1
2009-09-07 | 19:46 | 15.1 | 43 | 65 | 9.71 | 96.5 | 8.09
2009-09-07 | 20:01 | 15.05 | 43 | 65 | 9.7 | 96.3 | 8.08
2009-09-07 | 20:16 | 14.98 | 43 | 66 | 9.74 | 96.5 | 8.07
2009-09-07 | 20:31 | 14.9 | 43 | 66 | 9.76 | 96.6 | 8.06
2009-09-07 | 20:46 | 14.77 | 43 | 66 | 9.77 | 96.4 | 8.05
2009-09-07 | 21:01 | 14.64 | 43 | 66 | 9.82 | 96.7 | 8.05
2009-09-07 | 21:16 | 14.49 | 43 | 66 | 9.86 | 96.8 | 8.05
2009-09-07 | 21:31 | 14.34 | 43 | 66 | 9.91 | 96.8 | 8.04
2009-09-07 | 21:46 | 14.18 | 43 | 66 | 9.95 | 96.9 | 8.04
2009-09-07 | 22:01 | 14.02 | 43 | 66 | 9.96 | 96.7 | 8.04
2009-09-07 | 22:16 | 13.86 | 43 | 66 | 10 | 96.7 | 8.04
2009-09-07 | 22:31 | 13.7 | 43 | 66 | 10.03 | 96.7 | 8.04
2009-09-07 | 22:46 | 13.55 | 43 | 66 | 10.1 | 97.1 | 8.04
2009-09-07 | 23:01 | 13.4 | 43 | 66 | 10.12 | 96.9 | 8.04
2009-09-07 | 23:16 | 13.22 | 43 | 66 | 10.16 | 96.9 | 8.04
2009-09-07 | 23:31 | 13 | 43 | 66 | 10.21 | 96.9 | 8.04
2009-09-07 | 23:46 | 12.79 | 43 | 66 | 10.3 | 97.3 | 8.05
2009-09-08 | 0:01 | 12.58 | 43 | 66 | 10.34 | 97.2 | 8.04
2009-09-08 | 0:16 | 12.37 | 43 | 66 | 10.42 | 97.5 | 8.03
2009-09-08 | 0:31 | 12.18 | 43 | 66 | 10.47 | 97.6 | 8.03
2009-09-08 | 0:46 | 11.99 | 43 | 66 | 10.53 | 97.7 | 8.03
2009-09-08 | 1:01 | 11.81 | 43 | 66 | 10.56 | 97.6 | 8.02
2009-09-08 | 1:16 | 11.63 | 43 | 66 | 10.61 | 97.7 | 8.03
2009-09-08 | 1:31 | 11.47 | 43 | 66 | 10.68 | 97.9 | 8.03
2009-09-08 | 1:46 | 11.32 | 43 | 66 | 10.72 | 98 | 8.03
2009-09-08 | 2:01 | 11.17 | 43 | 66 | 10.76 | 98 | 8.03
2009-09-08 | 2:16 | 11.02 | 43 | 66 | 10.78 | 97.8 | 8.02
2009-09-08 | 2:31 | 10.89 | 43 | 66 | 10.82 | 97.8 | 8.03
2009-09-08 | 2:46 | 10.75 | 43 | 66 | 10.88 | 98.1 | 8.02
2009-09-08 | 3:01 | 10.61 | 43 | 66 | 10.9 | 98 | 8.02
2009-09-08 | 3:16 | 10.48 | 43 | 66 | 10.96 | 98.2 | 8.03
2009-09-08 | 3:31 | 10.35 | 43 | 66 | 11.01 | 98.3 | 8.02
2009-09-08 | 3:46 | 10.22 | 43 | 66 | 11.05 | 98.4 | 8.02
2009-09-08 | 4:01 | 10.1 | 43 | 66 | 11.06 | 98.2 | 8.03
2009-09-08 | 4:16 | 9.98 | 43 | 66 | 11.12 | 98.5 | 8.02
2009-09-08 | 4:31 | 9.86 | 43 | 67 | 11.16 | 98.5 | 8.03
2009-09-08 | 4:46 | 9.73 | 43 | 67 | 11.21 | 98.7 | 8.02
2009-09-08 | 5:01 | 9.61 | 43 | 67 | 11.21 | 98.4 | 8.02
2009-09-08 | 5:16 | 9.48 | 43 | 67 | 11.24 | 98.4 | 8.02
2009-09-08 | 5:31 | 9.36 | 43 | 67 | 11.31 | 98.7 | 8.02
2009-09-08 | 5:46 | 9.24 | 43 | 67 | 11.36 | 98.8 | 8.02
2009-09-08 | 6:01 | 9.12 | 44 | 67 | 11.4 | 98.9 | 8.02
2009-09-08 | 6:16 | 9.01 | 44 | 67 | 11.43 | 98.9 | 8.02
2009-09-08 | 6:31 | 8.9 | 44 | 67 | 11.45 | 98.8 | 8.02
2009-09-08 | 6:46 | 8.79 | 44 | 67 | 11.48 | 98.8 | 8.03
2009-09-08 | 7:01 | 8.68 | 44 | 68 | 11.56 | 99.2 | 8.03
2009-09-08 | 7:16 | 8.59 | 44 | 68 | 11.59 | 99.3 | 8.04
2009-09-08 | 7:31 | 8.5 | 44 | 68 | 11.66 | 99.7 | 8.05
2009-09-08 | 7:46 | 8.43 | 44 | 68 | 11.69 | 99.7 | 8.06
2009-09-08 | 8:01 | 8.37 | 44 | 69 | 11.72 | 99.9 | 8.08
2009-09-08 | 8:16 | 8.33 | 45 | 69 | 11.75 | 100 | 8.08
2009-09-08 | 8:31 | 8.31 | 45 | 69 | 11.81 | 100.5 | 8.09
2009-09-08 | 8:46 | 8.31 | 45 | 69 | 11.84 | 100.7 | 8.1
2009-09-08 | 9:01 | 8.35 | 45 | 70 | 11.82 | 100.7 | 8.11
2009-09-08 | 9:16 | 8.41 | 45 | 70 | 11.86 | 101.1 | 8.12
2009-09-08 | 9:31 | 8.49 | 45 | 70 | 11.86 | 101.4 | 8.14
2009-09-08 | 9:46 | 8.62 | 45 | 70 | 11.86 | 101.7 | 8.15
2009-09-08 | 10:01 | 8.75 | 46 | 70 | 11.83 | 101.7 | 8.17
2009-09-08 | 10:16 | 8.87 | 46 | 70 | 11.81 | 101.8 | 8.18
2009-09-08 | 10:31 | 8.98 | 46 | 71 | 11.8 | 102.1 | 8.2
2009-09-08 | 10:46 | 9.12 | 46 | 71 | 11.79 | 102.3 | 8.21
2009-09-08 | 11:01 | 9.25 | 46 | 71 | 11.77 | 102.4 | 8.22
2009-09-08 | 11:16 | 9.39 | 46 | 71 | 11.79 | 103 | 8.22
2009-09-08 | 11:30 | 9.54 | 46 | 71 | 11.39 | 99.8 | 8.22
2009-09-08 | 11:46 | 9.73 | 46 | 72 | 11.7 | 103 | 8.24
2009-09-08 | 12:01 | 9.97 | 46 | 72 | 11.61 | 102.8 | 8.25
2009-09-08 | 12:16 | 10.26 | 47 | 72 | 11.55 | 103 | 8.26
2009-09-08 | 12:31 | 10.65 | 47 | 72 | 11.42 | 102.7 | 8.27
2009-09-08 | 12:46 | 11.08 | 47 | 72 | 11.29 | 102.6 | 8.29
2009-09-08 | 13:01 | 11.57 | 47 | 72 | 11.13 | 102.3 | 8.3
2009-09-08 | 13:16 | 12.09 | 47 | 72 | 11.02 | 102.5 | 8.32
2009-09-08 | 13:31 | 12.62 | 47 | 73 | 10.85 | 102.2 | 8.33
2009-09-08 | 13:46 | 13.14 | 47 | 73 | 10.72 | 102.1 | 8.36
2009-09-08 | 14:01 | 13.66 | 47 | 73 | 10.58 | 101.9 | 8.38
2009-09-08 | 14:16 | 14.17 | 47 | 73 | 10.49 | 102.1 | 8.4
2009-09-08 | 14:31 | 14.71 | 48 | 73 | 10.37 | 102.2 | 8.43
2009-09-08 | 14:46 | 15.24 | 48 | 74 | 10.2 | 101.7 | 8.44
2009-09-08 | 15:01 | 15.69 | 48 | 74 | 10.1 | 101.7 | 8.46
2009-09-08 | 15:16 | 16.09 | 48 | 74 | 9.99 | 101.4 | 8.47
2009-09-08 | 15:31 | 16.43 | 48 | 74 | 9.89 | 101.2 | 8.48
2009-09-08 | 15:46 | 16.76 | 48 | 74 | 9.79 | 100.8 | 8.48
2009-09-08 | 16:01 | 17.02 | 48 | 74 | 9.71 | 100.5 | 8.48
2009-09-08 | 16:16 | 17.2 | 48 | 74 | 9.61 | 99.8 | 8.47
2009-09-08 | 16:31 | 17.25 | 48 | 74 | 9.56 | 99.4 | 8.46
2009-09-08 | 16:46 | 17.19 | 48 | 74 | 9.53 | 99 | 8.44
2009-09-08 | 17:01 | 17.06 | 48 | 75 | 9.52 | 98.6 | 8.4
2009-09-08 | 17:16 | 16.9 | 49 | 75 | 9.55 | 98.7 | 8.38
2009-09-08 | 17:31 | 16.72 | 49 | 75 | 9.53 | 98.1 | 8.35
2009-09-08 | 17:46 | 16.57 | 49 | 75 | 9.55 | 97.9 | 8.32
2009-09-08 | 18:01 | 16.45 | 49 | 75 | 9.55 | 97.7 | 8.28
2009-09-08 | 18:16 | 16.31 | 49 | 75 | 9.53 | 97.2 | 8.26
2009-09-08 | 18:31 | 16.19 | 49 | 75 | 9.53 | 96.9 | 8.23
2009-09-08 | 18:46 | 16.08 | 49 | 75 | 9.56 | 97 | 8.22
2009-09-08 | 19:01 | 16 | 49 | 75 | 9.54 | 96.7 | 8.2
2009-09-08 | 19:16 | 15.92 | 49 | 75 | 9.57 | 96.8 | 8.18
2009-09-08 | 19:31 | 15.86 | 49 | 75 | 9.53 | 96.3 | 8.17
2009-09-08 | 19:46 | 15.79 | 49 | 75 | 9.52 | 96 | 8.15
2009-09-08 | 20:01 | 15.7 | 49 | 75 | 9.52 | 95.9 | 8.14
2009-09-08 | 20:16 | 15.6 | 49 | 75 | 9.53 | 95.8 | 8.13
2009-09-08 | 20:31 | 15.51 | 49 | 75 | 9.58 | 96 | 8.12
2009-09-08 | 20:46 | 15.39 | 49 | 75 | 9.61 | 96.2 | 8.11
2009-09-08 | 21:01 | 15.27 | 49 | 75 | 9.62 | 96 | 8.11
2009-09-08 | 21:16 | 15.14 | 49 | 75 | 9.67 | 96.2 | 8.1
2009-09-08 | 21:31 | 15.01 | 48 | 75 | 9.71 | 96.3 | 8.1
2009-09-08 | 21:46 | 14.87 | 48 | 74 | 9.75 | 96.4 | 8.1
2009-09-08 | 22:01 | 14.73 | 48 | 74 | 9.78 | 96.5 | 8.1
2009-09-08 | 22:16 | 14.57 | 48 | 74 | 9.8 | 96.3 | 8.09
2009-09-08 | 22:31 | 14.4 | 48 | 74 | 9.87 | 96.6 | 8.1
2009-09-08 | 22:46 | 14.22 | 48 | 74 | 9.92 | 96.7 | 8.09
2009-09-08 | 23:01 | 14.04 | 48 | 74 | 9.97 | 96.8 | 8.1
2009-09-08 | 23:16 | 13.85 | 48 | 74 | 9.99 | 96.7 | 8.1
2009-09-08 | 23:31 | 13.67 | 48 | 74 | 10.06 | 96.9 | 8.09
2009-09-08 | 23:46 | 13.49 | 48 | 74 | 10.08 | 96.7 | 8.09
2009-09-09 | 0:01 | 13.31 | 48 | 74 | 10.12 | 96.8 | 8.09
2009-09-09 | 0:16 | 13.13 | 48 | 74 | 10.2 | 97.1 | 8.09
2009-09-09 | 0:31 | 12.95 | 48 | 74 | 10.23 | 97 | 8.09
2009-09-09 | 0:46 | 12.78 | 48 | 74 | 10.3 | 97.3 | 8.09
2009-09-09 | 1:01 | 12.6 | 48 | 74 | 10.36 | 97.4 | 8.08
2009-09-09 | 1:16 | 12.43 | 48 | 74 | 10.38 | 97.3 | 8.09
2009-09-09 | 1:31 | 12.27 | 48 | 74 | 10.41 | 97.2 | 8.08
2009-09-09 | 1:46 | 12.11 | 48 | 73 | 10.46 | 97.3 | 8.08
2009-09-09 | 2:01 | 11.95 | 48 | 73 | 10.53 | 97.6 | 8.08
2009-09-09 | 2:16 | 11.8 | 48 | 73 | 10.58 | 97.8 | 8.08
2009-09-09 | 2:31 | 11.66 | 48 | 73 | 10.63 | 97.8 | 8.08
2009-09-09 | 2:46 | 11.52 | 47 | 73 | 10.66 | 97.9 | 8.07
2009-09-09 | 3:01 | 11.39 | 47 | 73 | 10.71 | 98 | 8.07
2009-09-09 | 3:16 | 11.26 | 47 | 73 | 10.75 | 98 | 8.07
2009-09-09 | 3:31 | 11.13 | 47 | 73 | 10.78 | 98.1 | 8.07
2009-09-09 | 3:46 | 11.01 | 47 | 73 | 10.79 | 97.9 | 8.07
2009-09-09 | 4:01 | 10.89 | 47 | 72 | 10.85 | 98.1 | 8.06
2009-09-09 | 4:16 | 10.77 | 47 | 72 | 10.86 | 98 | 8.06
2009-09-09 | 4:31 | 10.67 | 47 | 72 | 10.89 | 98 | 8.06
2009-09-09 | 4:46 | 10.56 | 47 | 72 | 10.92 | 98 | 8.06
2009-09-09 | 5:01 | 10.46 | 47 | 72 | 10.98 | 98.3 | 8.06
2009-09-09 | 5:16 | 10.37 | 47 | 72 | 11.02 | 98.4 | 8.06
2009-09-09 | 5:31 | 10.28 | 47 | 72 | 11.05 | 98.5 | 8.06
2009-09-09 | 5:46 | 10.19 | 47 | 75 | 11.05 | 98.3 | 8.06
2009-09-09 | 6:01 | 10.11 | 49 | 75 | 11.07 | 98.3 | 8.06
2009-09-09 | 6:16 | 10.03 | 49 | 74 | 11.12 | 98.6 | 8.06
2009-09-09 | 6:31 | 9.95 | 48 | 74 | 11.13 | 98.5 | 8.06
2009-09-09 | 6:46 | 9.88 | 48 | 73 | 11.18 | 98.8 | 8.06
2009-09-09 | 7:01 | 9.81 | 48 | 73 | 11.2 | 98.8 | 8.07
2009-09-09 | 7:16 | 9.75 | 48 | 74 | 11.24 | 99 | 8.08
2009-09-09 | 7:31 | 9.72 | 48 | 74 | 11.32 | 99.7 | 8.09
2009-09-09 | 7:46 | 9.7 | 48 | 74 | 11.38 | 100.1 | 8.11
2009-09-09 | 8:01 | 9.69 | 48 | 74 | 11.42 | 100.5 | 8.13
2009-09-09 | 8:16 | 9.7 | 48 | 74 | 11.47 | 100.9 | 8.15
2009-09-09 | 8:31 | 9.73 | 48 | 74 | 11.48 | 101.1 | 8.17
2009-09-09 | 8:46 | 9.8 | 48 | 74 | 11.51 | 101.5 | 8.2
2009-09-09 | 9:01 | 9.87 | 48 | 74 | 11.57 | 102.2 | 8.23
2009-09-09 | 9:16 | 9.99 | 48 | 75 | 11.58 | 102.6 | 8.25
2009-09-09 | 9:31 | 10.13 | 49 | 75 | 11.57 | 102.8 | 8.28
2009-09-09 | 9:46 | 10.26 | 49 | 75 | 11.52 | 102.7 | 8.3
2009-09-09 | 10:01 | 10.42 | 49 | 76 | 11.53 | 103.2 | 8.32
2009-09-09 | 10:16 | 10.57 | 49 | 76 | 11.5 | 103.2 | 8.35
2009-09-09 | 10:31 | 10.73 | 49 | 76 | 11.47 | 103.4 | 8.36
2009-09-09 | 10:46 | 10.9 | 49 | 76 | 11.41 | 103.3 | 8.38
2009-09-09 | 11:01 | 11.07 | 50 | 76 | 11.39 | 103.5 | 8.4
2009-09-09 | 11:16 | 11.28 | 50 | 76 | 11.33 | 103.4 | 8.39
2009-09-09 | 11:31 | 11.41 | 50 | 77 | 11.27 | 103.2 | 8.38
2009-09-09 | 11:46 | 11.69 | 50 | 77 | 11.24 | 103.6 | 8.4
2009-09-09 | 12:01 | 11.9 | 50 | 77 | 11.22 | 103.9 | 8.41
2009-09-09 | 12:16 | 12.1 | 50 | 77 | 11.15 | 103.7 | 8.43
2009-09-09 | 12:31 | 12.37 | 50 | 77 | 11.1 | 103.9 | 8.45
2009-09-09 | 12:46 | 12.77 | 50 | 77 | 10.99 | 103.8 | 8.45
2009-09-09 | 13:01 | 13.1 | 50 | 78 | 10.89 | 103.6 | 8.45
2009-09-09 | 13:16 | 13.49 | 50 | 78 | 10.75 | 103.2 | 8.45
2009-09-09 | 13:31 | 13.95 | 50 | 78 | 10.61 | 102.9 | 8.45
2009-09-09 | 13:46 | 14.42 | 51 | 78 | 10.48 | 102.7 | 8.45
2009-09-09 | 14:01 | 14.83 | 51 | 78 | 10.37 | 102.5 | 8.47
2009-09-09 | 14:16 | 15.22 | 51 | 78 | 10.31 | 102.7 | 8.48
2009-09-09 | 14:31 | 15.67 | 51 | 78 | 10.2 | 102.6 | 8.48
2009-09-09 | 14:46 | 16.12 | 51 | 78 | 10.06 | 102.2 | 8.49
2009-09-09 | 15:01 | 16.54 | 51 | 79 | 10 | 102.4 | 8.51
2009-09-09 | 15:16 | 16.84 | 52 | 79 | 9.88 | 101.9 | 8.52
2009-09-09 | 15:31 | 17.01 | 52 | 80 | 9.82 | 101.6 | 8.54
2009-09-09 | 15:46 | 17.16 | 52 | 80 | 9.8 | 101.7 | 8.55
2009-09-09 | 16:01 | 17.38 | 52 | 80 | 9.75 | 101.7 | 8.56
2009-09-09 | 16:16 | 17.55 | 52 | 80 | 9.66 | 101.1 | 8.57
2009-09-09 | 16:31 | 17.66 | 52 | 79 | 9.6 | 100.7 | 8.54
2009-09-09 | 16:46 | 17.64 | 51 | 79 | 9.55 | 100.1 | 8.52
2009-09-09 | 17:01 | 17.56 | 51 | 79 | 9.57 | 100.2 | 8.49
2009-09-09 | 17:16 | 17.5 | 51 | 79 | 9.52 | 99.5 | 8.47
2009-09-09 | 17:31 | 17.45 | 51 | 79 | 9.51 | 99.3 | 8.44
2009-09-09 | 17:46 | 17.43 | 51 | 79 | 9.51 | 99.3 | 8.41
2009-09-09 | 18:01 | 17.42 | 51 | 79 | 9.43 | 98.5 | 8.38
2009-09-09 | 18:16 | 17.39 | 51 | 79 | 9.38 | 97.9 | 8.35
2009-09-09 | 18:31 | 17.36 | 51 | 79 | 9.39 | 97.9 | 8.32
2009-09-09 | 18:46 | 17.34 | 51 | 79 | 9.34 | 97.3 | 8.29
2009-09-09 | 19:01 | 17.31 | 51 | 79 | 9.3 | 96.8 | 8.25
2009-09-09 | 19:16 | 17.29 | 51 | 79 | 9.29 | 96.8 | 8.22
2009-09-09 | 19:31 | 17.25 | 51 | 80 | 9.26 | 96.3 | 8.2
2009-09-09 | 19:46 | 17.21 | 52 | 81 | 9.24 | 96 | 8.18
2009-09-09 | 20:01 | 17.16 | 52 | 80 | 9.27 | 96.3 | 8.17
2009-09-09 | 20:16 | 17.06 | 52 | 80 | 9.3 | 96.3 | 8.15
2009-09-09 | 20:31 | 16.94 | 52 | 81 | 9.3 | 96.2 | 8.14
2009-09-09 | 20:46 | 16.81 | 52 | 78 | 9.36 | 96.4 | 8.14
2009-09-09 | 21:01 | 16.65 | 51 | 78 | 9.4 | 96.5 | 8.13
2009-09-09 | 21:16 | 16.49 | 51 | 78 | 9.44 | 96.6 | 8.12
2009-09-09 | 21:31 | 16.33 | 51 | 78 | 9.45 | 96.5 | 8.11
2009-09-09 | 21:46 | 16.18 | 51 | 78 | 9.51 | 96.7 | 8.11
2009-09-09 | 22:01 | 16.02 | 51 | 80 | 9.55 | 96.8 | 8.1
2009-09-09 | 22:16 | 15.88 | 52 | 80 | 9.57 | 96.7 | 8.1
2009-09-09 | 22:31 | 15.75 | 52 | 76 | 9.62 | 96.9 | 8.1
2009-09-09 | 22:46 | 15.61 | 50 | 78 | 9.63 | 96.8 | 8.1
2009-09-09 | 23:01 | 15.48 | 50 | 77 | 9.69 | 97.1 | 8.09
2009-09-09 | 23:16 | 15.37 | 50 | 77 | 9.7 | 96.9 | 8.1
2009-09-09 | 23:31 | 15.26 | 50 | 77 | 9.71 | 96.9 | 8.09
2009-09-09 | 23:46 | 15.16 | 50 | 77 | 9.74 | 96.9 | 8.09
2009-09-10 | 0:01 | 15.06 | 50 | 76 | 9.76 | 96.9 | 8.09
2009-09-10 | 0:16 | 14.97 | 50 | 76 | 9.79 | 97 | 8.09
2009-09-10 | 0:31 | 14.88 | 50 | 76 | 9.81 | 97 | 8.09
2009-09-10 | 0:46 | 14.79 | 50 | 76 | 9.83 | 97.1 | 8.09
2009-09-10 | 1:01 | 14.7 | 49 | 76 | 9.88 | 97.4 | 8.09
2009-09-10 | 1:16 | 14.62 | 49 | 76 | 9.92 | 97.5 | 8.08
2009-09-10 | 1:31 | 14.53 | 49 | 77 | 9.91 | 97.3 | 8.09
2009-09-10 | 1:46 | 14.44 | 50 | 78 | 9.96 | 97.6 | 8.09
2009-09-10 | 2:01 | 14.34 | 51 | 78 | 9.97 | 97.4 | 8.09
2009-09-10 | 2:16 | 14.23 | 51 | 78 | 9.99 | 97.4 | 8.09
2009-09-10 | 2:31 | 14.1 | 51 | 76 | 10.05 | 97.7 | 8.09
2009-09-10 | 2:46 | 13.97 | 49 | 78 | 10.06 | 97.5 | 8.09
2009-09-10 | 3:01 | 13.85 | 50 | 77 | 10.09 | 97.6 | 8.09
2009-09-10 | 3:16 | 13.71 | 50 | 75 | 10.12 | 97.5 | 8.09
2009-09-10 | 3:31 | 13.59 | 49 | 75 | 10.18 | 97.9 | 8.09
2009-09-10 | 3:46 | 13.46 | 49 | 75 | 10.22 | 98 | 8.09
2009-09-10 | 4:01 | 13.33 | 49 | 77 | 10.23 | 97.8 | 8.08
2009-09-10 | 4:16 | 13.21 | 50 | 77 | 10.26 | 97.8 | 8.08
2009-09-10 | 4:31 | 13.08 | 50 | 77 | 10.32 | 98.1 | 8.08
2009-09-10 | 4:46 | 12.95 | 50 | 75 | 10.36 | 98.2 | 8.08
2009-09-10 | 5:01 | 12.82 | 49 | 76 | 10.37 | 98.1 | 8.08
2009-09-10 | 5:16 | 12.69 | 49 | 75 | 10.43 | 98.3 | 8.08
2009-09-10 | 5:31 | 12.56 | 49 | 75 | 10.45 | 98.2 | 8.07
2009-09-10 | 5:46 | 12.44 | 49 | 75 | 10.47 | 98.1 | 8.07
2009-09-10 | 6:01 | 12.33 | 48 | 75 | 10.5 | 98.2 | 8.08
2009-09-10 | 6:16 | 12.22 | 48 | 75 | 10.56 | 98.5 | 8.08
2009-09-10 | 6:31 | 12.1 | 48 | 75 | 10.58 | 98.4 | 8.08
2009-09-10 | 6:46 | 11.98 | 49 | 75 | 10.61 | 98.5 | 8.08
2009-09-10 | 7:01 | 11.87 | 49 | 75 | 10.69 | 98.9 | 8.09
2009-09-10 | 7:16 | 11.76 | 49 | 75 | 10.73 | 99 | 8.1
2009-09-10 | 7:31 | 11.66 | 49 | 75 | 10.77 | 99.2 | 8.1
2009-09-10 | 7:46 | 11.58 | 49 | 75 | 10.82 | 99.5 | 8.12
2009-09-10 | 8:01 | 11.49 | 49 | 75 | 10.9 | 100 | 8.12
2009-09-10 | 8:16 | 11.43 | 49 | 75 | 10.92 | 100 | 8.14
2009-09-10 | 8:31 | 11.37 | 49 | 75 | 10.98 | 100.4 | 8.15
2009-09-10 | 8:46 | 11.33 | 49 | 75 | 11.01 | 100.6 | 8.15
2009-09-10 | 9:01 | 11.33 | 49 | 76 | 11.04 | 100.9 | 8.17
2009-09-10 | 9:16 | 11.37 | 49 | 76 | 11.07 | 101.2 | 8.19
2009-09-10 | 9:31 | 11.44 | 49 | 76 | 11.08 | 101.5 | 8.2
2009-09-10 | 9:46 | 11.55 | 49 | 76 | 11.1 | 102 | 8.22
2009-09-10 | 10:01 | 11.67 | 50 | 77 | 11.11 | 102.3 | 8.23
2009-09-10 | 10:16 | 11.77 | 50 | 77 | 11.11 | 102.6 | 8.25
2009-09-10 | 10:31 | 11.89 | 50 | 78 | 11.12 | 103 | 8.27
2009-09-10 | 10:46 | 12 | 51 | 78 | 11.13 | 103.3 | 8.28
2009-09-10 | 11:01 | 12.11 | 51 | 79 | 11.1 | 103.3 | 8.29
2009-09-10 | 11:16 | 12.24 | 51 | 79 | 11.08 | 103.4 | 8.31
2009-09-10 | 11:31 | 12.37 | 51 | 79 | 11.09 | 103.8 | 8.31
2009-09-10 | 11:46 | 12.54 | 51 | 79 | 11.06 | 103.9 | 8.33
2009-09-10 | 12:01 | 12.74 | 51 | 79 | 10.99 | 103.7 | 8.35
2009-09-10 | 12:16 | 13.02 | 52 | 79 | 10.95 | 104 | 8.37
2009-09-10 | 12:31 | 13.39 | 52 | 79 | 10.84 | 103.7 | 8.39
2009-09-10 | 12:46 | 13.81 | 52 | 80 | 10.71 | 103.5 | 8.4
2009-09-10 | 13:01 | 14.28 | 52 | 80 | 10.58 | 103.3 | 8.41
2009-09-10 | 13:16 | 14.77 | 52 | 80 | 10.45 | 103.1 | 8.43
2009-09-10 | 13:31 | 15.27 | 52 | 80 | 10.37 | 103.4 | 8.45
2009-09-10 | 13:46 | 15.76 | 52 | 80 | 10.23 | 103.2 | 8.46
2009-09-10 | 14:01 | 16.23 | 52 | 80 | 10.15 | 103.3 | 8.48
2009-09-10 | 14:16 | 16.73 | 52 | 82 | 10.04 | 103.3 | 8.5
2009-09-10 | 14:31 | 17.26 | 53 | 82 | 9.88 | 102.8 | 8.52
2009-09-10 | 14:46 | 17.76 | 53 | 82 | 9.78 | 102.8 | 8.53
2009-09-10 | 15:01 | 18.21 | 53 | 81 | 9.67 | 102.6 | 8.55
2009-09-10 | 15:16 | 18.58 | 53 | 82 | 9.54 | 102 | 8.55
2009-09-10 | 15:31 | 18.89 | 53 | 82 | 9.44 | 101.6 | 8.56
2009-09-10 | 15:46 | 19.18 | 53 | 82 | 9.37 | 101.4 | 8.56
2009-09-10 | 16:01 | 19.4 | 53 | 82 | 9.33 | 101.4 | 8.56
2009-09-10 | 16:16 | 19.51 | 54 | 82 | 9.27 | 101 | 8.55
2009-09-10 | 16:31 | 19.52 | 54 | 83 | 9.21 | 100.3 | 8.53
2009-09-10 | 16:46 | 19.44 | 54 | 82 | 9.17 | 99.8 | 8.51
2009-09-10 | 17:01 | 19.31 | 54 | 82 | 9.16 | 99.3 | 8.47
2009-09-10 | 17:16 | 19.12 | 54 | 82 | 9.17 | 99.1 | 8.43
2009-09-10 | 17:31 | 18.94 | 54 | 82 | 9.15 | 98.5 | 8.39
2009-09-10 | 17:46 | 18.77 | 53 | 82 | 9.17 | 98.4 | 8.34
2009-09-10 | 18:01 | 18.61 | 53 | 82 | 9.17 | 98.1 | 8.32
2009-09-10 | 18:16 | 18.46 | 53 | 82 | 9.18 | 97.9 | 8.28
2009-09-10 | 18:31 | 18.34 | 53 | 81 | 9.19 | 97.8 | 8.26
2009-09-10 | 18:46 | 18.24 | 53 | 81 | 9.2 | 97.6 | 8.24
2009-09-10 | 19:01 | 18.15 | 52 | 80 | 9.2 | 97.5 | 8.22
2009-09-10 | 19:16 | 18.07 | 52 | 80 | 9.2 | 97.4 | 8.19
2009-09-10 | 19:31 | 18.01 | 52 | 78 | 9.16 | 96.8 | 8.18
2009-09-10 | 19:46 | 17.96 | 51 | 79 | 9.18 | 96.9 | 8.16
2009-09-10 | 20:01 | 17.9 | 51 | 79 | 9.16 | 96.5 | 8.15
2009-09-10 | 20:16 | 17.85 | 51 | 78 | 9.15 | 96.4 | 8.14
2009-09-10 | 20:31 | 17.78 | 51 | 76 | 9.17 | 96.4 | 8.13
2009-09-10 | 20:46 | 17.7 | 50 | 79 | 9.18 | 96.4 | 8.12
2009-09-10 | 21:01 | 17.62 | 51 | 77 | 9.2 | 96.5 | 8.12
2009-09-10 | 21:16 | 17.52 | 50 | 77 | 9.25 | 96.8 | 8.11
2009-09-10 | 21:31 | 17.41 | 50 | 77 | 9.29 | 96.9 | 8.11
2009-09-10 | 21:46 | 17.3 | 50 | 77 | 9.29 | 96.7 | 8.1
2009-09-10 | 22:01 | 17.17 | 50 | 77 | 9.31 | 96.7 | 8.11
2009-09-10 | 22:16 | 17.02 | 50 | 77 | 9.37 | 97 | 8.1
2009-09-10 | 22:31 | 16.87 | 50 | 76 | 9.39 | 96.9 | 8.1
2009-09-10 | 22:46 | 16.72 | 50 | 76 | 9.42 | 96.9 | 8.1
2009-09-10 | 23:01 | 16.58 | 50 | 77 | 9.48 | 97.2 | 8.09
2009-09-10 | 23:16 | 16.44 | 50 | 76 | 9.49 | 97.1 | 8.09
2009-09-10 | 23:31 | 16.3 | 49 | 77 | 9.55 | 97.4 | 8.09
2009-09-10 | 23:46 | 16.14 | 50 | 77 | 9.59 | 97.4 | 8.09
2009-09-11 | 0:01 | 15.99 | 50 | 77 | 9.64 | 97.6 | 8.09
2009-09-11 | 0:16 | 15.83 | 50 | 77 | 9.65 | 97.4 | 8.09
2009-09-11 | 0:31 | 15.67 | 50 | 77 | 9.71 | 97.7 | 8.09
2009-09-11 | 0:46 | 15.5 | 50 | 77 | 9.74 | 97.6 | 8.09
2009-09-11 | 1:01 | 15.32 | 50 | 76 | 9.77 | 97.6 | 8.09
2009-09-11 | 1:16 | 15.16 | 50 | 76 | 9.81 | 97.6 | 8.09
2009-09-11 | 1:31 | 14.99 | 50 | 76 | 9.88 | 97.9 | 8.09
2009-09-11 | 1:46 | 14.83 | 50 | 76 | 9.93 | 98.1 | 8.09
2009-09-11 | 2:01 | 14.66 | 50 | 76 | 9.98 | 98.3 | 8.09
2009-09-11 | 2:16 | 14.5 | 50 | 74 | 9.99 | 98 | 8.09
2009-09-11 | 2:31 | 14.34 | 48 | 74 | 10.05 | 98.3 | 8.08
2009-09-11 | 2:46 | 14.19 | 48 | 76 | 10.07 | 98.1 | 8.09
2009-09-11 | 3:01 | 14.05 | 49 | 76 | 10.1 | 98.1 | 8.08
2009-09-11 | 3:16 | 13.91 | 49 | 76 | 10.13 | 98.1 | 8.08
2009-09-11 | 3:31 | 13.78 | 49 | 76 | 10.17 | 98.2 | 8.08
2009-09-11 | 3:46 | 13.65 | 49 | 76 | 10.2 | 98.2 | 8.08
2009-09-11 | 4:01 | 13.52 | 50 | 76 | 10.23 | 98.2 | 8.08
2009-09-11 | 4:16 | 13.39 | 49 | 75 | 10.3 | 98.6 | 8.08
2009-09-11 | 4:31 | 13.26 | 49 | 75 | 10.34 | 98.7 | 8.08
2009-09-11 | 4:46 | 13.15 | 49 | 75 | 10.37 | 98.8 | 8.07
2009-09-11 | 5:01 | 13.03 | 49 | 76 | 10.4 | 98.8 | 8.07
2009-09-11 | 5:16 | 12.91 | 49 | 75 | 10.41 | 98.6 | 8.08
2009-09-11 | 5:31 | 12.8 | 49 | 76 | 10.46 | 98.8 | 8.07
2009-09-11 | 5:46 | 12.7 | 49 | 76 | 10.49 | 98.9 | 8.07
2009-09-11 | 6:01 | 12.6 | 49 | 75 | 10.5 | 98.7 | 8.07
2009-09-11 | 6:16 | 12.5 | 49 | 75 | 10.55 | 99 | 8.07
2009-09-11 | 6:31 | 12.4 | 49 | 74 | 10.56 | 98.9 | 8.07
2009-09-11 | 6:46 | 12.3 | 48 | 74 | 10.62 | 99.2 | 8.07
2009-09-11 | 7:01 | 12.21 | 48 | 74 | 10.67 | 99.5 | 8.08
2009-09-11 | 7:16 | 12.12 | 48 | 74 | 10.72 | 99.8 | 8.09
2009-09-11 | 7:31 | 12.05 | 48 | 75 | 10.76 | 100 | 8.1
2009-09-11 | 7:46 | 11.98 | 49 | 75 | 10.78 | 100 | 8.11
2009-09-11 | 8:01 | 11.93 | 49 | 76 | 10.8 | 100.1 | 8.12
2009-09-11 | 8:16 | 11.88 | 49 | 76 | 10.83 | 100.2 | 8.13
2009-09-11 | 8:31 | 11.84 | 49 | 76 | 10.89 | 100.7 | 8.14
2009-09-11 | 8:46 | 11.82 | 50 | 77 | 10.91 | 100.9 | 8.15
2009-09-11 | 9:01 | 11.82 | 50 | 77 | 10.91 | 100.8 | 8.16
2009-09-11 | 9:16 | 11.84 | 50 | 77 | 10.93 | 101 | 8.17
2009-09-11 | 9:31 | 11.88 | 50 | 78 | 10.98 | 101.6 | 8.19