## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA audited against Loner Pro Series, Sonde AB, NIST
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 11:30 |  | 11:30 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.08 |  | A
DO | In Situ |  | 9.85 |  | 10.09 |  | A
SpC | In Situ |  | 54.6 |  | 53 |  | A
Ph | In Situ |  | 8.23 |  | 8.4 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-05-19 12:10 | 17.2 | 8.41 | 10.06 | 104.5 | 52 | 33.8
Sonde | 2009-05-19 12:00 | 17.26 | 8.5 | 10.79 | 112.6 | 52 | 34
Grade |  | A | A | B |  | A
Audit 2 | 2009-05-22 17:59 | 16.4 | 8.95 | 10.48 | 107 | 54.4 | 35.1
Sonde | 2009-05-22 18:00 | 16.38 | 9.13 | 10.73 | 109.6 | 54 | 35
Grade |  | A | A | A |  | A
Audit 3 | 2009-05-31 11:00 | 15.9 | 7.56 | 9.76 | 98.8 | 41.6 | 27.3
Sonde | 2009-05-31 11:00 | 15.88 | 7.7 | 10.25 | 103.5 | 41 | 27
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA audited against Loner Pro Series, Sonde AB, NIST
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-31 |  | 2009-05-31 |  | A
Time | Computer set to atomic time |  | 12:33 |  | 12:43 |  | A
Temp | VWR NIST traceable surface water |  | 10.1 |  | 9.94 |  | A
DO | In Situ |  | 11.19 |  | 12.04 |  | B
SpC | In Situ |  | 29.6 |  | 29 |  | A
Ph | In Situ |  | 7.47 |  | 7.62 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner Pro Series
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 11:24 |  | 11:24 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17 |  | A
DO | Winkler titration of surface water |  | 10.09 |  | 9.85 |  | A
SpC | 148 µS/mL |  | 53 |  | 54.6 |  | A
Ph | pH 7 standard solution |  | 8.4 |  | 8.23 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-19 12:10 | 17.2 | 8.41 | 10.06 | 104.5 | 52 | 33.8
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 13:45 |  | 13:45 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.2 |  | A
DO | Winkler titration of surface water |  | 9.9 |  | 10.42 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 228 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner Pro Series
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 13:45 |  | 13:45 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.2 |  | A
DO | Winkler titration of surface water |  | 9.9 |  | 10.42 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 228 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-22 17:59 | 16.4 | 8.95 | 10.48 | 107 | 54.4 | 35.1
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-22 |  | 2009-05-22 |  | A
Time | Computer set to atomic time |  | 17:59 |  | 17:59 |  | A
Temp | VWR NIST traceable surface water |  | 16.5 |  | 16.4 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.48 |  | A
SpC | 148 µS/mL |  | 153.6 |  | 168 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-26 |  | 2009-05-26 |  | A
Time | Computer set to atomic time |  | 12:41 |  | 12:41 |  | A
Temp | VWR NIST traceable surface water |  | 21.5 |  | 21.3 |  | A
DO | Winkler titration of surface water |  | 9.3 |  | 8.79 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 153.9 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-31 11:00 | 15.9 | 7.56 | 9.76 | 98.8 | 41.6 | 27.3
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-31 |  | 2009-05-31 |  | A
Time | Computer set to atomic time |  | 15:15 |  | 15:15 |  | A
Temp | VWR NIST traceable surface water |  | 19.4 |  | 19.3 |  | A
DO | Winkler titration of surface water |  | 7.9 |  | 8.03 |  | A
SpC | 148 µS/mL |  | 153.6 |  | 177.1 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-05-19 | 12:01 | 17.26 | 34 | 52 | 10.79 | 112.3 | 8.5
2009-05-19 | 12:16 | 17.32 | 34 | 52 | 10.81 | 112.6 | 8.54
2009-05-19 | 12:31 | 17.36 | 34 | 52 | 10.7 | 111.6 | 8.6
2009-05-19 | 12:46 | 17.44 | 33 | 51 | 10.61 | 110.8 | 8.65
2009-05-19 | 13:00 | 17.52 | 33 | 51 | 10.61 | 110.9 | 8.7
2009-05-19 | 13:15 | 17.57 | 33 | 51 | 10.59 | 110.9 | 8.75
2009-05-19 | 13:30 | 17.63 | 33 | 51 | 10.57 | 110.8 | 8.78
2009-05-19 | 13:45 | 17.68 | 33 | 51 | 10.56 | 110.8 | 8.82
2009-05-19 | 14:00 | 17.71 | 33 | 51 | 10.56 | 110.8 | 8.85
2009-05-19 | 14:15 | 17.75 | 33 | 51 | 10.54 | 110.8 | 8.87
2009-05-19 | 14:30 | 17.78 | 33 | 51 | 10.48 | 110.2 | 8.9
2009-05-19 | 14:45 | 17.8 | 33 | 51 | 10.5 | 110.4 | 8.92
2009-05-19 | 15:00 | 17.82 | 33 | 51 | 10.45 | 110 | 8.94
2009-05-19 | 15:15 | 17.84 | 33 | 51 | 10.47 | 110.2 | 8.96
2009-05-19 | 15:30 | 17.85 | 33 | 51 | 10.41 | 109.6 | 8.97
2009-05-19 | 15:45 | 17.85 | 33 | 51 | 10.42 | 109.7 | 8.99
2009-05-19 | 16:00 | 17.84 | 33 | 51 | 10.42 | 109.7 | 9
2009-05-19 | 16:15 | 17.82 | 33 | 51 | 10.33 | 108.7 | 9
2009-05-19 | 16:30 | 17.79 | 33 | 51 | 10.28 | 108.1 | 9.01
2009-05-19 | 16:45 | 17.79 | 33 | 51 | 10.23 | 107.6 | 9.01
2009-05-19 | 17:00 | 17.78 | 33 | 52 | 10.19 | 107.2 | 9.01
2009-05-19 | 17:15 | 17.73 | 34 | 52 | 10.19 | 107.1 | 9.01
2009-05-19 | 17:30 | 17.68 | 34 | 52 | 10.12 | 106.2 | 9.01
2009-05-19 | 17:45 | 17.64 | 34 | 52 | 10.11 | 106 | 9
2009-05-19 | 18:00 | 17.54 | 34 | 52 | 10.04 | 105 | 8.99
2009-05-19 | 18:15 | 17.48 | 34 | 52 | 10.01 | 104.6 | 8.98
2009-05-19 | 18:30 | 17.4 | 34 | 52 | 9.94 | 103.8 | 8.96
2009-05-19 | 18:45 | 17.29 | 34 | 52 | 9.89 | 102.9 | 8.93
2009-05-19 | 19:00 | 17.19 | 34 | 52 | 9.82 | 102 | 8.91
2009-05-19 | 19:15 | 17.08 | 34 | 52 | 9.75 | 101.1 | 8.87
2009-05-19 | 19:30 | 16.98 | 34 | 52 | 9.66 | 99.9 | 8.82
2009-05-19 | 19:45 | 16.86 | 34 | 53 | 9.57 | 98.8 | 8.77
2009-05-19 | 20:00 | 16.77 | 34 | 53 | 9.47 | 97.6 | 8.71
2009-05-19 | 20:15 | 16.66 | 34 | 53 | 9.37 | 96.3 | 8.62
2009-05-19 | 20:30 | 16.55 | 35 | 53 | 9.28 | 95.1 | 8.52
2009-05-19 | 20:45 | 16.47 | 35 | 53 | 9.14 | 93.5 | 8.42
2009-05-19 | 21:00 | 16.34 | 35 | 54 | 9.08 | 92.7 | 8.29
2009-05-19 | 21:15 | 16.26 | 35 | 54 | 8.99 | 91.6 | 8.18
2009-05-19 | 21:30 | 16.21 | 35 | 54 | 8.88 | 90.3 | 8.09
2009-05-19 | 21:45 | 16.12 | 35 | 54 | 8.79 | 89.3 | 8.01
2009-05-19 | 22:00 | 16.05 | 35 | 54 | 8.71 | 88.4 | 7.94
2009-05-19 | 22:15 | 16 | 36 | 55 | 8.7 | 88.1 | 7.89
2009-05-19 | 22:30 | 15.94 | 36 | 55 | 8.62 | 87.2 | 7.84
2009-05-19 | 22:45 | 15.88 | 36 | 55 | 8.61 | 87 | 7.8
2009-05-19 | 23:00 | 15.83 | 36 | 55 | 8.6 | 86.8 | 7.77
2009-05-19 | 23:15 | 15.78 | 36 | 55 | 8.54 | 86.2 | 7.74
2009-05-19 | 23:30 | 15.73 | 35 | 55 | 8.55 | 86.2 | 7.72
2009-05-19 | 23:45 | 15.66 | 35 | 54 | 8.56 | 86.1 | 7.7
2009-05-20 | 0:00 | 15.6 | 35 | 54 | 8.54 | 85.8 | 7.68
2009-05-20 | 0:15 | 15.51 | 35 | 54 | 8.53 | 85.5 | 7.67
2009-05-20 | 0:30 | 15.45 | 35 | 54 | 8.53 | 85.4 | 7.65
2009-05-20 | 0:45 | 15.37 | 35 | 54 | 8.58 | 85.8 | 7.65
2009-05-20 | 1:01 | 15.29 | 35 | 53 | 8.57 | 85.5 | 7.64
2009-05-20 | 1:16 | 15.21 | 35 | 53 | 8.62 | 85.8 | 7.63
2009-05-20 | 1:31 | 15.12 | 34 | 53 | 8.62 | 85.7 | 7.62
2009-05-20 | 1:46 | 15.04 | 34 | 53 | 8.63 | 85.6 | 7.61
2009-05-20 | 2:01 | 14.95 | 34 | 52 | 8.68 | 86 | 7.61
2009-05-20 | 2:16 | 14.87 | 34 | 52 | 8.68 | 85.8 | 7.59
2009-05-20 | 2:31 | 14.79 | 34 | 52 | 8.69 | 85.8 | 7.59
2009-05-20 | 2:46 | 14.7 | 34 | 52 | 8.74 | 86.1 | 7.59
2009-05-20 | 3:01 | 14.63 | 33 | 51 | 8.75 | 86 | 7.58
2009-05-20 | 3:16 | 14.54 | 33 | 51 | 8.79 | 86.3 | 7.58
2009-05-20 | 3:31 | 14.44 | 33 | 51 | 8.84 | 86.6 | 7.57
2009-05-20 | 3:46 | 14.35 | 33 | 51 | 8.84 | 86.4 | 7.57
2009-05-20 | 4:01 | 14.25 | 33 | 50 | 8.85 | 86.3 | 7.56
2009-05-20 | 4:16 | 14.15 | 33 | 50 | 8.91 | 86.7 | 7.56
2009-05-20 | 4:31 | 14.05 | 33 | 50 | 8.92 | 86.7 | 7.56
2009-05-20 | 4:46 | 13.94 | 33 | 50 | 8.98 | 87 | 7.55
2009-05-20 | 5:01 | 13.84 | 32 | 50 | 9.02 | 87.2 | 7.56
2009-05-20 | 5:16 | 13.72 | 32 | 50 | 9.07 | 87.5 | 7.55
2009-05-20 | 5:31 | 13.62 | 32 | 50 | 9.12 | 87.8 | 7.55
2009-05-20 | 5:46 | 13.51 | 32 | 50 | 9.12 | 87.5 | 7.55
2009-05-20 | 6:01 | 13.4 | 32 | 49 | 9.19 | 88 | 7.55
2009-05-20 | 6:16 | 13.29 | 32 | 49 | 9.24 | 88.2 | 7.56
2009-05-20 | 6:31 | 13.21 | 32 | 49 | 9.34 | 89.1 | 7.57
2009-05-20 | 6:46 | 13.12 | 32 | 49 | 9.44 | 89.9 | 7.58
2009-05-20 | 7:01 | 13.05 | 32 | 49 | 9.5 | 90.3 | 7.59
2009-05-20 | 7:16 | 13 | 32 | 49 | 9.6 | 91.1 | 7.61
2009-05-20 | 7:31 | 12.96 | 32 | 49 | 9.7 | 92 | 7.63
2009-05-20 | 7:46 | 12.93 | 32 | 49 | 9.8 | 92.9 | 7.65
2009-05-20 | 8:01 | 12.92 | 32 | 49 | 9.92 | 94 | 7.67
2009-05-20 | 8:16 | 12.93 | 32 | 48 | 10.08 | 95.5 | 7.7
2009-05-20 | 8:31 | 12.96 | 31 | 48 | 10.17 | 96.4 | 7.73
2009-05-20 | 8:46 | 13 | 31 | 48 | 10.27 | 97.5 | 7.77
2009-05-20 | 9:01 | 13.06 | 31 | 48 | 10.39 | 98.7 | 7.81
2009-05-20 | 9:16 | 13.12 | 31 | 48 | 10.49 | 99.9 | 7.85
2009-05-20 | 9:31 | 13.2 | 31 | 48 | 10.61 | 101.1 | 7.9
2009-05-20 | 9:46 | 13.28 | 31 | 48 | 10.75 | 102.6 | 7.95
2009-05-20 | 10:01 | 13.37 | 31 | 48 | 10.88 | 104.2 | 8
2009-05-20 | 10:16 | 13.47 | 31 | 47 | 11 | 105.5 | 8.06
2009-05-20 | 10:31 | 13.58 | 31 | 47 | 11.08 | 106.5 | 8.11
2009-05-20 | 10:46 | 13.68 | 31 | 47 | 11.16 | 107.5 | 8.18
2009-05-20 | 11:01 | 13.78 | 31 | 47 | 11.23 | 108.5 | 8.24
2009-05-20 | 11:16 | 13.89 | 31 | 47 | 11.23 | 108.7 | 8.31
2009-05-20 | 11:31 | 14 | 30 | 47 | 11.26 | 109.2 | 8.38
2009-05-20 | 11:46 | 14.11 | 30 | 47 | 11.34 | 110.3 | 8.45
2009-05-20 | 12:01 | 14.23 | 30 | 47 | 11.38 | 111 | 8.51
2009-05-20 | 12:16 | 14.36 | 30 | 47 | 11.36 | 111.1 | 8.56
2009-05-20 | 12:31 | 14.5 | 30 | 47 | 11.34 | 111.2 | 8.62
2009-05-20 | 12:46 | 14.62 | 30 | 47 | 11.4 | 112.1 | 8.68
2009-05-20 | 13:00 | 14.76 | 30 | 47 | 11.42 | 112.6 | 8.72
2009-05-20 | 13:15 | 14.9 | 30 | 47 | 11.42 | 113 | 8.77
2009-05-20 | 13:30 | 15.04 | 30 | 47 | 11.36 | 112.8 | 8.81
2009-05-20 | 13:45 | 15.15 | 30 | 47 | 11.39 | 113.4 | 8.85
2009-05-20 | 14:00 | 15.28 | 30 | 47 | 11.37 | 113.5 | 8.87
2009-05-20 | 14:15 | 15.39 | 31 | 47 | 11.31 | 113.1 | 8.9
2009-05-20 | 14:30 | 15.5 | 31 | 47 | 11.27 | 112.9 | 8.93
2009-05-20 | 14:45 | 15.6 | 31 | 47 | 11.28 | 113.3 | 8.96
2009-05-20 | 15:00 | 15.69 | 31 | 47 | 11.28 | 113.5 | 8.98
2009-05-20 | 15:15 | 15.74 | 31 | 47 | 11.21 | 112.9 | 9
2009-05-20 | 15:30 | 15.82 | 31 | 47 | 11.21 | 113.1 | 9.01
2009-05-20 | 15:45 | 15.86 | 31 | 47 | 11.14 | 112.5 | 9.03
2009-05-20 | 16:00 | 15.91 | 31 | 47 | 11.14 | 112.6 | 9.04
2009-05-20 | 16:15 | 15.93 | 31 | 48 | 11.12 | 112.5 | 9.06
2009-05-20 | 16:30 | 15.95 | 31 | 47 | 11.09 | 112.2 | 9.06
2009-05-20 | 16:45 | 15.94 | 31 | 48 | 11.06 | 111.9 | 9.07
2009-05-20 | 17:00 | 15.93 | 31 | 48 | 11.02 | 111.5 | 9.07
2009-05-20 | 17:15 | 15.91 | 31 | 48 | 10.98 | 111 | 9.08
2009-05-20 | 17:30 | 15.87 | 31 | 48 | 10.94 | 110.5 | 9.08
2009-05-20 | 17:45 | 15.82 | 31 | 48 | 10.89 | 109.9 | 9.08
2009-05-20 | 18:00 | 15.77 | 31 | 48 | 10.8 | 108.8 | 9.07
2009-05-20 | 18:15 | 15.7 | 31 | 48 | 10.78 | 108.5 | 9.07
2009-05-20 | 18:30 | 15.62 | 31 | 48 | 10.68 | 107.3 | 9.05
2009-05-20 | 18:45 | 15.53 | 31 | 48 | 10.66 | 106.9 | 9.04
2009-05-20 | 19:00 | 15.44 | 31 | 48 | 10.6 | 106.1 | 9.02
2009-05-20 | 19:15 | 15.32 | 31 | 48 | 10.49 | 104.7 | 8.99
2009-05-20 | 19:30 | 15.23 | 31 | 48 | 10.38 | 103.5 | 8.96
2009-05-20 | 19:45 | 15.13 | 31 | 48 | 10.35 | 102.9 | 8.91
2009-05-20 | 20:00 | 15.02 | 31 | 48 | 10.22 | 101.4 | 8.87
2009-05-20 | 20:15 | 14.92 | 31 | 48 | 10.12 | 100.2 | 8.81
2009-05-20 | 20:30 | 14.82 | 32 | 49 | 10.07 | 99.5 | 8.73
2009-05-20 | 20:45 | 14.72 | 32 | 49 | 9.93 | 97.9 | 8.63
2009-05-20 | 21:00 | 14.62 | 32 | 49 | 9.83 | 96.7 | 8.53
2009-05-20 | 21:15 | 14.53 | 32 | 49 | 9.74 | 95.6 | 8.42
2009-05-20 | 21:30 | 14.44 | 32 | 49 | 9.66 | 94.6 | 8.29
2009-05-20 | 21:45 | 14.36 | 32 | 50 | 9.58 | 93.7 | 8.18
2009-05-20 | 22:00 | 14.29 | 32 | 50 | 9.51 | 92.9 | 8.09
2009-05-20 | 22:15 | 14.21 | 33 | 50 | 9.48 | 92.4 | 8.01
2009-05-20 | 22:30 | 14.16 | 33 | 50 | 9.4 | 91.6 | 7.94
2009-05-20 | 22:45 | 14.11 | 33 | 51 | 9.36 | 91 | 7.89
2009-05-20 | 23:00 | 14.07 | 33 | 51 | 9.32 | 90.6 | 7.84
2009-05-20 | 23:15 | 14.02 | 33 | 51 | 9.26 | 89.9 | 7.81
2009-05-20 | 23:30 | 13.98 | 33 | 51 | 9.2 | 89.2 | 7.77
2009-05-20 | 23:45 | 13.95 | 33 | 51 | 9.17 | 88.9 | 7.75
2009-05-21 | 0:00 | 13.9 | 33 | 51 | 9.19 | 88.9 | 7.72
2009-05-21 | 0:15 | 13.85 | 33 | 51 | 9.19 | 88.9 | 7.71
2009-05-21 | 0:30 | 13.8 | 33 | 51 | 9.19 | 88.8 | 7.69
2009-05-21 | 0:45 | 13.75 | 33 | 51 | 9.19 | 88.7 | 7.68
2009-05-21 | 1:01 | 13.71 | 33 | 51 | 9.19 | 88.6 | 7.66
2009-05-21 | 1:16 | 13.66 | 33 | 51 | 9.15 | 88.2 | 7.65
2009-05-21 | 1:31 | 13.62 | 33 | 51 | 9.17 | 88.3 | 7.64
2009-05-21 | 1:46 | 13.58 | 33 | 51 | 9.19 | 88.4 | 7.63
2009-05-21 | 2:01 | 13.53 | 33 | 51 | 9.16 | 88 | 7.63
2009-05-21 | 2:16 | 13.48 | 33 | 51 | 9.2 | 88.2 | 7.62
2009-05-21 | 2:31 | 13.43 | 33 | 51 | 9.22 | 88.3 | 7.61
2009-05-21 | 2:46 | 13.37 | 33 | 51 | 9.18 | 87.9 | 7.61
2009-05-21 | 3:01 | 13.31 | 33 | 51 | 9.19 | 87.9 | 7.6
2009-05-21 | 3:16 | 13.26 | 33 | 51 | 9.24 | 88.2 | 7.6
2009-05-21 | 3:31 | 13.19 | 33 | 51 | 9.23 | 88 | 7.6
2009-05-21 | 3:46 | 13.13 | 33 | 51 | 9.23 | 87.9 | 7.59
2009-05-21 | 4:01 | 13.07 | 33 | 51 | 9.25 | 88 | 7.59
2009-05-21 | 4:16 | 13 | 33 | 51 | 9.26 | 87.9 | 7.58
2009-05-21 | 4:31 | 12.94 | 33 | 51 | 9.27 | 87.9 | 7.59
2009-05-21 | 4:46 | 12.87 | 33 | 51 | 9.29 | 87.9 | 7.58
2009-05-21 | 5:01 | 12.8 | 33 | 51 | 9.35 | 88.4 | 7.58
2009-05-21 | 5:16 | 12.74 | 33 | 51 | 9.35 | 88.2 | 7.58
2009-05-21 | 5:31 | 12.67 | 33 | 51 | 9.4 | 88.5 | 7.58
2009-05-21 | 5:46 | 12.6 | 33 | 51 | 9.45 | 88.9 | 7.58
2009-05-21 | 6:01 | 12.53 | 33 | 51 | 9.5 | 89.2 | 7.58
2009-05-21 | 6:16 | 12.47 | 33 | 51 | 9.53 | 89.4 | 7.59
2009-05-21 | 6:31 | 12.42 | 33 | 51 | 9.62 | 90.1 | 7.6
2009-05-21 | 6:46 | 12.38 | 33 | 51 | 9.67 | 90.5 | 7.61
2009-05-21 | 7:01 | 12.35 | 33 | 51 | 9.8 | 91.6 | 7.63
2009-05-21 | 7:16 | 12.33 | 33 | 52 | 9.91 | 92.6 | 7.64
2009-05-21 | 7:31 | 12.33 | 33 | 51 | 10.01 | 93.6 | 7.67
2009-05-21 | 7:46 | 12.34 | 33 | 51 | 10.08 | 94.2 | 7.69
2009-05-21 | 8:01 | 12.36 | 34 | 52 | 10.17 | 95.1 | 7.71
2009-05-21 | 8:16 | 12.41 | 34 | 52 | 10.27 | 96.2 | 7.74
2009-05-21 | 8:31 | 12.47 | 33 | 52 | 10.37 | 97.3 | 7.78
2009-05-21 | 8:46 | 12.54 | 34 | 52 | 10.47 | 98.4 | 7.82
2009-05-21 | 9:01 | 12.64 | 33 | 51 | 10.62 | 100 | 7.85
2009-05-21 | 9:16 | 12.74 | 33 | 52 | 10.75 | 101.4 | 7.9
2009-05-21 | 9:31 | 12.86 | 33 | 51 | 10.8 | 102.2 | 7.94
2009-05-21 | 9:46 | 12.98 | 33 | 51 | 10.95 | 103.9 | 7.99
2009-05-21 | 10:01 | 13.11 | 33 | 51 | 11.07 | 105.3 | 8.06
2009-05-21 | 10:16 | 13.26 | 33 | 51 | 11.16 | 106.6 | 8.12
2009-05-21 | 10:31 | 13.4 | 33 | 51 | 11.25 | 107.7 | 8.18
2009-05-21 | 10:46 | 13.55 | 33 | 51 | 11.32 | 108.7 | 8.25
2009-05-21 | 11:01 | 13.71 | 33 | 51 | 11.32 | 109.1 | 8.32
2009-05-21 | 11:16 | 13.86 | 33 | 51 | 11.4 | 110.3 | 8.4
2009-05-21 | 11:31 | 14.02 | 33 | 51 | 11.45 | 111.1 | 8.47
2009-05-21 | 11:46 | 14.17 | 33 | 50 | 11.48 | 111.8 | 8.53
2009-05-21 | 12:01 | 14.32 | 33 | 50 | 11.44 | 111.8 | 8.6
2009-05-21 | 12:16 | 14.49 | 33 | 50 | 11.49 | 112.7 | 8.66
2009-05-21 | 12:31 | 14.64 | 33 | 50 | 11.46 | 112.8 | 8.71
2009-05-21 | 12:46 | 14.8 | 33 | 50 | 11.43 | 112.9 | 8.77
2009-05-21 | 13:00 | 14.95 | 33 | 50 | 11.42 | 113.1 | 8.82
2009-05-21 | 13:15 | 15.11 | 32 | 50 | 11.39 | 113.2 | 8.86
2009-05-21 | 13:30 | 15.26 | 32 | 50 | 11.38 | 113.4 | 8.9
2009-05-21 | 13:45 | 15.39 | 33 | 50 | 11.41 | 114.1 | 8.94
2009-05-21 | 14:00 | 15.53 | 33 | 50 | 11.39 | 114.3 | 8.97
2009-05-21 | 14:15 | 15.65 | 33 | 50 | 11.33 | 114 | 9
2009-05-21 | 14:30 | 15.76 | 33 | 50 | 11.31 | 114 | 9.02
2009-05-21 | 14:45 | 15.86 | 33 | 50 | 11.26 | 113.7 | 9.04
2009-05-21 | 15:00 | 15.93 | 33 | 50 | 11.27 | 114 | 9.06
2009-05-21 | 15:15 | 16.02 | 33 | 50 | 11.27 | 114.2 | 9.08
2009-05-21 | 15:30 | 16.08 | 33 | 50 | 11.18 | 113.5 | 9.1
2009-05-21 | 15:45 | 16.12 | 33 | 51 | 11.16 | 113.3 | 9.11
2009-05-21 | 16:00 | 16.15 | 33 | 51 | 11.1 | 112.8 | 9.12
2009-05-21 | 16:15 | 16.19 | 33 | 51 | 11.13 | 113.2 | 9.13
2009-05-21 | 16:30 | 16.2 | 33 | 51 | 11.06 | 112.6 | 9.14
2009-05-21 | 16:45 | 16.19 | 33 | 51 | 11.01 | 112 | 9.15
2009-05-21 | 17:00 | 16.18 | 33 | 51 | 11.02 | 112 | 9.15
2009-05-21 | 17:15 | 16.15 | 33 | 51 | 10.95 | 111.3 | 9.15
2009-05-21 | 17:30 | 16.08 | 33 | 51 | 10.88 | 110.5 | 9.15
2009-05-21 | 17:45 | 16.03 | 33 | 52 | 10.88 | 110.3 | 9.15
2009-05-21 | 18:00 | 15.97 | 34 | 52 | 10.85 | 109.9 | 9.15
2009-05-21 | 18:15 | 15.9 | 34 | 52 | 10.82 | 109.3 | 9.14
2009-05-21 | 18:30 | 15.75 | 34 | 52 | 10.72 | 108 | 9.12
2009-05-21 | 18:45 | 15.66 | 34 | 52 | 10.63 | 106.9 | 9.11
2009-05-21 | 19:00 | 15.58 | 34 | 52 | 10.6 | 106.4 | 9.1
2009-05-21 | 19:15 | 15.47 | 34 | 52 | 10.5 | 105.2 | 9.07
2009-05-21 | 19:30 | 15.35 | 34 | 52 | 10.4 | 103.9 | 9.04
2009-05-21 | 19:45 | 15.24 | 34 | 52 | 10.35 | 103.2 | 9.01
2009-05-21 | 20:00 | 15.14 | 34 | 53 | 10.28 | 102.3 | 8.97
2009-05-21 | 20:15 | 15.04 | 34 | 53 | 10.18 | 101.1 | 8.91
2009-05-21 | 20:30 | 14.95 | 34 | 53 | 10.05 | 99.6 | 8.85
2009-05-21 | 20:45 | 14.85 | 34 | 53 | 9.98 | 98.6 | 8.77
2009-05-21 | 21:00 | 14.78 | 35 | 53 | 9.87 | 97.5 | 8.67
2009-05-21 | 21:15 | 14.7 | 35 | 54 | 9.78 | 96.3 | 8.57
2009-05-21 | 21:30 | 14.64 | 35 | 54 | 9.7 | 95.4 | 8.45
2009-05-21 | 21:45 | 14.58 | 35 | 54 | 9.63 | 94.7 | 8.34
2009-05-21 | 22:00 | 14.53 | 35 | 54 | 9.49 | 93.2 | 8.23
2009-05-21 | 22:15 | 14.49 | 35 | 54 | 9.4 | 92.2 | 8.13
2009-05-21 | 22:30 | 14.45 | 35 | 55 | 9.34 | 91.5 | 8.05
2009-05-21 | 22:45 | 14.41 | 36 | 55 | 9.3 | 91.1 | 7.98
2009-05-21 | 23:00 | 14.38 | 36 | 55 | 9.22 | 90.2 | 7.92
2009-05-21 | 23:15 | 14.34 | 36 | 55 | 9.16 | 89.6 | 7.87
2009-05-21 | 23:30 | 14.31 | 36 | 55 | 9.11 | 89 | 7.83
2009-05-21 | 23:45 | 14.28 | 36 | 55 | 9.08 | 88.7 | 7.8
2009-05-22 | 0:00 | 14.23 | 36 | 55 | 9.07 | 88.5 | 7.78
2009-05-22 | 0:15 | 14.2 | 36 | 55 | 9.04 | 88.1 | 7.75
2009-05-22 | 0:30 | 14.17 | 36 | 55 | 9.05 | 88.1 | 7.73
2009-05-22 | 0:45 | 14.14 | 36 | 55 | 9.05 | 88.1 | 7.71
2009-05-22 | 1:01 | 14.11 | 36 | 55 | 9.05 | 88.1 | 7.7
2009-05-22 | 1:16 | 14.09 | 36 | 55 | 9.05 | 88 | 7.69
2009-05-22 | 1:31 | 14.07 | 36 | 55 | 9.02 | 87.7 | 7.68
2009-05-22 | 1:46 | 14.05 | 36 | 55 | 9.03 | 87.7 | 7.67
2009-05-22 | 2:01 | 14.03 | 36 | 55 | 9.03 | 87.7 | 7.66
2009-05-22 | 2:16 | 14.01 | 36 | 55 | 9 | 87.4 | 7.65
2009-05-22 | 2:31 | 13.99 | 36 | 55 | 9.02 | 87.5 | 7.64
2009-05-22 | 2:46 | 13.97 | 36 | 55 | 9 | 87.2 | 7.64
2009-05-22 | 3:01 | 13.94 | 36 | 55 | 8.98 | 87 | 7.63
2009-05-22 | 3:16 | 13.92 | 36 | 55 | 8.98 | 87 | 7.63
2009-05-22 | 3:31 | 13.89 | 36 | 55 | 8.98 | 87 | 7.62
2009-05-22 | 3:46 | 13.85 | 36 | 55 | 9.01 | 87.2 | 7.62
2009-05-22 | 4:01 | 13.82 | 36 | 55 | 9.05 | 87.5 | 7.61
2009-05-22 | 4:16 | 13.79 | 36 | 55 | 9.03 | 87.2 | 7.62
2009-05-22 | 4:31 | 13.75 | 36 | 55 | 9.06 | 87.5 | 7.61
2009-05-22 | 4:46 | 13.7 | 36 | 55 | 9.05 | 87.2 | 7.61
2009-05-22 | 5:01 | 13.66 | 36 | 55 | 9.09 | 87.5 | 7.61
2009-05-22 | 5:16 | 13.61 | 36 | 55 | 9.12 | 87.7 | 7.6
2009-05-22 | 5:31 | 13.54 | 36 | 55 | 9.1 | 87.4 | 7.6
2009-05-22 | 5:46 | 13.5 | 36 | 55 | 9.15 | 87.9 | 7.6
2009-05-22 | 6:01 | 13.44 | 36 | 55 | 9.22 | 88.4 | 7.6
2009-05-22 | 6:16 | 13.39 | 36 | 55 | 9.27 | 88.7 | 7.61
2009-05-22 | 6:31 | 13.35 | 36 | 55 | 9.3 | 89 | 7.62
2009-05-22 | 6:46 | 13.32 | 36 | 55 | 9.38 | 89.6 | 7.63
2009-05-22 | 7:01 | 13.3 | 36 | 55 | 9.5 | 90.7 | 7.65
2009-05-22 | 7:16 | 13.29 | 36 | 55 | 9.61 | 91.8 | 7.67
2009-05-22 | 7:31 | 13.29 | 36 | 55 | 9.68 | 92.5 | 7.69
2009-05-22 | 7:46 | 13.3 | 36 | 55 | 9.77 | 93.4 | 7.71
2009-05-22 | 8:01 | 13.33 | 36 | 55 | 9.88 | 94.4 | 7.73
2009-05-22 | 8:16 | 13.38 | 36 | 55 | 9.97 | 95.4 | 7.76
2009-05-22 | 8:31 | 13.44 | 36 | 55 | 10.08 | 96.6 | 7.8
2009-05-22 | 8:46 | 13.52 | 36 | 55 | 10.19 | 97.9 | 7.83
2009-05-22 | 9:01 | 13.61 | 36 | 55 | 10.34 | 99.5 | 7.88
2009-05-22 | 9:16 | 13.71 | 36 | 55 | 10.46 | 100.9 | 7.91
2009-05-22 | 9:31 | 13.82 | 36 | 55 | 10.58 | 102.3 | 7.96
2009-05-22 | 9:46 | 13.94 | 36 | 55 | 10.71 | 103.8 | 8.01
2009-05-22 | 10:01 | 14.07 | 35 | 55 | 10.81 | 105.1 | 8.08
2009-05-22 | 10:16 | 14.21 | 35 | 54 | 10.91 | 106.3 | 8.14
2009-05-22 | 10:31 | 14.35 | 35 | 54 | 10.99 | 107.5 | 8.2
2009-05-22 | 10:46 | 14.51 | 35 | 54 | 11.02 | 108.2 | 8.28
2009-05-22 | 11:01 | 14.67 | 35 | 54 | 11.11 | 109.4 | 8.35
2009-05-22 | 11:16 | 14.82 | 35 | 54 | 11.12 | 109.8 | 8.43
2009-05-22 | 11:31 | 14.97 | 35 | 54 | 11.2 | 111 | 8.51
2009-05-22 | 11:46 | 15.12 | 35 | 53 | 11.24 | 111.8 | 8.57
2009-05-22 | 12:01 | 15.28 | 35 | 53 | 11.26 | 112.4 | 8.64
2009-05-22 | 12:16 | 15.44 | 35 | 53 | 11.29 | 113 | 8.7
2009-05-22 | 12:31 | 15.6 | 34 | 53 | 11.3 | 113.5 | 8.76
2009-05-22 | 12:46 | 15.73 | 34 | 53 | 11.25 | 113.4 | 8.81
2009-05-22 | 13:00 | 15.89 | 34 | 53 | 11.28 | 114 | 8.85
2009-05-22 | 13:15 | 16.03 | 34 | 53 | 11.28 | 114.4 | 8.9
2009-05-22 | 13:30 | 16.18 | 34 | 53 | 11.22 | 114.1 | 8.93
2009-05-22 | 13:45 | 16.28 | 34 | 53 | 11.2 | 114.1 | 8.97
2009-05-22 | 14:00 | 16.36 | 34 | 53 | 11.22 | 114.5 | 8.99
2009-05-22 | 14:15 | 16.32 | 34 | 53 | 11.16 | 113.9 | 9.02
2009-05-22 | 14:30 | 16.33 | 34 | 53 | 11.18 | 114.1 | 9.04
2009-05-22 | 14:45 | 16.35 | 34 | 53 | 11.18 | 114.1 | 9.06
2009-05-22 | 15:00 | 16.34 | 34 | 53 | 11.17 | 114 | 9.08
2009-05-22 | 15:15 | 16.34 | 34 | 53 | 11.16 | 113.9 | 9.09
2009-05-22 | 15:30 | 16.4 | 34 | 53 | 11.12 | 113.6 | 9.11
2009-05-22 | 15:45 | 16.48 | 34 | 53 | 11.14 | 114 | 9.12
2009-05-22 | 16:00 | 16.52 | 35 | 53 | 11.09 | 113.6 | 9.14
2009-05-22 | 16:15 | 16.54 | 35 | 53 | 11.03 | 113.1 | 9.14
2009-05-22 | 16:30 | 16.55 | 35 | 53 | 11.04 | 113.2 | 9.14
2009-05-22 | 16:45 | 16.54 | 35 | 53 | 10.96 | 112.3 | 9.14
2009-05-22 | 17:00 | 16.54 | 35 | 53 | 10.95 | 112.2 | 9.14
2009-05-22 | 17:15 | 16.51 | 35 | 54 | 10.91 | 111.7 | 9.14
2009-05-22 | 17:30 | 16.48 | 35 | 54 | 10.82 | 110.8 | 9.14
2009-05-22 | 17:45 | 16.45 | 35 | 54 | 10.75 | 109.9 | 9.14
2009-05-22 | 18:00 | 16.38 | 35 | 54 | 10.73 | 109.6 | 9.13
2009-05-22 | 18:15 | 16.34 | 35 | 54 | 10.64 | 108.5 | 9.13
2009-05-22 | 18:30 | 16.3 | 35 | 54 | 10.56 | 107.7 | 9.12
2009-05-22 | 18:45 | 16.24 | 35 | 54 | 10.49 | 106.8 | 9.11
2009-05-22 | 19:00 | 16.15 | 35 | 54 | 10.45 | 106.2 | 9.09
2009-05-22 | 19:15 | 16.06 | 35 | 54 | 10.34 | 104.9 | 9.06
2009-05-22 | 19:30 | 15.96 | 35 | 54 | 10.28 | 104.1 | 9.03
2009-05-22 | 19:45 | 15.86 | 35 | 54 | 10.21 | 103.2 | 8.99
2009-05-22 | 20:00 | 15.77 | 35 | 54 | 10.06 | 101.4 | 8.94
2009-05-22 | 20:15 | 15.69 | 35 | 55 | 9.98 | 100.4 | 8.87
2009-05-22 | 20:30 | 15.62 | 36 | 55 | 9.84 | 98.9 | 8.81
2009-05-22 | 20:45 | 15.56 | 36 | 55 | 9.76 | 98 | 8.73
2009-05-22 | 21:00 | 15.51 | 36 | 55 | 9.67 | 97 | 8.62
2009-05-22 | 21:15 | 15.47 | 36 | 55 | 9.52 | 95.3 | 8.51
2009-05-22 | 21:30 | 15.42 | 36 | 56 | 9.42 | 94.2 | 8.39
2009-05-22 | 21:45 | 15.36 | 36 | 56 | 9.3 | 93 | 8.27
2009-05-22 | 22:00 | 15.32 | 37 | 56 | 9.27 | 92.6 | 8.16
2009-05-22 | 22:15 | 15.27 | 37 | 56 | 9.16 | 91.4 | 8.07
2009-05-22 | 22:30 | 15.23 | 37 | 57 | 9.11 | 90.8 | 8
2009-05-22 | 22:45 | 15.18 | 37 | 57 | 9.06 | 90.2 | 7.94
2009-05-22 | 23:00 | 15.14 | 37 | 57 | 8.99 | 89.4 | 7.89
2009-05-22 | 23:15 | 15.11 | 37 | 57 | 8.93 | 88.7 | 7.84
2009-05-22 | 23:30 | 15.07 | 37 | 57 | 8.92 | 88.6 | 7.81
2009-05-22 | 23:45 | 15.04 | 37 | 57 | 8.91 | 88.4 | 7.78
2009-05-23 | 0:00 | 15.01 | 37 | 57 | 8.84 | 87.7 | 7.75
2009-05-23 | 0:15 | 14.99 | 37 | 57 | 8.81 | 87.4 | 7.73
2009-05-23 | 0:30 | 14.96 | 37 | 57 | 8.79 | 87.1 | 7.72
2009-05-23 | 0:45 | 14.93 | 37 | 57 | 8.82 | 87.4 | 7.7
2009-05-23 | 1:01 | 14.91 | 37 | 57 | 8.79 | 87 | 7.69
2009-05-23 | 1:16 | 14.88 | 37 | 57 | 8.82 | 87.2 | 7.68
2009-05-23 | 1:31 | 14.85 | 37 | 57 | 8.83 | 87.3 | 7.67
2009-05-23 | 1:46 | 14.81 | 37 | 57 | 8.84 | 87.3 | 7.66
2009-05-23 | 2:01 | 14.77 | 37 | 57 | 8.82 | 87.1 | 7.65
2009-05-23 | 2:16 | 14.73 | 37 | 57 | 8.84 | 87.1 | 7.65
2009-05-23 | 2:31 | 14.69 | 37 | 57 | 8.87 | 87.3 | 7.64
2009-05-23 | 2:46 | 14.65 | 37 | 56 | 8.88 | 87.4 | 7.63
2009-05-23 | 3:01 | 14.6 | 37 | 56 | 8.85 | 87.1 | 7.63
2009-05-23 | 3:16 | 14.56 | 37 | 56 | 8.89 | 87.3 | 7.62
2009-05-23 | 3:31 | 14.51 | 37 | 56 | 8.87 | 87.1 | 7.62
2009-05-23 | 3:46 | 14.46 | 37 | 56 | 8.91 | 87.4 | 7.62
2009-05-23 | 4:01 | 14.41 | 37 | 56 | 8.9 | 87.1 | 7.61
2009-05-23 | 4:16 | 14.36 | 37 | 56 | 8.94 | 87.4 | 7.61
2009-05-23 | 4:31 | 14.3 | 36 | 56 | 8.94 | 87.3 | 7.61
2009-05-23 | 4:46 | 14.26 | 36 | 56 | 8.97 | 87.5 | 7.61
2009-05-23 | 5:01 | 14.19 | 36 | 56 | 8.96 | 87.3 | 7.61
2009-05-23 | 5:16 | 14.13 | 36 | 56 | 9 | 87.6 | 7.6
2009-05-23 | 5:31 | 14.08 | 36 | 56 | 9.03 | 87.8 | 7.6
2009-05-23 | 5:46 | 14.01 | 36 | 56 | 9.06 | 88 | 7.6
2009-05-23 | 6:01 | 13.95 | 36 | 56 | 9.11 | 88.3 | 7.6
2009-05-23 | 6:16 | 13.89 | 36 | 56 | 9.13 | 88.4 | 7.61
2009-05-23 | 6:31 | 13.85 | 36 | 56 | 9.22 | 89.2 | 7.61
2009-05-23 | 6:46 | 13.8 | 36 | 55 | 9.27 | 89.6 | 7.62
2009-05-23 | 7:01 | 13.77 | 36 | 55 | 9.34 | 90.2 | 7.64
2009-05-23 | 7:16 | 13.75 | 36 | 55 | 9.43 | 91 | 7.65
2009-05-23 | 7:31 | 13.74 | 36 | 55 | 9.53 | 92 | 7.67
2009-05-23 | 7:46 | 13.75 | 36 | 55 | 9.63 | 92.9 | 7.69
2009-05-23 | 8:01 | 13.78 | 36 | 55 | 9.77 | 94.4 | 7.71
2009-05-23 | 8:16 | 13.83 | 35 | 55 | 9.89 | 95.6 | 7.74
2009-05-23 | 8:31 | 13.89 | 35 | 55 | 10.02 | 97 | 7.77
2009-05-23 | 8:46 | 13.97 | 35 | 54 | 10.09 | 97.8 | 7.8
2009-05-23 | 9:01 | 14.06 | 35 | 54 | 10.18 | 98.9 | 7.83
2009-05-23 | 9:16 | 14.17 | 35 | 54 | 10.28 | 100.1 | 7.88
2009-05-23 | 9:31 | 14.28 | 35 | 54 | 10.42 | 101.7 | 7.92
2009-05-23 | 9:46 | 14.4 | 35 | 53 | 10.49 | 102.7 | 7.97
2009-05-23 | 10:01 | 14.53 | 35 | 53 | 10.57 | 103.8 | 8.02
2009-05-23 | 10:16 | 14.66 | 34 | 53 | 10.67 | 105 | 8.09
2009-05-23 | 10:31 | 14.79 | 34 | 53 | 10.75 | 106.1 | 8.15
2009-05-23 | 10:46 | 14.92 | 34 | 53 | 10.83 | 107.2 | 8.22
2009-05-23 | 11:01 | 15.05 | 34 | 53 | 10.88 | 108.1 | 8.28
2009-05-23 | 11:16 | 15.19 | 34 | 52 | 10.98 | 109.3 | 8.35
2009-05-23 | 11:31 | 15.32 | 34 | 52 | 11.04 | 110.2 | 8.42
2009-05-23 | 11:46 | 15.46 | 34 | 52 | 11.08 | 111 | 8.49
2009-05-23 | 12:01 | 15.6 | 34 | 52 | 11.06 | 111.1 | 8.55
2009-05-23 | 12:16 | 15.74 | 34 | 52 | 11.06 | 111.5 | 8.61
2009-05-23 | 12:31 | 15.89 | 34 | 52 | 11.06 | 111.8 | 8.67
2009-05-23 | 12:46 | 16.04 | 34 | 52 | 11.12 | 112.7 | 8.72
2009-05-23 | 13:00 | 16.18 | 34 | 52 | 11.14 | 113.3 | 8.77
2009-05-23 | 13:15 | 16.33 | 34 | 52 | 11.09 | 113.1 | 8.81
2009-05-23 | 13:30 | 16.45 | 34 | 52 | 11.12 | 113.7 | 8.85
2009-05-23 | 13:45 | 16.59 | 34 | 52 | 11.06 | 113.5 | 8.88
2009-05-23 | 14:00 | 16.72 | 34 | 52 | 11.07 | 113.9 | 8.91
2009-05-23 | 14:15 | 16.84 | 34 | 52 | 11.02 | 113.6 | 8.94
2009-05-23 | 14:30 | 16.94 | 34 | 52 | 11.03 | 114 | 8.96
2009-05-23 | 14:45 | 17.05 | 34 | 52 | 10.97 | 113.6 | 8.98
2009-05-23 | 15:00 | 17.14 | 34 | 53 | 10.97 | 113.8 | 9
2009-05-23 | 15:15 | 17.23 | 34 | 53 | 10.9 | 113.3 | 9.02
2009-05-23 | 15:30 | 17.31 | 34 | 53 | 10.85 | 113 | 9.03
2009-05-23 | 15:45 | 17.38 | 34 | 53 | 10.85 | 113.1 | 9.04
2009-05-23 | 16:00 | 17.43 | 34 | 53 | 10.77 | 112.5 | 9.05
2009-05-23 | 16:15 | 17.47 | 35 | 53 | 10.77 | 112.5 | 9.06
2009-05-23 | 16:30 | 17.51 | 35 | 53 | 10.68 | 111.7 | 9.06
2009-05-23 | 16:45 | 17.53 | 35 | 53 | 10.67 | 111.6 | 9.07
2009-05-23 | 17:00 | 17.54 | 35 | 53 | 10.59 | 110.8 | 9.07
2009-05-23 | 17:15 | 17.55 | 35 | 54 | 10.57 | 110.6 | 9.07
2009-05-23 | 17:30 | 17.51 | 35 | 54 | 10.51 | 109.9 | 9.07
2009-05-23 | 17:45 | 17.55 | 35 | 54 | 10.5 | 109.8 | 9.07
2009-05-23 | 18:00 | 17.55 | 35 | 54 | 10.45 | 109.4 | 9.07
2009-05-23 | 18:15 | 17.47 | 35 | 54 | 10.34 | 108 | 9.05
2009-05-23 | 18:30 | 17.44 | 35 | 54 | 10.29 | 107.5 | 9.04
2009-05-23 | 18:45 | 17.43 | 36 | 55 | 10.21 | 106.6 | 9.04
2009-05-23 | 19:00 | 17.36 | 36 | 55 | 10.06 | 104.8 | 9
2009-05-23 | 19:15 | 17.27 | 36 | 55 | 9.92 | 103.3 | 8.96
2009-05-23 | 19:30 | 17.22 | 36 | 55 | 9.85 | 102.3 | 8.92
2009-05-23 | 19:45 | 17.17 | 36 | 55 | 9.71 | 100.8 | 8.87
2009-05-23 | 20:00 | 17.11 | 36 | 56 | 9.59 | 99.5 | 8.81
2009-05-23 | 20:15 | 17.04 | 36 | 56 | 9.53 | 98.7 | 8.75
2009-05-23 | 20:30 | 16.97 | 36 | 56 | 9.44 | 97.6 | 8.68
2009-05-23 | 20:45 | 16.92 | 37 | 56 | 9.34 | 96.4 | 8.59
2009-05-23 | 21:00 | 16.86 | 37 | 57 | 9.23 | 95.3 | 8.49
2009-05-23 | 21:15 | 16.8 | 37 | 57 | 9.09 | 93.7 | 8.36
2009-05-23 | 21:30 | 16.74 | 37 | 57 | 8.97 | 92.3 | 8.24
2009-05-23 | 21:45 | 16.69 | 37 | 57 | 8.9 | 91.5 | 8.14
2009-05-23 | 22:00 | 16.65 | 38 | 58 | 8.78 | 90.2 | 8.05
2009-05-23 | 22:15 | 16.62 | 38 | 58 | 8.74 | 89.7 | 7.97
2009-05-23 | 22:30 | 16.6 | 38 | 58 | 8.66 | 88.9 | 7.91
2009-05-23 | 22:45 | 16.56 | 38 | 58 | 8.58 | 88 | 7.87
2009-05-23 | 23:00 | 16.53 | 38 | 58 | 8.53 | 87.4 | 7.83
2009-05-23 | 23:15 | 16.49 | 38 | 58 | 8.53 | 87.3 | 7.79
2009-05-23 | 23:30 | 16.46 | 38 | 58 | 8.51 | 87 | 7.77
2009-05-23 | 23:45 | 16.41 | 38 | 58 | 8.46 | 86.4 | 7.74
2009-05-24 | 0:00 | 16.36 | 38 | 58 | 8.43 | 86.1 | 7.72
2009-05-24 | 0:15 | 16.31 | 38 | 58 | 8.41 | 85.8 | 7.7
2009-05-24 | 0:30 | 16.26 | 38 | 58 | 8.41 | 85.6 | 7.69
2009-05-24 | 0:45 | 16.21 | 38 | 58 | 8.43 | 85.8 | 7.67
2009-05-24 | 1:01 | 16.16 | 38 | 58 | 8.41 | 85.5 | 7.66
2009-05-24 | 1:16 | 16.1 | 38 | 58 | 8.42 | 85.5 | 7.66
2009-05-24 | 1:31 | 16.04 | 38 | 58 | 8.4 | 85.2 | 7.64
2009-05-24 | 1:46 | 15.98 | 38 | 58 | 8.41 | 85.2 | 7.63
2009-05-24 | 2:01 | 15.92 | 38 | 58 | 8.42 | 85.2 | 7.63
2009-05-24 | 2:16 | 15.86 | 38 | 58 | 8.46 | 85.4 | 7.62
2009-05-24 | 2:31 | 15.8 | 38 | 58 | 8.44 | 85.1 | 7.61
2009-05-24 | 2:46 | 15.74 | 38 | 58 | 8.48 | 85.5 | 7.61
2009-05-24 | 3:01 | 15.68 | 38 | 58 | 8.5 | 85.6 | 7.61
2009-05-24 | 3:16 | 15.63 | 38 | 58 | 8.52 | 85.7 | 7.6
2009-05-24 | 3:31 | 15.58 | 38 | 58 | 8.51 | 85.4 | 7.6
2009-05-24 | 3:46 | 15.51 | 37 | 58 | 8.51 | 85.3 | 7.59
2009-05-24 | 4:01 | 15.46 | 37 | 58 | 8.55 | 85.6 | 7.59
2009-05-24 | 4:16 | 15.4 | 37 | 57 | 8.54 | 85.5 | 7.59
2009-05-24 | 4:31 | 15.34 | 37 | 57 | 8.55 | 85.4 | 7.58
2009-05-24 | 4:46 | 15.28 | 37 | 57 | 8.57 | 85.5 | 7.58
2009-05-24 | 5:01 | 15.21 | 37 | 57 | 8.58 | 85.5 | 7.58
2009-05-24 | 5:16 | 15.15 | 37 | 57 | 8.6 | 85.6 | 7.58
2009-05-24 | 5:31 | 15.08 | 37 | 57 | 8.65 | 86 | 7.58
2009-05-24 | 5:46 | 15.02 | 37 | 57 | 8.67 | 86 | 7.58
2009-05-24 | 6:01 | 14.95 | 37 | 57 | 8.73 | 86.5 | 7.58
2009-05-24 | 6:16 | 14.89 | 37 | 57 | 8.77 | 86.8 | 7.58
2009-05-24 | 6:31 | 14.84 | 37 | 56 | 8.83 | 87.3 | 7.59
2009-05-24 | 6:46 | 14.81 | 37 | 56 | 8.91 | 88 | 7.6
2009-05-24 | 7:01 | 14.78 | 37 | 56 | 9.03 | 89.1 | 7.61
2009-05-24 | 7:16 | 14.76 | 36 | 56 | 9.11 | 89.8 | 7.63
2009-05-24 | 7:31 | 14.75 | 36 | 56 | 9.24 | 91.1 | 7.65
2009-05-24 | 7:46 | 14.77 | 36 | 56 | 9.36 | 92.3 | 7.67
2009-05-24 | 8:01 | 14.79 | 36 | 56 | 9.47 | 93.5 | 7.69
2009-05-24 | 8:16 | 14.83 | 36 | 56 | 9.59 | 94.8 | 7.72
2009-05-24 | 8:31 | 14.89 | 36 | 55 | 9.67 | 95.7 | 7.75
2009-05-24 | 8:46 | 14.96 | 36 | 55 | 9.81 | 97.2 | 7.78
2009-05-24 | 9:01 | 15.05 | 36 | 55 | 9.93 | 98.6 | 7.82
2009-05-24 | 9:16 | 15.15 | 36 | 55 | 10.06 | 100.1 | 7.86
2009-05-24 | 9:31 | 15.26 | 36 | 55 | 10.16 | 101.4 | 7.91
2009-05-24 | 9:46 | 15.38 | 35 | 55 | 10.23 | 102.3 | 7.96
2009-05-24 | 10:01 | 15.51 | 35 | 54 | 10.32 | 103.5 | 8.01
2009-05-24 | 10:16 | 15.64 | 35 | 54 | 10.46 | 105.1 | 8.07
2009-05-24 | 10:31 | 15.76 | 35 | 54 | 10.56 | 106.5 | 8.13
2009-05-24 | 10:46 | 15.9 | 35 | 54 | 10.59 | 107.1 | 8.2
2009-05-24 | 11:01 | 16.04 | 35 | 53 | 10.64 | 107.9 | 8.27
2009-05-24 | 11:16 | 16.16 | 35 | 53 | 10.74 | 109.2 | 8.34
2009-05-24 | 11:31 | 16.3 | 34 | 53 | 10.8 | 110.1 | 8.41
2009-05-24 | 11:46 | 16.44 | 34 | 53 | 10.79 | 110.4 | 8.47
2009-05-24 | 12:01 | 16.57 | 34 | 53 | 10.86 | 111.3 | 8.54
2009-05-24 | 12:16 | 16.71 | 34 | 52 | 10.84 | 111.5 | 8.6
2009-05-24 | 12:31 | 16.84 | 34 | 52 | 10.84 | 111.8 | 8.66
2009-05-24 | 12:46 | 16.97 | 34 | 52 | 10.89 | 112.6 | 8.71
2009-05-24 | 13:00 | 17.11 | 34 | 52 | 10.92 | 113.3 | 8.76
2009-05-24 | 13:15 | 17.24 | 34 | 52 | 10.92 | 113.6 | 8.81
2009-05-24 | 13:30 | 17.37 | 34 | 52 | 10.92 | 113.8 | 8.84
2009-05-24 | 13:45 | 17.48 | 34 | 52 | 10.91 | 114.1 | 8.88
2009-05-24 | 14:00 | 17.59 | 34 | 52 | 10.89 | 114.1 | 8.9
2009-05-24 | 14:15 | 17.69 | 34 | 52 | 10.88 | 114.2 | 8.93
2009-05-24 | 14:30 | 17.78 | 34 | 52 | 10.86 | 114.2 | 8.96
2009-05-24 | 14:45 | 17.86 | 34 | 52 | 10.84 | 114.2 | 8.98
2009-05-24 | 15:00 | 17.93 | 34 | 53 | 10.77 | 113.6 | 9
2009-05-24 | 15:15 | 17.99 | 34 | 53 | 10.73 | 113.3 | 9.02
2009-05-24 | 15:30 | 18.05 | 34 | 53 | 10.73 | 113.5 | 9.03
2009-05-24 | 15:45 | 18.09 | 34 | 53 | 10.72 | 113.4 | 9.05
2009-05-24 | 16:00 | 18.12 | 34 | 53 | 10.69 | 113.2 | 9.06
2009-05-24 | 16:15 | 18.13 | 34 | 53 | 10.59 | 112.2 | 9.06
2009-05-24 | 16:30 | 18.14 | 35 | 53 | 10.55 | 111.7 | 9.07
2009-05-24 | 16:45 | 18.14 | 35 | 53 | 10.54 | 111.6 | 9.07
2009-05-24 | 17:00 | 18.13 | 35 | 53 | 10.5 | 111.2 | 9.08
2009-05-24 | 17:15 | 18.1 | 35 | 53 | 10.41 | 110.2 | 9.07
2009-05-24 | 17:30 | 18.07 | 35 | 53 | 10.39 | 109.9 | 9.07
2009-05-24 | 17:45 | 18.03 | 35 | 53 | 10.3 | 108.9 | 9.06
2009-05-24 | 18:00 | 17.99 | 35 | 54 | 10.23 | 108 | 9.06
2009-05-24 | 18:15 | 17.95 | 35 | 54 | 10.15 | 107.1 | 9.05
2009-05-24 | 18:30 | 17.9 | 35 | 54 | 10.07 | 106.2 | 9.04
2009-05-24 | 18:45 | 17.82 | 35 | 54 | 10 | 105.3 | 9.01
2009-05-24 | 19:00 | 17.74 | 35 | 54 | 9.97 | 104.7 | 9
2009-05-24 | 19:15 | 17.66 | 35 | 54 | 9.88 | 103.7 | 8.96
2009-05-24 | 19:30 | 17.57 | 35 | 54 | 9.79 | 102.5 | 8.92
2009-05-24 | 19:45 | 17.49 | 35 | 54 | 9.64 | 100.8 | 8.87
2009-05-24 | 20:00 | 17.4 | 35 | 54 | 9.58 | 99.9 | 8.82
2009-05-24 | 20:15 | 17.32 | 35 | 54 | 9.44 | 98.3 | 8.75
2009-05-24 | 20:30 | 17.25 | 35 | 55 | 9.32 | 97 | 8.67
2009-05-24 | 20:45 | 17.17 | 35 | 55 | 9.25 | 96 | 8.57
2009-05-24 | 21:00 | 17.09 | 36 | 55 | 9.16 | 95 | 8.46
2009-05-24 | 21:15 | 17.02 | 36 | 55 | 9.07 | 93.8 | 8.33
2009-05-24 | 21:30 | 16.95 | 36 | 55 | 8.94 | 92.4 | 8.21
2009-05-24 | 21:45 | 16.89 | 36 | 55 | 8.87 | 91.6 | 8.11
2009-05-24 | 22:00 | 16.83 | 36 | 56 | 8.77 | 90.4 | 8.02
2009-05-24 | 22:15 | 16.78 | 36 | 56 | 8.69 | 89.5 | 7.95
2009-05-24 | 22:30 | 16.72 | 36 | 56 | 8.65 | 88.9 | 7.88
2009-05-24 | 22:45 | 16.67 | 37 | 56 | 8.61 | 88.4 | 7.84
2009-05-24 | 23:00 | 16.62 | 37 | 56 | 8.52 | 87.5 | 7.79
2009-05-24 | 23:15 | 16.57 | 37 | 56 | 8.52 | 87.3 | 7.76
2009-05-24 | 23:30 | 16.52 | 37 | 56 | 8.47 | 86.8 | 7.73
2009-05-24 | 23:45 | 16.47 | 37 | 57 | 8.43 | 86.2 | 7.7
2009-05-25 | 0:00 | 16.42 | 37 | 56 | 8.45 | 86.4 | 7.69
2009-05-25 | 0:15 | 16.37 | 37 | 56 | 8.42 | 85.9 | 7.67
2009-05-25 | 0:30 | 16.33 | 37 | 56 | 8.4 | 85.7 | 7.65
2009-05-25 | 0:45 | 16.27 | 37 | 56 | 8.39 | 85.5 | 7.64
2009-05-25 | 1:01 | 16.23 | 37 | 56 | 8.39 | 85.4 | 7.63
2009-05-25 | 1:16 | 16.18 | 37 | 56 | 8.42 | 85.7 | 7.62
2009-05-25 | 1:31 | 16.13 | 37 | 56 | 8.41 | 85.4 | 7.61
2009-05-25 | 1:46 | 16.08 | 37 | 56 | 8.44 | 85.7 | 7.6
2009-05-25 | 2:01 | 16.03 | 36 | 56 | 8.43 | 85.5 | 7.6
2009-05-25 | 2:16 | 15.98 | 36 | 56 | 8.44 | 85.4 | 7.6
2009-05-25 | 2:31 | 15.92 | 36 | 56 | 8.47 | 85.7 | 7.59
2009-05-25 | 2:46 | 15.86 | 36 | 56 | 8.47 | 85.5 | 7.58
2009-05-25 | 3:01 | 15.8 | 36 | 56 | 8.47 | 85.5 | 7.58
2009-05-25 | 3:16 | 15.74 | 36 | 56 | 8.48 | 85.4 | 7.57
2009-05-25 | 3:31 | 15.67 | 36 | 56 | 8.53 | 85.8 | 7.57
2009-05-25 | 3:46 | 15.6 | 36 | 55 | 8.57 | 86.1 | 7.57
2009-05-25 | 4:01 | 15.54 | 36 | 55 | 8.59 | 86.2 | 7.57
2009-05-25 | 4:16 | 15.46 | 36 | 55 | 8.58 | 85.9 | 7.56
2009-05-25 | 4:31 | 15.39 | 36 | 55 | 8.62 | 86.2 | 7.56
2009-05-25 | 4:46 | 15.31 | 36 | 55 | 8.62 | 86 | 7.55
2009-05-25 | 5:01 | 15.23 | 36 | 55 | 8.62 | 86 | 7.56
2009-05-25 | 5:16 | 15.14 | 35 | 55 | 8.68 | 86.4 | 7.55
2009-05-25 | 5:31 | 15.06 | 35 | 54 | 8.69 | 86.3 | 7.55
2009-05-25 | 5:46 | 14.97 | 35 | 54 | 8.71 | 86.4 | 7.55
2009-05-25 | 6:01 | 14.88 | 35 | 54 | 8.76 | 86.7 | 7.55
2009-05-25 | 6:16 | 14.8 | 35 | 54 | 8.83 | 87.2 | 7.56
2009-05-25 | 6:31 | 14.73 | 35 | 54 | 8.9 | 87.8 | 7.56
2009-05-25 | 6:46 | 14.67 | 35 | 54 | 8.99 | 88.5 | 7.57
2009-05-25 | 7:01 | 14.62 | 35 | 54 | 9.09 | 89.4 | 7.58
2009-05-25 | 7:16 | 14.58 | 35 | 54 | 9.23 | 90.7 | 7.61
2009-05-25 | 7:31 | 14.56 | 35 | 53 | 9.36 | 92 | 7.63
2009-05-25 | 7:46 | 14.56 | 35 | 53 | 9.44 | 92.7 | 7.65
2009-05-25 | 8:01 | 14.57 | 35 | 53 | 9.55 | 93.8 | 7.67
2009-05-25 | 8:16 | 14.6 | 35 | 53 | 9.65 | 94.8 | 7.7
2009-05-25 | 8:31 | 14.64 | 34 | 53 | 9.75 | 95.9 | 7.73
2009-05-25 | 8:46 | 14.7 | 34 | 53 | 9.92 | 97.8 | 7.76
2009-05-25 | 9:01 | 14.79 | 34 | 53 | 10.04 | 99.1 | 7.8
2009-05-25 | 9:16 | 14.88 | 34 | 52 | 10.16 | 100.5 | 7.84
2009-05-25 | 9:31 | 14.98 | 34 | 52 | 10.27 | 101.8 | 7.88
2009-05-25 | 9:46 | 15.1 | 34 | 52 | 10.39 | 103.2 | 7.93
2009-05-25 | 10:01 | 15.21 | 34 | 52 | 10.44 | 104 | 7.98
2009-05-25 | 10:16 | 15.34 | 34 | 52 | 10.53 | 105.2 | 8.04
2009-05-25 | 10:31 | 15.46 | 34 | 52 | 10.59 | 106.1 | 8.1
2009-05-25 | 10:46 | 15.59 | 33 | 51 | 10.72 | 107.6 | 8.16
2009-05-25 | 11:01 | 15.72 | 33 | 51 | 10.75 | 108.3 | 8.23
2009-05-25 | 11:16 | 15.85 | 33 | 51 | 10.79 | 109 | 8.3
2009-05-25 | 11:31 | 15.97 | 33 | 51 | 10.88 | 110.2 | 8.36
2009-05-25 | 11:46 | 16.11 | 33 | 51 | 10.89 | 110.5 | 8.43
2009-05-25 | 12:01 | 16.23 | 33 | 50 | 10.91 | 111.1 | 8.5
2009-05-25 | 12:16 | 16.37 | 33 | 50 | 10.97 | 112 | 8.57
2009-05-25 | 12:31 | 16.5 | 33 | 50 | 10.99 | 112.6 | 8.62
2009-05-25 | 12:46 | 16.63 | 33 | 50 | 11.02 | 113.1 | 8.69
2009-05-25 | 13:00 | 16.77 | 32 | 50 | 10.97 | 113 | 8.72
2009-05-25 | 13:15 | 16.9 | 32 | 50 | 10.95 | 113 | 8.78
2009-05-25 | 13:30 | 17.02 | 32 | 50 | 10.99 | 113.7 | 8.82
2009-05-25 | 13:45 | 17.14 | 32 | 50 | 10.99 | 114 | 8.86
2009-05-25 | 14:00 | 17.26 | 32 | 50 | 10.99 | 114.3 | 8.89
2009-05-25 | 14:15 | 17.35 | 32 | 50 | 10.96 | 114.3 | 8.91
2009-05-25 | 14:30 | 17.44 | 33 | 50 | 10.95 | 114.4 | 8.94
2009-05-25 | 14:45 | 17.53 | 33 | 50 | 10.89 | 113.9 | 8.97
2009-05-25 | 15:00 | 17.61 | 33 | 50 | 10.84 | 113.6 | 8.98
2009-05-25 | 15:15 | 17.67 | 33 | 50 | 10.8 | 113.3 | 9
2009-05-25 | 15:30 | 17.72 | 33 | 50 | 10.77 | 113.1 | 9.01
2009-05-25 | 15:45 | 17.76 | 33 | 50 | 10.73 | 112.8 | 9.03
2009-05-25 | 16:00 | 17.8 | 33 | 51 | 10.69 | 112.5 | 9.04
2009-05-25 | 16:15 | 17.83 | 33 | 51 | 10.7 | 112.6 | 9.05
2009-05-25 | 16:30 | 17.85 | 33 | 51 | 10.68 | 112.4 | 9.06
2009-05-25 | 16:45 | 17.86 | 33 | 51 | 10.63 | 112 | 9.06
2009-05-25 | 17:00 | 17.86 | 33 | 51 | 10.55 | 111.1 | 9.06
2009-05-25 | 17:15 | 17.84 | 33 | 51 | 10.49 | 110.4 | 9.06
2009-05-25 | 17:30 | 17.82 | 33 | 51 | 10.47 | 110.2 | 9.07
2009-05-25 | 17:45 | 17.79 | 33 | 51 | 10.38 | 109.2 | 9.06
2009-05-25 | 18:00 | 17.74 | 33 | 51 | 10.37 | 108.9 | 9.05
2009-05-25 | 18:15 | 17.68 | 33 | 51 | 10.28 | 107.9 | 9.05
2009-05-25 | 18:30 | 17.61 | 33 | 51 | 10.23 | 107.2 | 9.03
2009-05-25 | 18:45 | 17.53 | 34 | 52 | 10.19 | 106.6 | 9.02
2009-05-25 | 19:00 | 17.45 | 34 | 52 | 10.11 | 105.6 | 8.99
2009-05-25 | 19:15 | 17.36 | 34 | 52 | 9.98 | 104.1 | 8.97
2009-05-25 | 19:30 | 17.27 | 34 | 52 | 9.88 | 102.8 | 8.93
2009-05-25 | 19:45 | 17.17 | 34 | 52 | 9.78 | 101.5 | 8.89
2009-05-25 | 20:00 | 17.07 | 34 | 52 | 9.72 | 100.8 | 8.84
2009-05-25 | 20:15 | 16.96 | 34 | 52 | 9.63 | 99.6 | 8.78
2009-05-25 | 20:30 | 16.88 | 34 | 53 | 9.49 | 98 | 8.71
2009-05-25 | 20:45 | 16.8 | 34 | 53 | 9.42 | 97 | 8.61
2009-05-25 | 21:00 | 16.71 | 34 | 53 | 9.33 | 95.9 | 8.51
2009-05-25 | 21:15 | 16.65 | 35 | 53 | 9.24 | 94.9 | 8.39
2009-05-25 | 21:30 | 16.57 | 35 | 54 | 9.12 | 93.5 | 8.27
2009-05-25 | 21:45 | 16.5 | 35 | 54 | 9.02 | 92.4 | 8.17
2009-05-25 | 22:00 | 16.43 | 35 | 54 | 8.94 | 91.4 | 8.06
2009-05-25 | 22:15 | 16.37 | 35 | 55 | 8.86 | 90.4 | 7.98
2009-05-25 | 22:30 | 16.31 | 36 | 55 | 8.79 | 89.6 | 7.91
2009-05-25 | 22:45 | 16.26 | 36 | 55 | 8.74 | 89 | 7.86
2009-05-25 | 23:00 | 16.21 | 36 | 55 | 8.71 | 88.7 | 7.81
2009-05-25 | 23:15 | 16.16 | 36 | 55 | 8.65 | 88 | 7.78
2009-05-25 | 23:30 | 16.11 | 36 | 55 | 8.61 | 87.4 | 7.74
2009-05-25 | 23:45 | 16.07 | 36 | 55 | 8.61 | 87.4 | 7.71
2009-05-26 | 0:00 | 16.02 | 36 | 55 | 8.61 | 87.2 | 7.69
2009-05-26 | 0:15 | 15.98 | 36 | 55 | 8.57 | 86.8 | 7.67
2009-05-26 | 0:30 | 15.94 | 36 | 55 | 8.54 | 86.4 | 7.65
2009-05-26 | 0:45 | 15.89 | 36 | 55 | 8.53 | 86.2 | 7.65
2009-05-26 | 1:01 | 15.85 | 36 | 55 | 8.56 | 86.4 | 7.63
2009-05-26 | 1:16 | 15.81 | 36 | 55 | 8.54 | 86.1 | 7.62
2009-05-26 | 1:31 | 15.77 | 36 | 55 | 8.53 | 86 | 7.6
2009-05-26 | 1:46 | 15.73 | 36 | 55 | 8.53 | 85.9 | 7.6
2009-05-26 | 2:01 | 15.7 | 36 | 55 | 8.53 | 85.8 | 7.59
2009-05-26 | 2:16 | 15.66 | 35 | 55 | 8.54 | 85.9 | 7.58
2009-05-26 | 2:31 | 15.63 | 35 | 55 | 8.58 | 86.2 | 7.58
2009-05-26 | 2:46 | 15.58 | 35 | 54 | 8.56 | 86 | 7.57
2009-05-26 | 3:01 | 15.53 | 35 | 54 | 8.6 | 86.3 | 7.57
2009-05-26 | 3:16 | 15.49 | 35 | 54 | 8.59 | 86.1 | 7.56
2009-05-26 | 3:31 | 15.44 | 35 | 54 | 8.63 | 86.4 | 7.55
2009-05-26 | 3:46 | 15.38 | 35 | 54 | 8.62 | 86.2 | 7.56
2009-05-26 | 4:01 | 15.34 | 35 | 54 | 8.62 | 86.1 | 7.55
2009-05-26 | 4:16 | 15.28 | 35 | 54 | 8.64 | 86.2 | 7.55
2009-05-26 | 4:31 | 15.22 | 35 | 54 | 8.69 | 86.5 | 7.54
2009-05-26 | 4:46 | 15.16 | 35 | 54 | 8.7 | 86.5 | 7.55
2009-05-26 | 5:01 | 15.09 | 35 | 53 | 8.69 | 86.4 | 7.54
2009-05-26 | 5:16 | 15.03 | 35 | 53 | 8.71 | 86.5 | 7.54
2009-05-26 | 5:31 | 14.96 | 35 | 53 | 8.77 | 86.9 | 7.54
2009-05-26 | 5:46 | 14.88 | 34 | 53 | 8.83 | 87.3 | 7.54
2009-05-26 | 6:01 | 14.82 | 34 | 53 | 8.84 | 87.3 | 7.54
2009-05-26 | 6:16 | 14.75 | 34 | 53 | 8.88 | 87.6 | 7.54
2009-05-26 | 6:31 | 14.69 | 34 | 53 | 8.99 | 88.6 | 7.55
2009-05-26 | 6:46 | 14.65 | 34 | 52 | 9.09 | 89.5 | 7.56
2009-05-26 | 7:01 | 14.62 | 34 | 52 | 9.2 | 90.5 | 7.58
2009-05-26 | 7:16 | 14.6 | 34 | 52 | 9.28 | 91.2 | 7.59
2009-05-26 | 7:31 | 14.59 | 34 | 52 | 9.4 | 92.4 | 7.61
2009-05-26 | 7:46 | 14.59 | 34 | 52 | 9.51 | 93.5 | 7.63
2009-05-26 | 8:01 | 14.61 | 34 | 52 | 9.6 | 94.4 | 7.66
2009-05-26 | 8:16 | 14.64 | 34 | 52 | 9.73 | 95.7 | 7.68
2009-05-26 | 8:31 | 14.69 | 33 | 51 | 9.85 | 97.1 | 7.71
2009-05-26 | 8:46 | 14.76 | 33 | 51 | 9.97 | 98.4 | 7.75
2009-05-26 | 9:01 | 14.83 | 33 | 51 | 10.08 | 99.6 | 7.78
2009-05-26 | 9:16 | 14.92 | 33 | 51 | 10.17 | 100.7 | 7.82
2009-05-26 | 9:31 | 15.01 | 33 | 51 | 10.26 | 101.8 | 7.86
2009-05-26 | 9:46 | 15.12 | 33 | 50 | 10.4 | 103.4 | 7.91
2009-05-26 | 10:01 | 15.23 | 33 | 50 | 10.52 | 104.9 | 7.96
2009-05-26 | 10:16 | 15.36 | 32 | 50 | 10.62 | 106.1 | 8.01
2009-05-26 | 10:31 | 15.49 | 32 | 50 | 10.69 | 107.2 | 8.07
2009-05-26 | 10:46 | 15.63 | 32 | 50 | 10.79 | 108.4 | 8.13
2009-05-26 | 11:01 | 15.77 | 32 | 49 | 10.85 | 109.4 | 8.2
2009-05-26 | 11:16 | 15.91 | 32 | 49 | 10.92 | 110.5 | 8.26
2009-05-26 | 11:31 | 16.05 | 32 | 49 | 10.98 | 111.4 | 8.34
2009-05-26 | 11:46 | 16.19 | 32 | 49 | 10.97 | 111.6 | 8.4
2009-05-26 | 12:01 | 16.32 | 31 | 48 | 11.01 | 112.4 | 8.47
2009-05-26 | 12:16 | 16.46 | 31 | 48 | 11.04 | 112.9 | 8.54
2009-05-26 | 12:31 | 16.57 | 31 | 48 | 11.02 | 113 | 8.59
2009-05-26 | 12:46 | 16.69 | 31 | 48 | 11.05 | 113.6 | 8.65
2009-05-26 | 13:00 | 16.83 | 31 | 48 | 11.02 | 113.6 | 8.7
2009-05-26 | 13:15 | 16.95 | 31 | 48 | 11.05 | 114.2 | 8.75
2009-05-26 | 13:30 | 17.07 | 31 | 48 | 11.05 | 114.5 | 8.79
2009-05-26 | 13:45 | 17.19 | 31 | 48 | 11.05 | 114.8 | 8.83
2009-05-26 | 14:00 | 17.28 | 31 | 48 | 11.04 | 114.9 | 8.86
2009-05-26 | 14:15 | 17.37 | 31 | 48 | 11.04 | 115.1 | 8.9
2009-05-26 | 14:30 | 17.46 | 31 | 48 | 10.97 | 114.6 | 8.92
2009-05-26 | 14:45 | 17.52 | 31 | 48 | 10.93 | 114.3 | 8.95
2009-05-26 | 15:00 | 17.57 | 31 | 48 | 10.9 | 114.1 | 8.97
2009-05-26 | 15:15 | 17.64 | 31 | 48 | 10.91 | 114.4 | 8.99
2009-05-26 | 15:30 | 17.69 | 31 | 48 | 10.91 | 114.5 | 9
2009-05-26 | 15:45 | 17.71 | 31 | 48 | 10.87 | 114.2 | 9.02
2009-05-26 | 16:00 | 17.73 | 31 | 48 | 10.81 | 113.6 | 9.04
2009-05-26 | 16:15 | 17.75 | 31 | 48 | 10.76 | 113.1 | 9.04
2009-05-26 | 16:30 | 17.73 | 32 | 48 | 10.76 | 113 | 9.05
2009-05-26 | 16:45 | 17.74 | 32 | 49 | 10.73 | 112.8 | 9.06
2009-05-26 | 17:00 | 17.72 | 32 | 49 | 10.7 | 112.4 | 9.06
2009-05-26 | 17:15 | 17.68 | 32 | 49 | 10.62 | 111.4 | 9.07
2009-05-26 | 17:30 | 17.65 | 32 | 49 | 10.6 | 111.1 | 9.06
2009-05-26 | 17:45 | 17.6 | 32 | 49 | 10.56 | 110.6 | 9.05
2009-05-26 | 18:00 | 17.55 | 32 | 49 | 10.51 | 110 | 9.05
2009-05-26 | 18:15 | 17.5 | 32 | 49 | 10.46 | 109.3 | 9.04
2009-05-26 | 18:30 | 17.43 | 32 | 49 | 10.39 | 108.5 | 9.03
2009-05-26 | 18:45 | 17.35 | 32 | 49 | 10.33 | 107.7 | 9.01
2009-05-26 | 19:00 | 17.28 | 32 | 49 | 10.24 | 106.6 | 8.99
2009-05-26 | 19:15 | 17.18 | 32 | 49 | 10.15 | 105.4 | 8.96
2009-05-26 | 19:30 | 17.1 | 32 | 50 | 10.06 | 104.3 | 8.92
2009-05-26 | 19:45 | 17.02 | 32 | 49 | 9.97 | 103.2 | 8.88
2009-05-26 | 20:00 | 16.94 | 32 | 50 | 9.83 | 101.6 | 8.82
2009-05-26 | 20:15 | 16.86 | 32 | 50 | 9.72 | 100.3 | 8.77
2009-05-26 | 20:30 | 16.78 | 33 | 50 | 9.6 | 98.8 | 8.69
2009-05-26 | 20:45 | 16.71 | 33 | 50 | 9.48 | 97.5 | 8.6
2009-05-26 | 21:00 | 16.64 | 33 | 51 | 9.43 | 96.8 | 8.49
2009-05-26 | 21:15 | 16.58 | 33 | 51 | 9.35 | 95.9 | 8.38
2009-05-26 | 21:30 | 16.52 | 33 | 51 | 9.26 | 94.8 | 8.26
2009-05-26 | 21:45 | 16.46 | 34 | 52 | 9.2 | 94.1 | 8.16
2009-05-26 | 22:00 | 16.41 | 34 | 52 | 9.1 | 93 | 8.05
2009-05-26 | 22:15 | 16.36 | 34 | 52 | 8.99 | 91.8 | 7.97
2009-05-26 | 22:30 | 16.32 | 34 | 53 | 8.94 | 91.2 | 7.9
2009-05-26 | 22:45 | 16.29 | 34 | 53 | 8.89 | 90.6 | 7.84
2009-05-26 | 23:00 | 16.25 | 35 | 53 | 8.79 | 89.5 | 7.79
2009-05-26 | 23:15 | 16.22 | 35 | 53 | 8.77 | 89.3 | 7.76
2009-05-26 | 23:30 | 16.19 | 35 | 53 | 8.7 | 88.5 | 7.73
2009-05-26 | 23:45 | 16.16 | 35 | 54 | 8.65 | 87.9 | 7.69
2009-05-27 | 0:00 | 16.13 | 35 | 54 | 8.65 | 87.9 | 7.67
2009-05-27 | 0:15 | 16.1 | 35 | 54 | 8.6 | 87.4 | 7.65
2009-05-27 | 0:30 | 16.08 | 35 | 54 | 8.58 | 87.1 | 7.63
2009-05-27 | 0:45 | 16.06 | 35 | 54 | 8.56 | 86.8 | 7.62
2009-05-27 | 1:01 | 16.03 | 35 | 54 | 8.58 | 87 | 7.6
2009-05-27 | 1:16 | 16.01 | 35 | 54 | 8.55 | 86.7 | 7.59
2009-05-27 | 1:31 | 15.98 | 35 | 54 | 8.57 | 86.8 | 7.58
2009-05-27 | 1:46 | 15.97 | 35 | 54 | 8.55 | 86.5 | 7.57
2009-05-27 | 2:01 | 15.95 | 35 | 54 | 8.56 | 86.7 | 7.56
2009-05-27 | 2:16 | 15.93 | 35 | 54 | 8.54 | 86.4 | 7.56
2009-05-27 | 2:31 | 15.9 | 35 | 54 | 8.54 | 86.3 | 7.55
2009-05-27 | 2:46 | 15.88 | 35 | 54 | 8.53 | 86.2 | 7.55
2009-05-27 | 3:01 | 15.85 | 35 | 54 | 8.54 | 86.3 | 7.55
2009-05-27 | 3:16 | 15.82 | 35 | 54 | 8.54 | 86.2 | 7.54
2009-05-27 | 3:31 | 15.79 | 35 | 54 | 8.55 | 86.3 | 7.54
2009-05-27 | 3:46 | 15.75 | 35 | 54 | 8.56 | 86.2 | 7.53
2009-05-27 | 4:01 | 15.71 | 35 | 54 | 8.6 | 86.6 | 7.53
2009-05-27 | 4:16 | 15.67 | 35 | 54 | 8.63 | 86.8 | 7.53
2009-05-27 | 4:31 | 15.62 | 35 | 53 | 8.61 | 86.6 | 7.52
2009-05-27 | 4:46 | 15.57 | 35 | 53 | 8.62 | 86.5 | 7.52
2009-05-27 | 5:01 | 15.51 | 35 | 53 | 8.66 | 86.9 | 7.52
2009-05-27 | 5:16 | 15.45 | 35 | 53 | 8.67 | 86.8 | 7.52
2009-05-27 | 5:31 | 15.39 | 35 | 53 | 8.7 | 87 | 7.52
2009-05-27 | 5:46 | 15.33 | 35 | 53 | 8.72 | 87.1 | 7.52
2009-05-27 | 6:01 | 15.28 | 35 | 53 | 8.78 | 87.6 | 7.52
2009-05-27 | 6:16 | 15.22 | 35 | 53 | 8.81 | 87.8 | 7.53
2009-05-27 | 6:31 | 15.17 | 35 | 53 | 8.91 | 88.7 | 7.53
2009-05-27 | 6:46 | 15.13 | 34 | 53 | 9.01 | 89.6 | 7.54
2009-05-27 | 7:01 | 15.11 | 34 | 53 | 9.07 | 90.1 | 7.56
2009-05-27 | 7:16 | 15.09 | 34 | 53 | 9.17 | 91.1 | 7.58
2009-05-27 | 7:31 | 15.09 | 34 | 53 | 9.3 | 92.4 | 7.6
2009-05-27 | 7:46 | 15.11 | 34 | 53 | 9.37 | 93.2 | 7.62
2009-05-27 | 8:01 | 15.14 | 34 | 53 | 9.47 | 94.2 | 7.64
2009-05-27 | 8:16 | 15.19 | 34 | 52 | 9.58 | 95.4 | 7.66
2009-05-27 | 8:31 | 15.26 | 34 | 52 | 9.69 | 96.6 | 7.69
2009-05-27 | 8:46 | 15.33 | 34 | 52 | 9.83 | 98.2 | 7.72
2009-05-27 | 9:01 | 15.42 | 34 | 52 | 9.95 | 99.6 | 7.76
2009-05-27 | 9:16 | 15.52 | 34 | 52 | 10.02 | 100.5 | 7.79
2009-05-27 | 9:31 | 15.63 | 34 | 52 | 10.15 | 102.1 | 7.83
2009-05-27 | 9:46 | 15.75 | 33 | 51 | 10.27 | 103.5 | 7.88
2009-05-27 | 10:01 | 15.87 | 33 | 51 | 10.33 | 104.4 | 7.92
2009-05-27 | 10:16 | 16.01 | 33 | 51 | 10.41 | 105.5 | 7.98
2009-05-27 | 10:31 | 16.14 | 33 | 51 | 10.52 | 106.9 | 8.03
2009-05-27 | 10:46 | 16.28 | 33 | 50 | 10.58 | 107.8 | 8.09
2009-05-27 | 11:01 | 16.4 | 33 | 50 | 10.67 | 109 | 8.15
2009-05-27 | 11:16 | 16.53 | 32 | 50 | 10.74 | 110 | 8.21
2009-05-27 | 11:31 | 16.66 | 32 | 50 | 10.79 | 110.8 | 8.28
2009-05-27 | 11:46 | 16.79 | 32 | 49 | 10.82 | 111.5 | 8.33
2009-05-27 | 12:01 | 16.91 | 32 | 49 | 10.85 | 112.1 | 8.4
2009-05-27 | 12:16 | 17.04 | 32 | 49 | 10.88 | 112.7 | 8.45
2009-05-27 | 12:31 | 17.17 | 31 | 48 | 10.89 | 113.1 | 8.51
2009-05-27 | 12:46 | 17.29 | 31 | 48 | 10.9 | 113.5 | 8.57
2009-05-27 | 13:00 | 17.41 | 31 | 48 | 10.85 | 113.2 | 8.62
2009-05-27 | 13:15 | 17.54 | 31 | 48 | 10.87 | 113.8 | 8.67
2009-05-27 | 13:30 | 17.67 | 31 | 48 | 10.88 | 114.2 | 8.71
2009-05-27 | 13:45 | 17.78 | 31 | 48 | 10.83 | 113.9 | 8.75
2009-05-27 | 14:00 | 17.88 | 31 | 48 | 10.8 | 113.8 | 8.78
2009-05-27 | 14:15 | 17.96 | 31 | 48 | 10.83 | 114.3 | 8.81
2009-05-27 | 14:30 | 18.04 | 31 | 48 | 10.83 | 114.5 | 8.84
2009-05-27 | 14:45 | 18.12 | 31 | 48 | 10.82 | 114.5 | 8.87
2009-05-27 | 15:00 | 18.17 | 31 | 48 | 10.8 | 114.5 | 8.88
2009-05-27 | 15:15 | 18.22 | 31 | 48 | 10.78 | 114.4 | 8.9
2009-05-27 | 15:30 | 18.27 | 31 | 48 | 10.75 | 114.2 | 8.92
2009-05-27 | 15:45 | 18.3 | 31 | 48 | 10.72 | 114 | 8.94
2009-05-27 | 16:00 | 18.32 | 31 | 48 | 10.69 | 113.7 | 8.95
2009-05-27 | 16:15 | 18.35 | 31 | 48 | 10.66 | 113.4 | 8.96
2009-05-27 | 16:30 | 18.36 | 31 | 48 | 10.61 | 112.9 | 8.97
2009-05-27 | 16:45 | 18.34 | 32 | 49 | 10.56 | 112.3 | 8.97
2009-05-27 | 17:00 | 18.33 | 32 | 49 | 10.47 | 111.3 | 8.98
2009-05-27 | 17:15 | 18.3 | 32 | 49 | 10.44 | 111 | 8.97
2009-05-27 | 17:30 | 18.21 | 32 | 49 | 10.37 | 110.1 | 8.96
2009-05-27 | 17:45 | 18.16 | 32 | 49 | 10.28 | 108.9 | 8.95
2009-05-27 | 18:00 | 18.15 | 32 | 49 | 10.27 | 108.8 | 8.95
2009-05-27 | 18:15 | 18.1 | 32 | 49 | 10.18 | 107.8 | 8.94
2009-05-27 | 18:30 | 18.06 | 32 | 49 | 10.16 | 107.5 | 8.93
2009-05-27 | 18:45 | 18.01 | 32 | 49 | 10.12 | 106.9 | 8.91
2009-05-27 | 19:00 | 17.96 | 32 | 49 | 10.01 | 105.6 | 8.88
2009-05-27 | 19:15 | 17.89 | 32 | 50 | 9.93 | 104.7 | 8.86
2009-05-27 | 19:30 | 17.81 | 32 | 50 | 9.85 | 103.7 | 8.81
2009-05-27 | 19:45 | 17.75 | 33 | 50 | 9.77 | 102.6 | 8.75
2009-05-27 | 20:00 | 17.67 | 33 | 50 | 9.66 | 101.3 | 8.68
2009-05-27 | 20:15 | 17.6 | 33 | 51 | 9.5 | 99.6 | 8.6
2009-05-27 | 20:30 | 17.52 | 33 | 51 | 9.39 | 98.2 | 8.51
2009-05-27 | 20:45 | 17.45 | 33 | 51 | 9.26 | 96.7 | 8.4
2009-05-27 | 21:00 | 17.4 | 33 | 52 | 9.15 | 95.4 | 8.29
2009-05-27 | 21:15 | 17.35 | 34 | 52 | 9.09 | 94.7 | 8.19
2009-05-27 | 21:30 | 17.3 | 34 | 52 | 9 | 93.7 | 8.09
2009-05-27 | 21:45 | 17.25 | 34 | 52 | 8.89 | 92.5 | 8
2009-05-27 | 22:00 | 17.21 | 34 | 53 | 8.81 | 91.5 | 7.93
2009-05-27 | 22:15 | 17.17 | 34 | 53 | 8.72 | 90.6 | 7.87
2009-05-27 | 22:30 | 17.13 | 35 | 53 | 8.68 | 90.1 | 7.82
2009-05-27 | 22:45 | 17.09 | 35 | 53 | 8.6 | 89.1 | 7.77
2009-05-27 | 23:00 | 17.05 | 35 | 54 | 8.54 | 88.4 | 7.73
2009-05-27 | 23:15 | 17.03 | 35 | 54 | 8.48 | 87.8 | 7.7
2009-05-27 | 23:30 | 16.99 | 35 | 54 | 8.44 | 87.3 | 7.67
2009-05-27 | 23:45 | 16.95 | 35 | 54 | 8.46 | 87.4 | 7.65
2009-05-28 | 0:00 | 16.92 | 35 | 54 | 8.44 | 87.2 | 7.63
2009-05-28 | 0:15 | 16.88 | 35 | 54 | 8.41 | 86.8 | 7.61
2009-05-28 | 0:30 | 16.85 | 35 | 54 | 8.41 | 86.7 | 7.6
2009-05-28 | 0:45 | 16.82 | 35 | 54 | 8.41 | 86.7 | 7.59
2009-05-28 | 1:01 | 16.78 | 35 | 54 | 8.39 | 86.5 | 7.58
2009-05-28 | 1:16 | 16.75 | 35 | 54 | 8.37 | 86.2 | 7.57
2009-05-28 | 1:31 | 16.72 | 35 | 53 | 8.39 | 86.3 | 7.55
2009-05-28 | 1:46 | 16.68 | 35 | 53 | 8.38 | 86.1 | 7.55
2009-05-28 | 2:01 | 16.63 | 35 | 53 | 8.4 | 86.3 | 7.54
2009-05-28 | 2:16 | 16.59 | 35 | 53 | 8.39 | 86.1 | 7.53
2009-05-28 | 2:31 | 16.54 | 35 | 53 | 8.42 | 86.3 | 7.53
2009-05-28 | 2:46 | 16.49 | 35 | 53 | 8.44 | 86.4 | 7.53
2009-05-28 | 3:01 | 16.44 | 35 | 53 | 8.42 | 86.1 | 7.52
2009-05-28 | 3:16 | 16.39 | 34 | 53 | 8.45 | 86.4 | 7.52
2009-05-28 | 3:31 | 16.34 | 34 | 53 | 8.47 | 86.4 | 7.51
2009-05-28 | 3:46 | 16.29 | 34 | 53 | 8.48 | 86.4 | 7.51
2009-05-28 | 4:01 | 16.23 | 34 | 53 | 8.49 | 86.5 | 7.5
2009-05-28 | 4:16 | 16.18 | 34 | 53 | 8.49 | 86.3 | 7.5
2009-05-28 | 4:31 | 16.12 | 34 | 53 | 8.54 | 86.7 | 7.5
2009-05-28 | 4:46 | 16.07 | 34 | 53 | 8.52 | 86.4 | 7.49
2009-05-28 | 5:01 | 16.01 | 34 | 52 | 8.52 | 86.4 | 7.49
2009-05-28 | 5:16 | 15.94 | 34 | 52 | 8.53 | 86.3 | 7.49
2009-05-28 | 5:31 | 15.88 | 34 | 52 | 8.56 | 86.5 | 7.49
2009-05-28 | 5:46 | 15.82 | 34 | 52 | 8.59 | 86.7 | 7.49
2009-05-28 | 6:01 | 15.76 | 34 | 52 | 8.65 | 87.2 | 7.49
2009-05-28 | 6:16 | 15.7 | 34 | 52 | 8.68 | 87.4 | 7.5
2009-05-28 | 6:31 | 15.65 | 34 | 52 | 8.77 | 88.2 | 7.5
2009-05-28 | 6:46 | 15.63 | 34 | 52 | 8.85 | 88.9 | 7.52
2009-05-28 | 7:01 | 15.61 | 34 | 52 | 8.98 | 90.3 | 7.54
2009-05-28 | 7:16 | 15.6 | 34 | 52 | 9.07 | 91.1 | 7.56
2009-05-28 | 7:31 | 15.61 | 34 | 52 | 9.17 | 92.1 | 7.58
2009-05-28 | 7:46 | 15.62 | 33 | 52 | 9.32 | 93.7 | 7.6
2009-05-28 | 8:01 | 15.64 | 33 | 51 | 9.41 | 94.6 | 7.63
2009-05-28 | 8:16 | 15.68 | 33 | 51 | 9.55 | 96.1 | 7.65
2009-05-28 | 8:31 | 15.72 | 33 | 51 | 9.63 | 97 | 7.68
2009-05-28 | 8:46 | 15.78 | 33 | 51 | 9.72 | 98 | 7.7
2009-05-28 | 9:01 | 15.86 | 33 | 51 | 9.87 | 99.7 | 7.74
2009-05-28 | 9:16 | 15.95 | 33 | 51 | 10 | 101.2 | 7.78
2009-05-28 | 9:31 | 16.05 | 33 | 50 | 10.08 | 102.2 | 7.82
2009-05-28 | 9:46 | 16.16 | 33 | 50 | 10.22 | 103.9 | 7.88
2009-05-28 | 10:01 | 16.26 | 32 | 50 | 10.27 | 104.6 | 7.93
2009-05-28 | 10:16 | 16.38 | 32 | 50 | 10.4 | 106.2 | 7.98
2009-05-28 | 10:31 | 16.5 | 32 | 49 | 10.44 | 106.8 | 8.04
2009-05-28 | 10:46 | 16.61 | 32 | 49 | 10.49 | 107.7 | 8.09
2009-05-28 | 11:01 | 16.73 | 32 | 49 | 10.55 | 108.5 | 8.14
2009-05-28 | 11:16 | 16.85 | 32 | 49 | 10.59 | 109.3 | 8.2
2009-05-28 | 11:31 | 16.98 | 32 | 49 | 10.66 | 110.3 | 8.27
2009-05-28 | 11:46 | 17.11 | 31 | 48 | 10.72 | 111.2 | 8.32
2009-05-28 | 12:01 | 17.24 | 31 | 48 | 10.71 | 111.4 | 8.38
2009-05-28 | 12:16 | 17.37 | 31 | 48 | 10.76 | 112.2 | 8.45
2009-05-28 | 12:31 | 17.49 | 31 | 48 | 10.79 | 112.8 | 8.51
2009-05-28 | 12:46 | 17.63 | 31 | 48 | 10.81 | 113.3 | 8.57
2009-05-28 | 13:00 | 17.76 | 31 | 48 | 10.76 | 113.1 | 8.62
2009-05-28 | 13:15 | 17.88 | 31 | 48 | 10.79 | 113.7 | 8.66
2009-05-28 | 13:30 | 18.01 | 31 | 48 | 10.79 | 114 | 8.7
2009-05-28 | 13:45 | 18.12 | 31 | 48 | 10.74 | 113.8 | 8.74
2009-05-28 | 14:00 | 18.22 | 31 | 48 | 10.77 | 114.2 | 8.77
2009-05-28 | 14:15 | 18.33 | 31 | 48 | 10.76 | 114.4 | 8.8
2009-05-28 | 14:30 | 18.43 | 31 | 48 | 10.69 | 114 | 8.83
2009-05-28 | 14:45 | 18.52 | 31 | 48 | 10.69 | 114.1 | 8.85
2009-05-28 | 15:00 | 18.61 | 31 | 48 | 10.64 | 113.8 | 8.87
2009-05-28 | 15:15 | 18.68 | 31 | 48 | 10.59 | 113.4 | 8.9
2009-05-28 | 15:30 | 18.75 | 31 | 48 | 10.58 | 113.5 | 8.9
2009-05-28 | 15:45 | 18.81 | 31 | 48 | 10.52 | 113 | 8.92
2009-05-28 | 16:00 | 18.87 | 32 | 48 | 10.46 | 112.5 | 8.93
2009-05-28 | 16:15 | 18.92 | 32 | 49 | 10.47 | 112.7 | 8.93
2009-05-28 | 16:30 | 18.96 | 32 | 49 | 10.38 | 111.8 | 8.93
2009-05-28 | 16:45 | 19.01 | 32 | 49 | 10.31 | 111.2 | 8.94
2009-05-28 | 17:00 | 19.03 | 32 | 49 | 10.3 | 111.1 | 8.94
2009-05-28 | 17:15 | 19.05 | 32 | 49 | 10.21 | 110.2 | 8.93
2009-05-28 | 17:30 | 19.06 | 32 | 49 | 10.13 | 109.4 | 8.94
2009-05-28 | 17:45 | 19.07 | 32 | 49 | 10.07 | 108.7 | 8.93
2009-05-28 | 18:00 | 19.04 | 32 | 49 | 9.99 | 107.8 | 8.91
2009-05-28 | 18:15 | 19.01 | 32 | 49 | 9.96 | 107.4 | 8.89
2009-05-28 | 18:30 | 18.99 | 32 | 49 | 9.89 | 106.6 | 8.87
2009-05-28 | 18:45 | 18.92 | 32 | 50 | 9.74 | 104.8 | 8.84
2009-05-28 | 19:00 | 18.85 | 32 | 50 | 9.64 | 103.7 | 8.79
2009-05-28 | 19:15 | 18.82 | 32 | 50 | 9.59 | 103 | 8.76
2009-05-28 | 19:30 | 18.77 | 32 | 50 | 9.46 | 101.5 | 8.7
2009-05-28 | 19:45 | 18.7 | 33 | 50 | 9.38 | 100.5 | 8.63
2009-05-28 | 20:00 | 18.64 | 33 | 50 | 9.27 | 99.2 | 8.56
2009-05-28 | 20:15 | 18.58 | 33 | 51 | 9.13 | 97.6 | 8.46
2009-05-28 | 20:30 | 18.52 | 33 | 51 | 9.03 | 96.4 | 8.36
2009-05-28 | 20:45 | 18.45 | 33 | 51 | 8.94 | 95.3 | 8.24
2009-05-28 | 21:00 | 18.39 | 33 | 51 | 8.79 | 93.6 | 8.12
2009-05-28 | 21:15 | 18.34 | 34 | 52 | 8.71 | 92.7 | 8.02
2009-05-28 | 21:30 | 18.3 | 34 | 52 | 8.63 | 91.7 | 7.94
2009-05-28 | 21:45 | 18.25 | 34 | 52 | 8.52 | 90.4 | 7.87
2009-05-28 | 22:00 | 18.22 | 34 | 52 | 8.44 | 89.5 | 7.81
2009-05-28 | 22:15 | 18.19 | 34 | 52 | 8.39 | 89 | 7.76
2009-05-28 | 22:30 | 18.17 | 34 | 52 | 8.31 | 88.1 | 7.73
2009-05-28 | 22:45 | 18.15 | 34 | 53 | 8.25 | 87.4 | 7.69
2009-05-28 | 23:00 | 18.12 | 34 | 53 | 8.23 | 87.2 | 7.66
2009-05-28 | 23:15 | 18.1 | 34 | 53 | 8.2 | 86.8 | 7.63
2009-05-28 | 23:30 | 18.08 | 34 | 53 | 8.17 | 86.5 | 7.61
2009-05-28 | 23:45 | 18.05 | 34 | 53 | 8.16 | 86.3 | 7.6
2009-05-29 | 0:00 | 18.01 | 34 | 53 | 8.14 | 86 | 7.58
2009-05-29 | 0:15 | 17.97 | 34 | 53 | 8.1 | 85.5 | 7.57
2009-05-29 | 0:30 | 17.96 | 34 | 53 | 8.11 | 85.6 | 7.56
2009-05-29 | 0:45 | 17.96 | 34 | 53 | 8.08 | 85.3 | 7.55
2009-05-29 | 1:01 | 17.95 | 34 | 53 | 8.08 | 85.2 | 7.54
2009-05-29 | 1:16 | 17.94 | 34 | 53 | 8.09 | 85.4 | 7.53
2009-05-29 | 1:31 | 17.93 | 34 | 52 | 8.09 | 85.4 | 7.52
2009-05-29 | 1:46 | 17.91 | 34 | 52 | 8.1 | 85.4 | 7.51
2009-05-29 | 2:01 | 17.9 | 34 | 52 | 8.08 | 85.2 | 7.51
2009-05-29 | 2:16 | 17.88 | 34 | 52 | 8.09 | 85.2 | 7.5
2009-05-29 | 2:31 | 17.87 | 34 | 52 | 8.06 | 84.9 | 7.49
2009-05-29 | 2:46 | 17.85 | 34 | 52 | 8.08 | 85.1 | 7.49
2009-05-29 | 3:01 | 17.83 | 34 | 52 | 8.1 | 85.3 | 7.49
2009-05-29 | 3:16 | 17.81 | 34 | 52 | 8.11 | 85.3 | 7.48
2009-05-29 | 3:31 | 17.78 | 34 | 52 | 8.12 | 85.4 | 7.48
2009-05-29 | 3:46 | 17.75 | 33 | 51 | 8.12 | 85.3 | 7.47
2009-05-29 | 4:01 | 17.72 | 33 | 51 | 8.1 | 85 | 7.46
2009-05-29 | 4:16 | 17.67 | 33 | 51 | 8.1 | 85 | 7.47
2009-05-29 | 4:31 | 17.62 | 33 | 51 | 8.1 | 84.9 | 7.46
2009-05-29 | 4:46 | 17.57 | 33 | 51 | 8.14 | 85.3 | 7.46
2009-05-29 | 5:01 | 17.52 | 33 | 51 | 8.17 | 85.4 | 7.45
2009-05-29 | 5:16 | 17.45 | 33 | 51 | 8.16 | 85.2 | 7.45
2009-05-29 | 5:31 | 17.39 | 33 | 50 | 8.18 | 85.3 | 7.45
2009-05-29 | 5:46 | 17.33 | 33 | 50 | 8.2 | 85.5 | 7.45
2009-05-29 | 6:01 | 17.27 | 33 | 50 | 8.27 | 86.1 | 7.45
2009-05-29 | 6:16 | 17.2 | 33 | 50 | 8.3 | 86.2 | 7.46
2009-05-29 | 6:31 | 17.14 | 33 | 50 | 8.35 | 86.6 | 7.46
2009-05-29 | 6:46 | 17.08 | 32 | 50 | 8.41 | 87.2 | 7.46
2009-05-29 | 7:01 | 17.04 | 32 | 50 | 8.54 | 88.4 | 7.48
2009-05-29 | 7:16 | 17.02 | 32 | 50 | 8.66 | 89.7 | 7.49
2009-05-29 | 7:31 | 17 | 32 | 50 | 8.75 | 90.5 | 7.52
2009-05-29 | 7:46 | 16.99 | 32 | 49 | 8.86 | 91.6 | 7.53
2009-05-29 | 8:01 | 17 | 32 | 49 | 8.96 | 92.7 | 7.56
2009-05-29 | 8:16 | 17.03 | 32 | 49 | 9.07 | 93.9 | 7.59
2009-05-29 | 8:31 | 17.07 | 32 | 49 | 9.23 | 95.7 | 7.61
2009-05-29 | 8:46 | 17.13 | 32 | 49 | 9.36 | 97.1 | 7.65
2009-05-29 | 9:01 | 17.2 | 32 | 49 | 9.48 | 98.5 | 7.68
2009-05-29 | 9:16 | 17.27 | 31 | 48 | 9.6 | 99.9 | 7.72
2009-05-29 | 9:31 | 17.36 | 31 | 48 | 9.71 | 101.2 | 7.76
2009-05-29 | 9:46 | 17.45 | 31 | 48 | 9.77 | 102.1 | 7.8
2009-05-29 | 10:01 | 17.54 | 31 | 48 | 9.86 | 103.2 | 7.85
2009-05-29 | 10:16 | 17.64 | 31 | 48 | 9.98 | 104.6 | 7.89
2009-05-29 | 10:31 | 17.74 | 31 | 48 | 10.07 | 105.8 | 7.93
2009-05-29 | 10:46 | 17.84 | 31 | 47 | 10.15 | 106.8 | 7.98
2009-05-29 | 11:01 | 17.94 | 31 | 47 | 10.22 | 107.9 | 8.03
2009-05-29 | 11:16 | 18.05 | 31 | 47 | 10.27 | 108.7 | 8.08
2009-05-29 | 11:31 | 18.16 | 30 | 47 | 10.28 | 108.9 | 8.13
2009-05-29 | 11:46 | 18.29 | 30 | 47 | 10.31 | 109.5 | 8.2
2009-05-29 | 12:01 | 18.41 | 30 | 47 | 10.37 | 110.5 | 8.25
2009-05-29 | 12:16 | 18.54 | 30 | 47 | 10.41 | 111.2 | 8.31
2009-05-29 | 12:31 | 18.69 | 30 | 47 | 10.44 | 111.9 | 8.37
2009-05-29 | 12:46 | 18.66 | 30 | 47 | 10.35 | 110.8 | 8.35
2009-05-29 | 13:00 | 18.71 | 30 | 47 | 10.28 | 110.2 | 8.39
2009-05-29 | 13:15 | 18.78 | 30 | 47 | 10.28 | 110.3 | 8.41
2009-05-29 | 13:30 | 18.81 | 30 | 47 | 10.27 | 110.3 | 8.43
2009-05-29 | 13:45 | 18.71 | 30 | 47 | 10.14 | 108.7 | 8.41
2009-05-29 | 14:00 | 18.65 | 31 | 47 | 10.1 | 108.1 | 8.41
2009-05-29 | 14:15 | 18.82 | 31 | 47 | 10.18 | 109.3 | 8.47
2009-05-29 | 14:30 | 18.84 | 31 | 47 | 10.12 | 108.8 | 8.48
2009-05-29 | 14:45 | 18.96 | 31 | 47 | 10.17 | 109.5 | 8.51
2009-05-29 | 15:00 | 19.07 | 31 | 48 | 10.12 | 109.2 | 8.52
2009-05-29 | 15:15 | 19.02 | 31 | 48 | 10.05 | 108.4 | 8.48
2009-05-29 | 15:30 | 18.92 | 31 | 48 | 9.91 | 106.6 | 8.41
2009-05-29 | 15:45 | 18.89 | 31 | 48 | 9.78 | 105.2 | 8.36
2009-05-29 | 16:00 | 18.86 | 31 | 48 | 9.7 | 104.3 | 8.33
2009-05-29 | 16:15 | 18.83 | 31 | 48 | 9.66 | 103.8 | 8.29
2009-05-29 | 16:30 | 18.82 | 31 | 48 | 9.58 | 102.9 | 8.25
2009-05-29 | 16:45 | 18.81 | 31 | 48 | 9.48 | 101.8 | 8.23
2009-05-29 | 17:00 | 18.8 | 31 | 48 | 9.35 | 100.4 | 8.18
2009-05-29 | 17:15 | 18.77 | 32 | 49 | 9.23 | 99.1 | 8.1
2009-05-29 | 17:30 | 18.7 | 32 | 49 | 9.13 | 97.8 | 8.03
2009-05-29 | 17:45 | 18.64 | 32 | 49 | 9.04 | 96.8 | 7.97
2009-05-29 | 18:00 | 18.61 | 32 | 49 | 8.98 | 96 | 7.92
2009-05-29 | 18:15 | 18.6 | 32 | 49 | 8.95 | 95.7 | 7.89
2009-05-29 | 18:30 | 18.61 | 32 | 49 | 8.91 | 95.2 | 7.86
2009-05-29 | 18:45 | 18.63 | 32 | 50 | 8.88 | 95 | 7.82
2009-05-29 | 19:00 | 18.68 | 32 | 50 | 8.83 | 94.6 | 7.8
2009-05-29 | 19:15 | 18.72 | 32 | 50 | 8.75 | 93.7 | 7.77
2009-05-29 | 19:30 | 18.73 | 32 | 50 | 8.68 | 93.1 | 7.74
2009-05-29 | 19:45 | 18.74 | 32 | 50 | 8.64 | 92.7 | 7.72
2009-05-29 | 20:00 | 18.75 | 32 | 50 | 8.58 | 92 | 7.7
2009-05-29 | 20:15 | 18.74 | 33 | 50 | 8.55 | 91.7 | 7.68
2009-05-29 | 20:30 | 18.74 | 33 | 50 | 8.54 | 91.6 | 7.66
2009-05-29 | 20:45 | 18.73 | 33 | 50 | 8.49 | 91 | 7.64
2009-05-29 | 21:00 | 18.72 | 33 | 50 | 8.44 | 90.5 | 7.62
2009-05-29 | 21:15 | 18.72 | 33 | 50 | 8.36 | 89.6 | 7.6
2009-05-29 | 21:30 | 18.7 | 33 | 50 | 8.29 | 88.8 | 7.58
2009-05-29 | 21:45 | 18.67 | 33 | 50 | 8.26 | 88.5 | 7.56
2009-05-29 | 22:00 | 18.65 | 33 | 50 | 8.2 | 87.8 | 7.54
2009-05-29 | 22:15 | 18.61 | 32 | 50 | 8.15 | 87.2 | 7.53
2009-05-29 | 22:30 | 18.58 | 32 | 50 | 8.16 | 87.2 | 7.52
2009-05-29 | 22:45 | 18.54 | 32 | 50 | 8.14 | 86.9 | 7.51
2009-05-29 | 23:00 | 18.49 | 32 | 50 | 8.12 | 86.7 | 7.49
2009-05-29 | 23:15 | 18.44 | 32 | 50 | 8.12 | 86.5 | 7.49
2009-05-29 | 23:30 | 18.39 | 32 | 49 | 8.13 | 86.5 | 7.48
2009-05-29 | 23:45 | 18.33 | 32 | 49 | 8.11 | 86.3 | 7.47
2009-05-30 | 0:00 | 18.27 | 32 | 49 | 8.12 | 86.3 | 7.46
2009-05-30 | 0:15 | 18.2 | 32 | 49 | 8.1 | 85.9 | 7.46
2009-05-30 | 0:30 | 18.13 | 32 | 49 | 8.13 | 86.1 | 7.45
2009-05-30 | 0:45 | 18.06 | 32 | 49 | 8.15 | 86.2 | 7.44
2009-05-30 | 1:01 | 17.98 | 31 | 48 | 8.14 | 85.9 | 7.44
2009-05-30 | 1:16 | 17.91 | 31 | 48 | 8.17 | 86.1 | 7.44
2009-05-30 | 1:31 | 17.83 | 31 | 48 | 8.19 | 86.2 | 7.44
2009-05-30 | 1:46 | 17.75 | 31 | 48 | 8.23 | 86.4 | 7.44
2009-05-30 | 2:01 | 17.66 | 31 | 48 | 8.24 | 86.5 | 7.43
2009-05-30 | 2:16 | 17.58 | 31 | 48 | 8.27 | 86.6 | 7.43
2009-05-30 | 2:31 | 17.5 | 31 | 48 | 8.26 | 86.3 | 7.43
2009-05-30 | 2:46 | 17.42 | 31 | 47 | 8.27 | 86.3 | 7.42
2009-05-30 | 3:01 | 17.34 | 31 | 47 | 8.29 | 86.4 | 7.43
2009-05-30 | 3:16 | 17.27 | 31 | 47 | 8.3 | 86.4 | 7.42
2009-05-30 | 3:31 | 17.19 | 31 | 47 | 8.32 | 86.4 | 7.42
2009-05-30 | 3:46 | 17.12 | 31 | 47 | 8.31 | 86.2 | 7.42
2009-05-30 | 4:01 | 17.04 | 31 | 47 | 8.35 | 86.5 | 7.42
2009-05-30 | 4:16 | 16.97 | 31 | 47 | 8.4 | 86.9 | 7.42
2009-05-30 | 4:31 | 16.89 | 30 | 47 | 8.43 | 87 | 7.41
2009-05-30 | 4:46 | 16.81 | 30 | 47 | 8.47 | 87.3 | 7.42
2009-05-30 | 5:01 | 16.74 | 30 | 47 | 8.5 | 87.5 | 7.41
2009-05-30 | 5:16 | 16.66 | 30 | 47 | 8.48 | 87.1 | 7.42
2009-05-30 | 5:31 | 16.6 | 30 | 47 | 8.54 | 87.6 | 7.42
2009-05-30 | 5:46 | 16.53 | 30 | 47 | 8.54 | 87.5 | 7.42
2009-05-30 | 6:01 | 16.46 | 30 | 46 | 8.58 | 87.8 | 7.42
2009-05-30 | 6:16 | 16.41 | 30 | 46 | 8.66 | 88.5 | 7.43
2009-05-30 | 6:31 | 16.35 | 30 | 46 | 8.74 | 89.2 | 7.43
2009-05-30 | 6:46 | 16.31 | 30 | 46 | 8.82 | 90 | 7.45
2009-05-30 | 7:01 | 16.28 | 30 | 46 | 8.87 | 90.4 | 7.46
2009-05-30 | 7:16 | 16.26 | 30 | 46 | 8.99 | 91.6 | 7.47
2009-05-30 | 7:31 | 16.25 | 30 | 46 | 9.1 | 92.6 | 7.49
2009-05-30 | 7:46 | 16.25 | 30 | 46 | 9.17 | 93.4 | 7.51
2009-05-30 | 8:01 | 16.26 | 30 | 46 | 9.27 | 94.4 | 7.54
2009-05-30 | 8:16 | 16.28 | 30 | 46 | 9.37 | 95.6 | 7.55
2009-05-30 | 8:31 | 16.32 | 30 | 46 | 9.51 | 97 | 7.58
2009-05-30 | 8:46 | 16.37 | 30 | 45 | 9.64 | 98.4 | 7.61
2009-05-30 | 9:01 | 16.42 | 30 | 45 | 9.71 | 99.2 | 7.64
2009-05-30 | 9:16 | 16.49 | 29 | 45 | 9.8 | 100.3 | 7.67
2009-05-30 | 9:31 | 16.56 | 29 | 45 | 9.93 | 101.8 | 7.71
2009-05-30 | 9:46 | 16.64 | 29 | 45 | 10.03 | 103.1 | 7.74
2009-05-30 | 10:01 | 16.74 | 29 | 45 | 10.08 | 103.7 | 7.78
2009-05-30 | 10:16 | 16.85 | 29 | 45 | 10.14 | 104.6 | 7.81
2009-05-30 | 10:31 | 16.96 | 29 | 45 | 10.25 | 106 | 7.85
2009-05-30 | 10:46 | 17.1 | 29 | 45 | 10.26 | 106.4 | 7.89
2009-05-30 | 11:01 | 17.24 | 29 | 45 | 10.3 | 107.1 | 7.94
2009-05-30 | 11:16 | 17.35 | 29 | 45 | 10.34 | 107.8 | 7.98
2009-05-30 | 11:31 | 17.49 | 29 | 45 | 10.37 | 108.4 | 8.03
2009-05-30 | 11:46 | 17.66 | 29 | 45 | 10.41 | 109.1 | 8.08
2009-05-30 | 12:01 | 17.81 | 29 | 45 | 10.47 | 110.1 | 8.13
2009-05-30 | 12:16 | 17.95 | 29 | 45 | 10.5 | 110.8 | 8.18
2009-05-30 | 12:31 | 18.11 | 29 | 45 | 10.47 | 110.9 | 8.22
2009-05-30 | 12:46 | 18.27 | 29 | 45 | 10.45 | 111 | 8.27
2009-05-30 | 13:00 | 18.41 | 29 | 45 | 10.49 | 111.7 | 8.31
2009-05-30 | 13:15 | 18.55 | 29 | 45 | 10.45 | 111.7 | 8.35
2009-05-30 | 13:30 | 18.67 | 29 | 45 | 10.44 | 111.8 | 8.39
2009-05-30 | 13:45 | 18.79 | 29 | 45 | 10.47 | 112.4 | 8.43
2009-05-30 | 14:00 | 18.91 | 29 | 45 | 10.46 | 112.5 | 8.46
2009-05-30 | 14:15 | 19.02 | 29 | 45 | 10.45 | 112.7 | 8.49
2009-05-30 | 14:30 | 19.12 | 29 | 45 | 10.39 | 112.2 | 8.52
2009-05-30 | 14:45 | 19.22 | 29 | 45 | 10.34 | 112 | 8.54
2009-05-30 | 15:00 | 19.32 | 29 | 45 | 10.3 | 111.8 | 8.56
2009-05-30 | 15:15 | 19.39 | 30 | 45 | 10.26 | 111.5 | 8.58
2009-05-30 | 15:30 | 19.34 | 30 | 46 | 10.14 | 110.1 | 8.55
2009-05-30 | 15:45 | 19.34 | 30 | 46 | 10.05 | 109.1 | 8.54
2009-05-30 | 16:00 | 19.44 | 30 | 46 | 10.09 | 109.8 | 8.56
2009-05-30 | 16:15 | 19.38 | 30 | 46 | 9.97 | 108.3 | 8.53
2009-05-30 | 16:30 | 19.32 | 30 | 46 | 9.83 | 106.7 | 8.48
2009-05-30 | 16:45 | 19.32 | 30 | 46 | 9.82 | 106.5 | 8.47
2009-05-30 | 17:00 | 19.29 | 30 | 46 | 9.68 | 105 | 8.42
2009-05-30 | 17:15 | 19.21 | 30 | 46 | 9.54 | 103.3 | 8.34
2009-05-30 | 17:30 | 19.14 | 30 | 46 | 9.44 | 102.1 | 8.25
2009-05-30 | 17:45 | 19.08 | 30 | 46 | 9.36 | 101 | 8.18
2009-05-30 | 18:00 | 19.03 | 30 | 46 | 9.28 | 100.2 | 8.12
2009-05-30 | 18:15 | 18.96 | 30 | 46 | 9.18 | 98.9 | 8.06
2009-05-30 | 18:30 | 18.89 | 30 | 46 | 9.12 | 98.1 | 7.99
2009-05-30 | 18:45 | 18.81 | 30 | 46 | 9.04 | 97.1 | 7.94
2009-05-30 | 19:00 | 18.75 | 30 | 46 | 8.92 | 95.7 | 7.87
2009-05-30 | 19:15 | 18.7 | 30 | 46 | 8.85 | 94.8 | 7.81
2009-05-30 | 19:30 | 18.64 | 30 | 46 | 8.79 | 94.1 | 7.77
2009-05-30 | 19:45 | 18.59 | 30 | 46 | 8.72 | 93.2 | 7.74
2009-05-30 | 20:00 | 18.56 | 30 | 46 | 8.68 | 92.8 | 7.71
2009-05-30 | 20:15 | 18.51 | 30 | 46 | 8.61 | 91.9 | 7.68
2009-05-30 | 20:30 | 18.46 | 30 | 46 | 8.54 | 91.1 | 7.65
2009-05-30 | 20:45 | 18.4 | 30 | 46 | 8.52 | 90.7 | 7.62
2009-05-30 | 21:00 | 18.35 | 29 | 45 | 8.45 | 89.9 | 7.59
2009-05-30 | 21:15 | 18.3 | 29 | 45 | 8.43 | 89.5 | 7.57
2009-05-30 | 21:30 | 18.24 | 29 | 45 | 8.4 | 89.1 | 7.55
2009-05-30 | 21:45 | 18.21 | 29 | 45 | 8.35 | 88.5 | 7.53
2009-05-30 | 22:00 | 18.16 | 29 | 45 | 8.35 | 88.4 | 7.5
2009-05-30 | 22:15 | 18.12 | 29 | 44 | 8.3 | 87.9 | 7.49
2009-05-30 | 22:30 | 18.07 | 29 | 44 | 8.28 | 87.6 | 7.48
2009-05-30 | 22:45 | 18.02 | 29 | 44 | 8.26 | 87.3 | 7.47
2009-05-30 | 23:00 | 17.97 | 29 | 44 | 8.25 | 87.1 | 7.46
2009-05-30 | 23:15 | 17.92 | 28 | 44 | 8.28 | 87.4 | 7.44
2009-05-30 | 23:30 | 17.86 | 28 | 44 | 8.27 | 87.1 | 7.44
2009-05-30 | 23:45 | 17.8 | 28 | 43 | 8.26 | 86.9 | 7.42
2009-05-31 | 0:00 | 17.75 | 28 | 43 | 8.25 | 86.7 | 7.42
2009-05-31 | 0:15 | 17.69 | 28 | 43 | 8.31 | 87.2 | 7.41
2009-05-31 | 0:30 | 17.63 | 28 | 43 | 8.3 | 87 | 7.41
2009-05-31 | 0:45 | 17.57 | 28 | 43 | 8.34 | 87.3 | 7.4
2009-05-31 | 1:01 | 17.51 | 28 | 43 | 8.36 | 87.4 | 7.4
2009-05-31 | 1:16 | 17.43 | 28 | 43 | 8.35 | 87.1 | 7.39
2009-05-31 | 1:31 | 17.35 | 28 | 42 | 8.39 | 87.5 | 7.39
2009-05-31 | 1:46 | 17.27 | 28 | 42 | 8.42 | 87.7 | 7.39
2009-05-31 | 2:01 | 17.19 | 27 | 42 | 8.44 | 87.7 | 7.38
2009-05-31 | 2:16 | 17.11 | 27 | 42 | 8.44 | 87.5 | 7.38
2009-05-31 | 2:31 | 17.02 | 27 | 42 | 8.48 | 87.8 | 7.38
2009-05-31 | 2:46 | 16.94 | 27 | 42 | 8.52 | 88 | 7.38
2009-05-31 | 3:01 | 16.86 | 27 | 42 | 8.55 | 88.2 | 7.37
2009-05-31 | 3:16 | 16.78 | 27 | 42 | 8.55 | 88 | 7.38
2009-05-31 | 3:31 | 16.7 | 27 | 42 | 8.59 | 88.4 | 7.37
2009-05-31 | 3:46 | 16.62 | 27 | 42 | 8.63 | 88.6 | 7.37
2009-05-31 | 4:01 | 16.54 | 27 | 42 | 8.65 | 88.6 | 7.37
2009-05-31 | 4:16 | 16.46 | 27 | 42 | 8.68 | 88.8 | 7.36
2009-05-31 | 4:31 | 16.38 | 27 | 42 | 8.66 | 88.5 | 7.37
2009-05-31 | 4:46 | 16.3 | 27 | 42 | 8.73 | 89 | 7.36
2009-05-31 | 5:01 | 16.22 | 27 | 42 | 8.72 | 88.8 | 7.36
2009-05-31 | 5:16 | 16.14 | 27 | 42 | 8.74 | 88.8 | 7.36
2009-05-31 | 5:31 | 16.06 | 27 | 42 | 8.8 | 89.3 | 7.37
2009-05-31 | 5:46 | 15.97 | 27 | 42 | 8.82 | 89.3 | 7.37
2009-05-31 | 6:01 | 15.89 | 27 | 42 | 8.84 | 89.4 | 7.37
2009-05-31 | 6:16 | 15.81 | 27 | 42 | 8.9 | 89.8 | 7.37
2009-05-31 | 6:31 | 15.75 | 27 | 42 | 8.99 | 90.6 | 7.39
2009-05-31 | 6:46 | 15.68 | 27 | 42 | 9.08 | 91.4 | 7.39
2009-05-31 | 7:01 | 15.62 | 27 | 42 | 9.16 | 92.1 | 7.4
2009-05-31 | 7:16 | 15.56 | 27 | 42 | 9.21 | 92.5 | 7.42
2009-05-31 | 7:31 | 15.51 | 27 | 42 | 9.31 | 93.4 | 7.42
2009-05-31 | 7:46 | 15.47 | 27 | 42 | 9.35 | 93.6 | 7.43
2009-05-31 | 8:01 | 15.43 | 27 | 41 | 9.46 | 94.7 | 7.45
2009-05-31 | 8:16 | 15.41 | 27 | 41 | 9.52 | 95.2 | 7.46
2009-05-31 | 8:31 | 15.39 | 27 | 41 | 9.58 | 95.8 | 7.47
2009-05-31 | 8:46 | 15.4 | 27 | 41 | 9.68 | 96.8 | 7.5
2009-05-31 | 9:01 | 15.43 | 27 | 41 | 9.83 | 98.4 | 7.53
2009-05-31 | 9:16 | 15.45 | 27 | 41 | 9.86 | 98.8 | 7.55
2009-05-31 | 9:31 | 15.44 | 27 | 41 | 9.87 | 98.8 | 7.55
2009-05-31 | 9:46 | 15.45 | 27 | 41 | 9.9 | 99.2 | 7.56
2009-05-31 | 10:01 | 15.46 | 27 | 41 | 9.94 | 99.5 | 7.56
2009-05-31 | 10:16 | 15.49 | 27 | 41 | 9.95 | 99.7 | 7.58
2009-05-31 | 10:31 | 15.58 | 27 | 41 | 10.07 | 101.1 | 7.62
2009-05-31 | 10:46 | 15.71 | 27 | 41 | 10.16 | 102.3 | 7.66
2009-05-31 | 11:01 | 15.88 | 27 | 41 | 10.25 | 103.5 | 7.7