## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited against Pro-Series (loner), Sonde AA, and NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-04 |  | 2009-08-04 |  | A
Time | Computer set to atomic time |  | 10:28 |  | 10:25 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.43 |  | A
DO | In Situ |  | 8.85 |  | 9.15 |  | A
SpC | In Situ |  | 57.8 |  | 58 |  | A
Ph | In Situ |  | 7.98 |  | 8.16 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-08-04 10:31 | 17.4 | 7.99 | 8.88 | 92.7 | 57.4 | 37
Sonde | 2009-08-04 10:30 | 17.46 | 8.17 | 9.08 | 94.8 | 58 | 38
Grade |  | A | A | A |  | A
Audit 2 | 2009-08-12 09:15 | 16.5 | 7.96 | 9.35 | 95.8 | 70.4 | 45.5
Sonde | 2009-08-12 09:15 | 16.54 | 8.2 | 9.39 | 96.3 | 69 | 45
Grade |  | A | A | A |  | A
Audit 3 | 2009-08-24 12:35 | 13.6 | 8.15 | 10.29 | 99.1 | 71 | 46.2
Sonde | 2009-08-24 12:30 | 13.58 | 8.27 | 10.5 | 100.9 | 71 | 46
Grade |  | A | A | A |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited against Pro-Series (loner), Sonde AA, and NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-24 |  | 2009-08-24 |  | A
Time | Computer set to atomic time |  | 12:41 |  | 12:38 |  | A
Temp | VWR NIST traceable surface water |  | 13.6 |  | 13.83 |  | A
DO | In Situ |  | 9.99 |  | 10.25 |  | A
SpC | In Situ |  | 70.8 |  | 71 |  | A
Ph | In Situ |  | 8.18 |  | 8.29 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-27 |  | 2009-07-27 |  | A
Time | Computer set to atomic time |  | 15:31 |  | 15:31 |  | A
Temp | VWR NIST traceable surface water |  | 22.1 |  | 21.7 |  | A
DO | Winkler titration of surface water |  | 8.75 |  | 8.92 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 158.6 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-08-04 10:31 | 17.4 | 7.99 | 8.88 | 92.7 | 57.4 | 37
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-04 |  | 2009-08-04 |  | A
Time | Computer set to atomic time |  | 11:35 |  | 11:35 |  | A
Temp | VWR NIST traceable surface water |  | 13.1 |  | 13.3 |  | A
DO | Winkler titration of surface water |  | 9.6 |  | 9.79 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 184.6 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
for the Pre-Audit Failing SpC, I'm sure the SpC is working properly, we had to use old solution for this audit |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-04 |  | 2009-08-04 |  | A
Time | Computer set to atomic time |  | 11:35 |  | 11:35 |  | A
Temp | VWR NIST traceable surface water |  | 13.1 |  | 13.3 |  | A
DO | Winkler titration of surface water |  | 9.6 |  | 9.79 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 184.6 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-08-12 09:15 | 16.5 | 7.96 | 9.35 | 95.8 | 70.4 | 45.5
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-12 |  | 2009-08-12 |  | A
Time | Computer set to atomic time |  | 10:15 |  | 10:15 |  | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.54 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 150.6 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.1 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
for the Pre-Audit Failing SpC, I'm sure the SpC is working properly, we had to use old solution for this audit |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-21 |  | 2009-08-21 |  | A
Time | Computer set to atomic time |  | 13:39 |  | 13:39 |  | A
Temp | VWR NIST traceable surface water |  | 15.1 |  | 15.3 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 9.47 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 162.3 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-08-24 12:35 | 13.6 | 8.15 | 10.29 | 99.1 | 71 | 46.2
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-08-24 |  | 2009-08-24 |  | A
Time | Computer set to atomic time |  | 12:47 |  | 12:47 |  | A
Temp | VWR NIST traceable surface water |  | 13.9 |  | 14 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 9.86 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 146.4 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.16 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | SpCond | TDS | DO Conc | DO% | pH
M/D/Y |  | °C | uS/cm | mg/L | mg/L | %
2009-08-04 | 10:30 | 17.46 | 58 | 38 | 9.08 | 94.8 | 8.17
2009-08-04 | 10:46 | 17.62 | 57 | 37 | 9.23 | 96.8 | 8.18
2009-08-04 | 11:01 | 17.85 | 57 | 37 | 9.26 | 97.5 | 8.19
2009-08-04 | 11:16 | 18.11 | 57 | 37 | 9.21 | 97.5 | 8.21
2009-08-04 | 11:31 | 18.43 | 56 | 37 | 9.19 | 97.9 | 8.22
2009-08-04 | 11:46 | 18.78 | 56 | 36 | 9.12 | 97.9 | 8.24
2009-08-04 | 12:01 | 19.17 | 56 | 36 | 9.06 | 98 | 8.23
2009-08-04 | 12:16 | 19.57 | 56 | 36 | 8.98 | 97.9 | 8.25
2009-08-04 | 12:31 | 19.97 | 55 | 36 | 8.87 | 97.5 | 8.27
2009-08-04 | 12:46 | 20.42 | 55 | 36 | 8.82 | 97.8 | 8.29
2009-08-04 | 13:01 | 20.81 | 55 | 36 | 8.71 | 97.3 | 8.29
2009-08-04 | 13:16 | 21.28 | 55 | 36 | 8.65 | 97.6 | 8.31
2009-08-04 | 13:31 | 21.59 | 55 | 36 | 8.5 | 96.4 | 8.31
2009-08-04 | 13:46 | 21.86 | 55 | 36 | 8.43 | 96.1 | 8.29
2009-08-04 | 14:01 | 22.1 | 55 | 36 | 8.33 | 95.5 | 8.28
2009-08-04 | 14:16 | 22.06 | 55 | 35 | 8.31 | 95.2 | 8.25
2009-08-04 | 14:31 | 22.14 | 54 | 35 | 8.23 | 94.4 | 8.23
2009-08-04 | 14:46 | 22.23 | 54 | 35 | 8.26 | 94.9 | 8.25
2009-08-04 | 15:01 | 22.23 | 54 | 35 | 8.32 | 95.6 | 8.27
2009-08-04 | 15:16 | 22.39 | 54 | 35 | 8.33 | 96 | 8.28
2009-08-04 | 15:31 | 22.63 | 54 | 35 | 8.29 | 96 | 8.31
2009-08-04 | 15:46 | 22.83 | 54 | 35 | 8.28 | 96.2 | 8.32
2009-08-04 | 16:01 | 23.03 | 54 | 35 | 8.25 | 96.2 | 8.34
2009-08-04 | 16:16 | 23.12 | 54 | 35 | 8.17 | 95.4 | 8.33
2009-08-04 | 16:31 | 23.15 | 54 | 35 | 8.18 | 95.6 | 8.33
2009-08-04 | 16:46 | 23.11 | 54 | 35 | 8.18 | 95.6 | 8.33
2009-08-04 | 17:01 | 23.01 | 54 | 35 | 8.18 | 95.4 | 8.31
2009-08-04 | 17:16 | 22.87 | 54 | 35 | 8.18 | 95.1 | 8.3
2009-08-04 | 17:31 | 22.7 | 54 | 35 | 8.2 | 95.1 | 8.27
2009-08-04 | 17:46 | 22.58 | 54 | 35 | 8.19 | 94.7 | 8.25
2009-08-04 | 18:01 | 22.45 | 54 | 35 | 8.19 | 94.4 | 8.24
2009-08-04 | 18:16 | 22.3 | 54 | 35 | 8.15 | 93.8 | 8.22
2009-08-04 | 18:31 | 22.22 | 54 | 35 | 8.14 | 93.5 | 8.2
2009-08-04 | 18:46 | 22.12 | 54 | 35 | 8.12 | 93.1 | 8.19
2009-08-04 | 19:01 | 22.05 | 54 | 35 | 8.11 | 92.9 | 8.17
2009-08-04 | 19:16 | 21.96 | 54 | 35 | 8.1 | 92.6 | 8.16
2009-08-04 | 19:31 | 21.87 | 54 | 35 | 8.08 | 92.2 | 8.14
2009-08-04 | 19:46 | 21.79 | 54 | 35 | 8.09 | 92.1 | 8.13
2009-08-04 | 20:01 | 21.7 | 54 | 35 | 8.08 | 91.9 | 8.11
2009-08-04 | 20:16 | 21.58 | 54 | 35 | 8.12 | 92.1 | 8.09
2009-08-04 | 20:31 | 21.45 | 54 | 35 | 8.14 | 92.1 | 8.08
2009-08-04 | 20:46 | 21.34 | 55 | 35 | 8.14 | 91.9 | 8.06
2009-08-04 | 21:01 | 21.18 | 55 | 35 | 8.13 | 91.5 | 8.05
2009-08-04 | 21:16 | 21.04 | 55 | 36 | 8.14 | 91.4 | 8.04
2009-08-04 | 21:31 | 20.89 | 55 | 36 | 8.19 | 91.7 | 8.04
2009-08-04 | 21:46 | 20.74 | 55 | 36 | 8.23 | 91.9 | 8.03
2009-08-04 | 22:01 | 20.58 | 55 | 36 | 8.26 | 91.9 | 8.03
2009-08-04 | 22:16 | 20.4 | 55 | 36 | 8.3 | 92 | 8.02
2009-08-04 | 22:31 | 20.21 | 55 | 36 | 8.3 | 91.7 | 8.02
2009-08-04 | 22:46 | 20.02 | 55 | 36 | 8.32 | 91.6 | 8.02
2009-08-04 | 23:01 | 19.84 | 55 | 36 | 8.38 | 91.9 | 8.02
2009-08-04 | 23:16 | 19.64 | 55 | 36 | 8.43 | 92.1 | 8.02
2009-08-04 | 23:31 | 19.45 | 55 | 36 | 8.48 | 92.3 | 8.01
2009-08-04 | 23:46 | 19.27 | 55 | 36 | 8.48 | 91.9 | 8.01
2009-08-05 | 0:01 | 19.08 | 56 | 36 | 8.5 | 91.8 | 8.01
2009-08-05 | 0:16 | 18.91 | 56 | 36 | 8.53 | 91.8 | 8.01
2009-08-05 | 0:31 | 18.73 | 56 | 36 | 8.59 | 92.1 | 8.01
2009-08-05 | 0:46 | 18.57 | 56 | 36 | 8.65 | 92.4 | 8.01
2009-08-05 | 1:01 | 18.4 | 56 | 36 | 8.67 | 92.3 | 8.01
2009-08-05 | 1:16 | 18.24 | 56 | 36 | 8.71 | 92.4 | 8.02
2009-08-05 | 1:31 | 18.09 | 56 | 36 | 8.7 | 92.1 | 8.02
2009-08-05 | 1:46 | 17.93 | 56 | 37 | 8.75 | 92.3 | 8.02
2009-08-05 | 2:01 | 17.77 | 56 | 37 | 8.81 | 92.6 | 8.02
2009-08-05 | 2:16 | 17.61 | 57 | 37 | 8.84 | 92.6 | 8.01
2009-08-05 | 2:31 | 17.45 | 57 | 37 | 8.83 | 92.3 | 8.02
2009-08-05 | 2:46 | 17.29 | 57 | 37 | 8.88 | 92.5 | 8.02
2009-08-05 | 3:01 | 17.13 | 57 | 37 | 8.93 | 92.7 | 8.02
2009-08-05 | 3:16 | 16.97 | 57 | 37 | 8.93 | 92.4 | 8.02
2009-08-05 | 3:31 | 16.82 | 58 | 37 | 8.98 | 92.6 | 8.02
2009-08-05 | 3:46 | 16.66 | 58 | 37 | 8.99 | 92.4 | 8.02
2009-08-05 | 4:01 | 16.51 | 58 | 38 | 9.01 | 92.3 | 8.02
2009-08-05 | 4:16 | 16.36 | 58 | 38 | 9.04 | 92.2 | 8.02
2009-08-05 | 4:31 | 16.2 | 58 | 38 | 9.1 | 92.6 | 8.02
2009-08-05 | 4:46 | 16.06 | 58 | 38 | 9.11 | 92.4 | 8.02
2009-08-05 | 5:01 | 15.92 | 59 | 38 | 9.18 | 92.8 | 8.02
2009-08-05 | 5:16 | 15.77 | 59 | 38 | 9.17 | 92.5 | 8.02
2009-08-05 | 5:31 | 15.62 | 59 | 38 | 9.23 | 92.8 | 8.02
2009-08-05 | 5:46 | 15.48 | 59 | 38 | 9.25 | 92.7 | 8.03
2009-08-05 | 6:01 | 15.34 | 59 | 39 | 9.31 | 93 | 8.03
2009-08-05 | 6:16 | 15.21 | 60 | 39 | 9.36 | 93.2 | 8.03
2009-08-05 | 6:31 | 15.08 | 60 | 39 | 9.41 | 93.5 | 8.04
2009-08-05 | 6:46 | 14.96 | 60 | 39 | 9.46 | 93.7 | 8.05
2009-08-05 | 7:01 | 14.84 | 60 | 39 | 9.48 | 93.7 | 8.06
2009-08-05 | 7:16 | 14.74 | 60 | 39 | 9.55 | 94.1 | 8.07
2009-08-05 | 7:31 | 14.66 | 61 | 39 | 9.6 | 94.5 | 8.08
2009-08-05 | 7:46 | 14.6 | 61 | 39 | 9.59 | 94.3 | 8.09
2009-08-05 | 8:01 | 14.55 | 61 | 40 | 9.65 | 94.7 | 8.1
2009-08-05 | 8:16 | 14.53 | 61 | 40 | 9.63 | 94.6 | 8.11
2009-08-05 | 8:31 | 14.53 | 61 | 40 | 9.68 | 95 | 8.12
2009-08-05 | 8:46 | 14.55 | 61 | 40 | 9.72 | 95.4 | 8.13
2009-08-05 | 9:01 | 14.58 | 61 | 40 | 9.74 | 95.7 | 8.14
2009-08-05 | 9:16 | 14.61 | 61 | 40 | 9.71 | 95.5 | 8.15
2009-08-05 | 9:31 | 14.66 | 61 | 39 | 9.75 | 96 | 8.15
2009-08-05 | 9:46 | 14.74 | 61 | 39 | 9.76 | 96.3 | 8.16
2009-08-05 | 10:01 | 14.83 | 60 | 39 | 9.73 | 96.1 | 8.17
2009-08-05 | 10:16 | 14.93 | 60 | 39 | 9.71 | 96.1 | 8.18
2009-08-05 | 10:31 | 15.06 | 60 | 39 | 9.69 | 96.3 | 8.19
2009-08-05 | 10:46 | 15.22 | 60 | 39 | 9.67 | 96.4 | 8.21
2009-08-05 | 11:01 | 15.4 | 60 | 39 | 9.65 | 96.5 | 8.21
2009-08-05 | 11:16 | 15.64 | 60 | 39 | 9.61 | 96.7 | 8.22
2009-08-05 | 11:31 | 15.94 | 59 | 39 | 9.6 | 97.1 | 8.22
2009-08-05 | 11:46 | 16.3 | 59 | 39 | 9.52 | 97.1 | 8.23
2009-08-05 | 12:01 | 16.7 | 59 | 38 | 9.44 | 97 | 8.24
2009-08-05 | 12:16 | 17.14 | 59 | 38 | 9.34 | 96.9 | 8.25
2009-08-05 | 12:31 | 17.59 | 59 | 38 | 9.29 | 97.3 | 8.27
2009-08-05 | 12:46 | 18.08 | 59 | 38 | 9.17 | 97.1 | 8.28
2009-08-05 | 13:01 | 18.57 | 59 | 38 | 9.07 | 97 | 8.3
2009-08-05 | 13:16 | 19.05 | 59 | 38 | 9.01 | 97.2 | 8.31
2009-08-05 | 13:31 | 19.55 | 58 | 38 | 8.93 | 97.4 | 8.32
2009-08-05 | 13:46 | 20.07 | 58 | 38 | 8.85 | 97.4 | 8.33
2009-08-05 | 14:01 | 20.54 | 58 | 38 | 8.77 | 97.5 | 8.34
2009-08-05 | 14:16 | 21.06 | 58 | 38 | 8.68 | 97.5 | 8.36
2009-08-05 | 14:31 | 21.49 | 58 | 38 | 8.56 | 96.9 | 8.38
2009-08-05 | 14:46 | 21.91 | 58 | 38 | 8.46 | 96.6 | 8.39
2009-08-05 | 15:01 | 22.24 | 58 | 38 | 8.38 | 96.3 | 8.4
2009-08-05 | 15:16 | 22.5 | 58 | 38 | 8.34 | 96.3 | 8.4
2009-08-05 | 15:31 | 22.7 | 58 | 38 | 8.3 | 96.2 | 8.4
2009-08-05 | 15:46 | 22.88 | 58 | 38 | 8.22 | 95.5 | 8.39
2009-08-05 | 16:01 | 23 | 58 | 38 | 8.2 | 95.6 | 8.39
2009-08-05 | 16:16 | 22.95 | 58 | 38 | 8.2 | 95.5 | 8.38
2009-08-05 | 16:31 | 22.9 | 58 | 38 | 8.18 | 95.2 | 8.36
2009-08-05 | 16:46 | 22.8 | 58 | 38 | 8.14 | 94.6 | 8.35
2009-08-05 | 17:01 | 22.62 | 58 | 38 | 8.17 | 94.6 | 8.33
2009-08-05 | 17:16 | 22.43 | 58 | 38 | 8.18 | 94.3 | 8.31
2009-08-05 | 17:31 | 22.24 | 58 | 38 | 8.18 | 94 | 8.28
2009-08-05 | 17:46 | 22.04 | 58 | 38 | 8.16 | 93.4 | 8.27
2009-08-05 | 18:01 | 21.89 | 58 | 38 | 8.18 | 93.4 | 8.24
2009-08-05 | 18:16 | 21.68 | 58 | 38 | 8.2 | 93.2 | 8.22
2009-08-05 | 18:31 | 21.5 | 58 | 38 | 8.18 | 92.7 | 8.2
2009-08-05 | 18:46 | 21.32 | 58 | 38 | 8.21 | 92.7 | 8.18
2009-08-05 | 19:01 | 21.18 | 58 | 38 | 8.22 | 92.6 | 8.17
2009-08-05 | 19:16 | 21.08 | 58 | 38 | 8.2 | 92.2 | 8.15
2009-08-05 | 19:31 | 21.03 | 58 | 38 | 8.18 | 91.8 | 8.14
2009-08-05 | 19:46 | 20.95 | 58 | 38 | 8.17 | 91.5 | 8.13
2009-08-05 | 20:01 | 20.82 | 58 | 38 | 8.21 | 91.7 | 8.11
2009-08-05 | 20:16 | 20.67 | 58 | 38 | 8.22 | 91.7 | 8.1
2009-08-05 | 20:31 | 20.52 | 58 | 38 | 8.24 | 91.5 | 8.09
2009-08-05 | 20:46 | 20.39 | 58 | 38 | 8.22 | 91.1 | 8.08
2009-08-05 | 21:01 | 20.26 | 58 | 38 | 8.23 | 91 | 8.07
2009-08-05 | 21:16 | 20.09 | 58 | 38 | 8.29 | 91.3 | 8.06
2009-08-05 | 21:31 | 19.97 | 58 | 38 | 8.32 | 91.4 | 8.06
2009-08-05 | 21:46 | 19.84 | 59 | 38 | 8.34 | 91.5 | 8.05
2009-08-05 | 22:01 | 19.68 | 59 | 38 | 8.35 | 91.3 | 8.06
2009-08-05 | 22:16 | 19.54 | 59 | 38 | 8.37 | 91.2 | 8.05
2009-08-05 | 22:31 | 19.39 | 59 | 38 | 8.4 | 91.2 | 8.05
2009-08-05 | 22:46 | 19.24 | 59 | 38 | 8.42 | 91.2 | 8.05
2009-08-05 | 23:01 | 19.08 | 59 | 38 | 8.45 | 91.2 | 8.05
2009-08-05 | 23:16 | 18.92 | 59 | 38 | 8.47 | 91.2 | 8.04
2009-08-05 | 23:31 | 18.74 | 59 | 38 | 8.53 | 91.5 | 8.04
2009-08-05 | 23:46 | 18.58 | 59 | 38 | 8.55 | 91.4 | 8.04
2009-08-06 | 0:01 | 18.42 | 59 | 38 | 8.59 | 91.6 | 8.04
2009-08-06 | 0:16 | 18.25 | 59 | 38 | 8.64 | 91.7 | 8.04
2009-08-06 | 0:31 | 18.08 | 59 | 38 | 8.64 | 91.5 | 8.03
2009-08-06 | 0:46 | 17.92 | 59 | 38 | 8.66 | 91.3 | 8.04
2009-08-06 | 1:01 | 17.75 | 59 | 38 | 8.72 | 91.7 | 8.03
2009-08-06 | 1:16 | 17.59 | 59 | 38 | 8.76 | 91.8 | 8.03
2009-08-06 | 1:31 | 17.44 | 59 | 38 | 8.81 | 92 | 8.03
2009-08-06 | 1:46 | 17.3 | 59 | 38 | 8.81 | 91.7 | 8.03
2009-08-06 | 2:01 | 17.16 | 59 | 39 | 8.85 | 91.9 | 8.03
2009-08-06 | 2:16 | 17.03 | 59 | 39 | 8.86 | 91.7 | 8.03
2009-08-06 | 2:31 | 16.91 | 59 | 39 | 8.91 | 92 | 8.03
2009-08-06 | 2:46 | 16.79 | 60 | 39 | 8.9 | 91.7 | 8.03
2009-08-06 | 3:01 | 16.68 | 60 | 39 | 8.95 | 92 | 8.04
2009-08-06 | 3:16 | 16.57 | 60 | 39 | 8.95 | 91.8 | 8.04
2009-08-06 | 3:31 | 16.47 | 60 | 39 | 8.99 | 92 | 8.04
2009-08-06 | 3:46 | 16.37 | 60 | 39 | 9.02 | 92.1 | 8.03
2009-08-06 | 4:01 | 16.28 | 60 | 39 | 9.05 | 92.3 | 8.03
2009-08-06 | 4:16 | 16.19 | 60 | 39 | 9.04 | 91.9 | 8.03
2009-08-06 | 4:31 | 16.11 | 60 | 39 | 9.08 | 92.2 | 8.04
2009-08-06 | 4:46 | 16.03 | 61 | 39 | 9.1 | 92.3 | 8.04
2009-08-06 | 5:01 | 15.95 | 61 | 39 | 9.12 | 92.3 | 8.04
2009-08-06 | 5:16 | 15.88 | 61 | 40 | 9.11 | 92.1 | 8.04
2009-08-06 | 5:31 | 15.83 | 61 | 40 | 9.11 | 92 | 8.04
2009-08-06 | 5:46 | 15.77 | 61 | 40 | 9.15 | 92.3 | 8.04
2009-08-06 | 6:01 | 15.72 | 62 | 40 | 9.18 | 92.5 | 8.04
2009-08-06 | 6:16 | 15.69 | 62 | 40 | 9.19 | 92.6 | 8.05
2009-08-06 | 6:31 | 15.65 | 62 | 40 | 9.17 | 92.3 | 8.06
2009-08-06 | 6:46 | 15.61 | 63 | 41 | 9.21 | 92.6 | 8.06
2009-08-06 | 7:01 | 15.56 | 63 | 41 | 9.25 | 92.9 | 8.07
2009-08-06 | 7:16 | 15.52 | 63 | 41 | 9.25 | 92.8 | 8.07
2009-08-06 | 7:31 | 15.49 | 64 | 41 | 9.32 | 93.4 | 8.1
2009-08-06 | 7:46 | 15.46 | 64 | 42 | 9.3 | 93.2 | 8.11
2009-08-06 | 8:01 | 15.44 | 64 | 42 | 9.34 | 93.5 | 8.11
2009-08-06 | 8:16 | 15.42 | 65 | 42 | 9.33 | 93.3 | 8.12
2009-08-06 | 8:31 | 15.44 | 65 | 42 | 9.36 | 93.8 | 8.15
2009-08-06 | 8:46 | 15.52 | 65 | 42 | 9.43 | 94.5 | 8.19
2009-08-06 | 9:01 | 15.6 | 66 | 43 | 9.5 | 95.5 | 8.22
2009-08-06 | 9:16 | 15.69 | 66 | 43 | 9.53 | 95.9 | 8.24
2009-08-06 | 9:31 | 15.79 | 66 | 43 | 9.49 | 95.7 | 8.27
2009-08-06 | 9:46 | 15.88 | 67 | 43 | 9.51 | 96.1 | 8.28
2009-08-06 | 10:01 | 15.92 | 67 | 43 | 9.46 | 95.7 | 8.29
2009-08-06 | 10:16 | 15.9 | 67 | 44 | 9.48 | 95.8 | 8.29
2009-08-06 | 10:31 | 15.94 | 67 | 44 | 9.46 | 95.7 | 8.3
2009-08-06 | 10:46 | 16.03 | 68 | 44 | 9.49 | 96.2 | 8.32
2009-08-06 | 11:01 | 16.13 | 68 | 44 | 9.49 | 96.5 | 8.33
2009-08-06 | 11:16 | 16.25 | 68 | 44 | 9.51 | 96.8 | 8.35
2009-08-06 | 11:31 | 16.37 | 69 | 45 | 9.5 | 97 | 8.36
2009-08-06 | 11:46 | 16.5 | 69 | 45 | 9.48 | 97 | 8.37
2009-08-06 | 12:01 | 16.66 | 69 | 45 | 9.47 | 97.3 | 8.39
2009-08-06 | 12:16 | 16.81 | 69 | 45 | 9.44 | 97.4 | 8.4
2009-08-06 | 12:31 | 16.94 | 70 | 45 | 9.38 | 96.9 | 8.4
2009-08-06 | 12:46 | 17.14 | 70 | 45 | 9.37 | 97.3 | 8.42
2009-08-06 | 13:01 | 17.3 | 70 | 45 | 9.3 | 96.9 | 8.43
2009-08-06 | 13:16 | 17.43 | 70 | 45 | 9.25 | 96.6 | 8.43
2009-08-06 | 13:31 | 17.56 | 70 | 45 | 9.2 | 96.3 | 8.43
2009-08-06 | 13:46 | 17.66 | 70 | 45 | 9.13 | 95.8 | 8.42
2009-08-06 | 14:01 | 17.68 | 70 | 45 | 9.1 | 95.5 | 8.4
2009-08-06 | 14:16 | 17.63 | 69 | 45 | 9.05 | 94.9 | 8.36
2009-08-06 | 14:31 | 17.6 | 69 | 45 | 9.04 | 94.7 | 8.35
2009-08-06 | 14:46 | 17.56 | 69 | 45 | 9.09 | 95.2 | 8.34
2009-08-06 | 15:01 | 17.5 | 69 | 45 | 9.14 | 95.5 | 8.34
2009-08-06 | 15:16 | 17.47 | 69 | 45 | 9.17 | 95.8 | 8.34
2009-08-06 | 15:31 | 17.41 | 68 | 44 | 9.18 | 95.8 | 8.34
2009-08-06 | 15:46 | 17.37 | 68 | 44 | 9.14 | 95.3 | 8.32
2009-08-06 | 16:01 | 17.3 | 68 | 44 | 9.12 | 94.9 | 8.29
2009-08-06 | 16:16 | 17.24 | 67 | 44 | 9.13 | 95 | 8.29
2009-08-06 | 16:31 | 17.22 | 67 | 44 | 9.14 | 95 | 8.28
2009-08-06 | 16:46 | 17.16 | 67 | 44 | 9.14 | 94.9 | 8.27
2009-08-06 | 17:01 | 17.08 | 67 | 43 | 9.15 | 94.8 | 8.26
2009-08-06 | 17:16 | 17 | 66 | 43 | 9.12 | 94.4 | 8.25
2009-08-06 | 17:31 | 16.92 | 66 | 43 | 9.14 | 94.4 | 8.23
2009-08-06 | 17:46 | 16.83 | 66 | 43 | 9.16 | 94.4 | 8.22
2009-08-06 | 18:01 | 16.75 | 66 | 43 | 9.18 | 94.5 | 8.22
2009-08-06 | 18:16 | 16.67 | 66 | 43 | 9.21 | 94.6 | 8.22
2009-08-06 | 18:31 | 16.55 | 66 | 43 | 9.22 | 94.5 | 8.21
2009-08-06 | 18:46 | 16.48 | 65 | 43 | 9.23 | 94.4 | 8.2
2009-08-06 | 19:01 | 16.42 | 65 | 43 | 9.23 | 94.3 | 8.19
2009-08-06 | 19:16 | 16.32 | 65 | 42 | 9.2 | 93.8 | 8.18
2009-08-06 | 19:31 | 16.24 | 65 | 42 | 9.19 | 93.6 | 8.17
2009-08-06 | 19:46 | 16.16 | 65 | 42 | 9.18 | 93.3 | 8.16
2009-08-06 | 20:01 | 16.06 | 65 | 42 | 9.17 | 93.1 | 8.16
2009-08-06 | 20:16 | 15.92 | 65 | 42 | 9.2 | 93.1 | 8.14
2009-08-06 | 20:31 | 15.8 | 65 | 42 | 9.19 | 92.7 | 8.12
2009-08-06 | 20:46 | 15.67 | 65 | 42 | 9.22 | 92.8 | 8.11
2009-08-06 | 21:01 | 15.53 | 65 | 42 | 9.22 | 92.5 | 8.11
2009-08-06 | 21:16 | 15.43 | 65 | 42 | 9.26 | 92.7 | 8.1
2009-08-06 | 21:31 | 15.32 | 65 | 42 | 9.26 | 92.5 | 8.09
2009-08-06 | 21:46 | 15.22 | 65 | 42 | 9.31 | 92.7 | 8.09
2009-08-06 | 22:01 | 15.12 | 65 | 42 | 9.35 | 93 | 8.09
2009-08-06 | 22:16 | 15.05 | 65 | 42 | 9.38 | 93.1 | 8.09
2009-08-06 | 22:31 | 14.98 | 65 | 42 | 9.36 | 92.8 | 8.08
2009-08-06 | 22:46 | 14.91 | 65 | 42 | 9.35 | 92.6 | 8.08
2009-08-06 | 23:01 | 14.85 | 65 | 42 | 9.36 | 92.5 | 8.08
2009-08-06 | 23:16 | 14.8 | 65 | 42 | 9.41 | 92.9 | 8.08
2009-08-06 | 23:31 | 14.73 | 65 | 42 | 9.4 | 92.6 | 8.08
2009-08-06 | 23:46 | 14.64 | 65 | 42 | 9.45 | 93 | 8.08
2009-08-07 | 0:01 | 14.57 | 65 | 42 | 9.48 | 93.1 | 8.08
2009-08-07 | 0:16 | 14.49 | 65 | 42 | 9.46 | 92.8 | 8.08
2009-08-07 | 0:31 | 14.43 | 65 | 43 | 9.51 | 93.1 | 8.09
2009-08-07 | 0:46 | 14.39 | 66 | 43 | 9.53 | 93.2 | 8.09
2009-08-07 | 1:01 | 14.33 | 66 | 43 | 9.54 | 93.3 | 8.08
2009-08-07 | 1:16 | 14.29 | 66 | 43 | 9.56 | 93.4 | 8.09
2009-08-07 | 1:31 | 14.24 | 66 | 43 | 9.57 | 93.4 | 8.08
2009-08-07 | 1:46 | 14.16 | 66 | 43 | 9.6 | 93.5 | 8.08
2009-08-07 | 2:01 | 14.11 | 66 | 43 | 9.58 | 93.1 | 8.09
2009-08-07 | 2:16 | 14.06 | 66 | 43 | 9.61 | 93.4 | 8.08
2009-08-07 | 2:31 | 14.02 | 66 | 43 | 9.62 | 93.4 | 8.08
2009-08-07 | 2:46 | 13.96 | 66 | 43 | 9.6 | 93.1 | 8.08
2009-08-07 | 3:01 | 13.9 | 66 | 43 | 9.6 | 93 | 8.08
2009-08-07 | 3:16 | 13.85 | 66 | 43 | 9.64 | 93.3 | 8.08
2009-08-07 | 3:31 | 13.8 | 66 | 43 | 9.63 | 93 | 8.09
2009-08-07 | 3:46 | 13.76 | 67 | 43 | 9.62 | 92.9 | 8.08
2009-08-07 | 4:01 | 13.71 | 67 | 43 | 9.63 | 92.9 | 8.09
2009-08-07 | 4:16 | 13.66 | 67 | 43 | 9.69 | 93.3 | 8.09
2009-08-07 | 4:31 | 13.6 | 67 | 43 | 9.67 | 93 | 8.09
2009-08-07 | 4:46 | 13.55 | 67 | 44 | 9.68 | 93 | 8.09
2009-08-07 | 5:01 | 13.5 | 67 | 44 | 9.69 | 92.9 | 8.09
2009-08-07 | 5:16 | 13.45 | 67 | 44 | 9.73 | 93.3 | 8.09
2009-08-07 | 5:31 | 13.4 | 68 | 44 | 9.73 | 93.1 | 8.09
2009-08-07 | 5:46 | 13.35 | 68 | 44 | 9.72 | 93 | 8.09
2009-08-07 | 6:01 | 13.29 | 68 | 44 | 9.77 | 93.4 | 8.09
2009-08-07 | 6:16 | 13.23 | 68 | 44 | 9.81 | 93.6 | 8.1
2009-08-07 | 6:31 | 13.17 | 68 | 44 | 9.81 | 93.4 | 8.1
2009-08-07 | 6:46 | 13.1 | 69 | 45 | 9.82 | 93.4 | 8.11
2009-08-07 | 7:01 | 13.06 | 69 | 45 | 9.87 | 93.9 | 8.12
2009-08-07 | 7:16 | 13.04 | 69 | 45 | 9.93 | 94.4 | 8.14
2009-08-07 | 7:31 | 13.03 | 69 | 45 | 9.93 | 94.4 | 8.15
2009-08-07 | 7:46 | 13.02 | 69 | 45 | 9.99 | 94.9 | 8.17
2009-08-07 | 8:01 | 13.01 | 70 | 45 | 10.03 | 95.2 | 8.18
2009-08-07 | 8:16 | 13.03 | 70 | 45 | 10.02 | 95.2 | 8.2
2009-08-07 | 8:31 | 13.04 | 70 | 46 | 10.03 | 95.3 | 8.21
2009-08-07 | 8:46 | 13.05 | 70 | 46 | 10.07 | 95.7 | 8.23
2009-08-07 | 9:01 | 13.1 | 71 | 46 | 10.07 | 95.8 | 8.24
2009-08-07 | 9:16 | 13.17 | 71 | 46 | 10.12 | 96.4 | 8.26
2009-08-07 | 9:31 | 13.23 | 71 | 46 | 10.14 | 96.8 | 8.28
2009-08-07 | 9:46 | 13.31 | 71 | 46 | 10.11 | 96.6 | 8.29
2009-08-07 | 10:01 | 13.36 | 71 | 46 | 10.08 | 96.5 | 8.3
2009-08-07 | 10:16 | 13.43 | 72 | 47 | 10.11 | 96.9 | 8.3
2009-08-07 | 10:31 | 13.62 | 72 | 47 | 10.08 | 97 | 8.32
2009-08-07 | 10:46 | 13.71 | 72 | 47 | 10.06 | 97 | 8.33
2009-08-07 | 11:01 | 13.86 | 72 | 47 | 10.09 | 97.6 | 8.34
2009-08-07 | 11:16 | 14.06 | 72 | 47 | 10.08 | 98 | 8.35
2009-08-07 | 11:31 | 14.29 | 72 | 47 | 10.06 | 98.3 | 8.36
2009-08-07 | 11:46 | 14.53 | 72 | 47 | 9.98 | 98 | 8.36
2009-08-07 | 12:01 | 14.79 | 72 | 47 | 9.91 | 97.9 | 8.36
2009-08-07 | 12:16 | 15.11 | 71 | 46 | 9.9 | 98.4 | 8.36
2009-08-07 | 12:31 | 15.45 | 71 | 46 | 9.84 | 98.6 | 8.37
2009-08-07 | 12:46 | 15.81 | 71 | 46 | 9.79 | 98.8 | 8.37
2009-08-07 | 13:01 | 16.04 | 70 | 46 | 9.68 | 98.1 | 8.38
2009-08-07 | 13:16 | 16.23 | 70 | 45 | 9.64 | 98.1 | 8.38
2009-08-07 | 13:31 | 16.57 | 70 | 45 | 9.54 | 97.9 | 8.4
2009-08-07 | 13:46 | 16.96 | 69 | 45 | 9.5 | 98.2 | 8.4
2009-08-07 | 14:01 | 17.3 | 69 | 45 | 9.41 | 98 | 8.4
2009-08-07 | 14:16 | 17.7 | 69 | 45 | 9.33 | 97.9 | 8.4
2009-08-07 | 14:31 | 18.01 | 69 | 45 | 9.22 | 97.4 | 8.42
2009-08-07 | 14:46 | 18.24 | 69 | 45 | 9.19 | 97.5 | 8.41
2009-08-07 | 15:01 | 18.54 | 69 | 45 | 9.14 | 97.6 | 8.42
2009-08-07 | 15:16 | 18.85 | 68 | 44 | 9.05 | 97.3 | 8.42
2009-08-07 | 15:31 | 19.11 | 68 | 44 | 9.01 | 97.3 | 8.43
2009-08-07 | 15:46 | 19.33 | 68 | 44 | 8.97 | 97.4 | 8.42
2009-08-07 | 16:01 | 19.56 | 68 | 44 | 8.91 | 97.2 | 8.43
2009-08-07 | 16:16 | 19.53 | 68 | 44 | 8.88 | 96.7 | 8.41
2009-08-07 | 16:31 | 19.4 | 68 | 44 | 8.83 | 96 | 8.41
2009-08-07 | 16:46 | 19.3 | 68 | 44 | 8.83 | 95.8 | 8.4
2009-08-07 | 17:01 | 19.19 | 68 | 44 | 8.89 | 96.2 | 8.39
2009-08-07 | 17:16 | 19.07 | 68 | 44 | 8.88 | 95.9 | 8.39
2009-08-07 | 17:31 | 18.92 | 68 | 44 | 8.87 | 95.5 | 8.37
2009-08-07 | 17:46 | 18.79 | 68 | 44 | 8.9 | 95.5 | 8.36
2009-08-07 | 18:01 | 18.65 | 68 | 44 | 8.87 | 95 | 8.34
2009-08-07 | 18:16 | 18.48 | 68 | 44 | 8.9 | 95 | 8.31
2009-08-07 | 18:31 | 18.31 | 67 | 44 | 8.92 | 94.8 | 8.29
2009-08-07 | 18:46 | 18.15 | 67 | 44 | 8.89 | 94.2 | 8.26
2009-08-07 | 19:01 | 17.98 | 67 | 44 | 8.92 | 94.2 | 8.24
2009-08-07 | 19:16 | 17.84 | 67 | 44 | 8.94 | 94.1 | 8.22
2009-08-07 | 19:31 | 17.7 | 67 | 44 | 8.95 | 94 | 8.2
2009-08-07 | 19:46 | 17.57 | 67 | 44 | 8.93 | 93.5 | 8.19
2009-08-07 | 20:01 | 17.42 | 67 | 43 | 8.96 | 93.5 | 8.17
2009-08-07 | 20:16 | 17.28 | 67 | 43 | 8.98 | 93.5 | 8.15
2009-08-07 | 20:31 | 17.14 | 67 | 43 | 8.96 | 93 | 8.14
2009-08-07 | 20:46 | 17.02 | 67 | 43 | 8.96 | 92.8 | 8.13
2009-08-07 | 21:01 | 16.89 | 67 | 43 | 9.01 | 93 | 8.12
2009-08-07 | 21:16 | 16.77 | 67 | 43 | 9.01 | 92.8 | 8.11
2009-08-07 | 21:31 | 16.65 | 66 | 43 | 9.02 | 92.6 | 8.11
2009-08-07 | 21:46 | 16.53 | 66 | 43 | 9.04 | 92.6 | 8.1
2009-08-07 | 22:01 | 16.4 | 66 | 43 | 9.1 | 93 | 8.1
2009-08-07 | 22:16 | 16.28 | 66 | 43 | 9.11 | 92.8 | 8.09
2009-08-07 | 22:31 | 16.15 | 66 | 43 | 9.16 | 93.1 | 8.09
2009-08-07 | 22:46 | 16 | 66 | 43 | 9.21 | 93.3 | 8.09
2009-08-07 | 23:01 | 15.86 | 66 | 43 | 9.2 | 93 | 8.09
2009-08-07 | 23:16 | 15.72 | 66 | 43 | 9.26 | 93.3 | 8.08
2009-08-07 | 23:31 | 15.56 | 66 | 43 | 9.31 | 93.5 | 8.08
2009-08-07 | 23:46 | 15.42 | 66 | 43 | 9.31 | 93.2 | 8.08
2009-08-08 | 0:01 | 15.27 | 66 | 43 | 9.37 | 93.5 | 8.08
2009-08-08 | 0:16 | 15.12 | 66 | 43 | 9.42 | 93.7 | 8.08
2009-08-08 | 0:31 | 14.97 | 66 | 43 | 9.46 | 93.8 | 8.07
2009-08-08 | 0:46 | 14.82 | 66 | 43 | 9.5 | 93.8 | 8.07
2009-08-08 | 1:01 | 14.69 | 66 | 43 | 9.53 | 93.9 | 8.07
2009-08-08 | 1:16 | 14.56 | 66 | 43 | 9.56 | 93.9 | 8.07
2009-08-08 | 1:31 | 14.42 | 66 | 43 | 9.56 | 93.6 | 8.07
2009-08-08 | 1:46 | 14.29 | 66 | 43 | 9.61 | 93.8 | 8.07
2009-08-08 | 2:01 | 14.15 | 66 | 43 | 9.65 | 93.9 | 8.07
2009-08-08 | 2:16 | 14.02 | 66 | 43 | 9.64 | 93.6 | 8.07
2009-08-08 | 2:31 | 13.88 | 66 | 43 | 9.7 | 93.9 | 8.07
2009-08-08 | 2:46 | 13.74 | 66 | 43 | 9.75 | 94 | 8.07
2009-08-08 | 3:01 | 13.61 | 66 | 43 | 9.78 | 94.1 | 8.07
2009-08-08 | 3:16 | 13.47 | 66 | 43 | 9.82 | 94.2 | 8.07
2009-08-08 | 3:31 | 13.34 | 66 | 43 | 9.87 | 94.4 | 8.07
2009-08-08 | 3:46 | 13.21 | 66 | 43 | 9.9 | 94.4 | 8.07
2009-08-08 | 4:01 | 13.09 | 66 | 43 | 9.89 | 94 | 8.07
2009-08-08 | 4:16 | 12.97 | 66 | 43 | 9.94 | 94.3 | 8.07
2009-08-08 | 4:31 | 12.85 | 66 | 43 | 9.99 | 94.5 | 8.06
2009-08-08 | 4:46 | 12.73 | 66 | 43 | 9.98 | 94.1 | 8.07
2009-08-08 | 5:01 | 12.61 | 66 | 43 | 10.03 | 94.4 | 8.07
2009-08-08 | 5:16 | 12.5 | 66 | 43 | 10.08 | 94.6 | 8.07
2009-08-08 | 5:31 | 12.38 | 66 | 43 | 10.11 | 94.6 | 8.07
2009-08-08 | 5:46 | 12.27 | 66 | 43 | 10.09 | 94.3 | 8.07
2009-08-08 | 6:01 | 12.16 | 67 | 43 | 10.11 | 94.2 | 8.07
2009-08-08 | 6:16 | 12.06 | 67 | 43 | 10.13 | 94.1 | 8.06
2009-08-08 | 6:31 | 11.97 | 67 | 43 | 10.16 | 94.3 | 8.07
2009-08-08 | 6:46 | 11.89 | 67 | 44 | 10.2 | 94.4 | 8.09
2009-08-08 | 7:01 | 11.82 | 67 | 44 | 10.28 | 95 | 8.09
2009-08-08 | 7:16 | 11.76 | 67 | 44 | 10.34 | 95.4 | 8.11
2009-08-08 | 7:31 | 11.73 | 67 | 44 | 10.39 | 95.8 | 8.11
2009-08-08 | 7:46 | 11.72 | 68 | 44 | 10.42 | 96 | 8.13
2009-08-08 | 8:01 | 11.72 | 68 | 44 | 10.4 | 95.9 | 8.14
2009-08-08 | 8:16 | 11.75 | 68 | 44 | 10.4 | 96 | 8.15
2009-08-08 | 8:31 | 11.8 | 68 | 44 | 10.44 | 96.4 | 8.16
2009-08-08 | 8:46 | 11.86 | 68 | 44 | 10.42 | 96.4 | 8.18
2009-08-08 | 9:01 | 11.93 | 68 | 44 | 10.46 | 96.9 | 8.19
2009-08-08 | 9:16 | 12 | 69 | 45 | 10.42 | 96.7 | 8.19
2009-08-08 | 9:31 | 12.06 | 69 | 45 | 10.45 | 97.1 | 8.2
2009-08-08 | 9:46 | 12.18 | 69 | 45 | 10.42 | 97.1 | 8.21
2009-08-08 | 10:01 | 12.31 | 69 | 45 | 10.44 | 97.6 | 8.22
2009-08-08 | 10:16 | 12.44 | 69 | 45 | 10.44 | 97.8 | 8.23
2009-08-08 | 10:31 | 12.58 | 69 | 45 | 10.43 | 98.1 | 8.23
2009-08-08 | 10:46 | 12.74 | 70 | 45 | 10.37 | 97.9 | 8.24
2009-08-08 | 11:01 | 12.91 | 70 | 45 | 10.33 | 97.8 | 8.24
2009-08-08 | 11:16 | 13.13 | 70 | 45 | 10.28 | 97.9 | 8.25
2009-08-08 | 11:31 | 13.44 | 70 | 46 | 10.26 | 98.3 | 8.26
2009-08-08 | 11:46 | 13.76 | 70 | 46 | 10.17 | 98.2 | 8.27
2009-08-08 | 12:01 | 14.14 | 70 | 46 | 10.12 | 98.5 | 8.27
2009-08-08 | 12:16 | 14.58 | 71 | 46 | 10.03 | 98.6 | 8.28
2009-08-08 | 12:31 | 15.05 | 71 | 46 | 9.94 | 98.7 | 8.29
2009-08-08 | 12:46 | 15.52 | 71 | 46 | 9.84 | 98.7 | 8.3
2009-08-08 | 13:01 | 16.03 | 71 | 46 | 9.73 | 98.6 | 8.31
2009-08-08 | 13:16 | 16.56 | 71 | 46 | 9.62 | 98.7 | 8.32
2009-08-08 | 13:31 | 17.08 | 71 | 46 | 9.46 | 98.1 | 8.33
2009-08-08 | 13:46 | 17.63 | 71 | 46 | 9.38 | 98.3 | 8.35
2009-08-08 | 14:01 | 18.15 | 71 | 46 | 9.28 | 98.3 | 8.36
2009-08-08 | 14:16 | 18.71 | 71 | 46 | 9.13 | 97.8 | 8.36
2009-08-08 | 14:31 | 19.19 | 71 | 46 | 9.04 | 97.9 | 8.37
2009-08-08 | 14:46 | 19.54 | 71 | 46 | 8.92 | 97.2 | 8.38
2009-08-08 | 15:01 | 19.58 | 71 | 46 | 8.84 | 96.4 | 8.36
2009-08-08 | 15:16 | 19.7 | 71 | 46 | 8.78 | 96 | 8.37
2009-08-08 | 15:31 | 20.23 | 71 | 46 | 8.68 | 95.9 | 8.38
2009-08-08 | 15:46 | 20.61 | 70 | 46 | 8.64 | 96.2 | 8.38
2009-08-08 | 16:01 | 20.73 | 70 | 46 | 8.62 | 96.2 | 8.37
2009-08-08 | 16:16 | 20.71 | 70 | 45 | 8.58 | 95.7 | 8.37
2009-08-08 | 16:31 | 20.69 | 70 | 45 | 8.59 | 95.8 | 8.36
2009-08-08 | 16:46 | 20.83 | 69 | 45 | 8.57 | 95.8 | 8.36
2009-08-08 | 17:01 | 20.92 | 69 | 45 | 8.55 | 95.8 | 8.36
2009-08-08 | 17:16 | 20.94 | 69 | 45 | 8.53 | 95.5 | 8.35
2009-08-08 | 17:31 | 20.82 | 69 | 45 | 8.53 | 95.3 | 8.34
2009-08-08 | 17:46 | 20.68 | 69 | 45 | 8.5 | 94.7 | 8.32
2009-08-08 | 18:01 | 20.53 | 68 | 44 | 8.51 | 94.5 | 8.3
2009-08-08 | 18:16 | 20.39 | 68 | 44 | 8.49 | 94.1 | 8.28
2009-08-08 | 18:31 | 20.24 | 68 | 44 | 8.52 | 94.1 | 8.25
2009-08-08 | 18:46 | 20.12 | 68 | 44 | 8.5 | 93.7 | 8.24
2009-08-08 | 19:01 | 19.94 | 68 | 44 | 8.54 | 93.8 | 8.22
2009-08-08 | 19:16 | 19.78 | 68 | 44 | 8.53 | 93.5 | 8.21
2009-08-08 | 19:31 | 19.66 | 68 | 44 | 8.56 | 93.6 | 8.2
2009-08-08 | 19:46 | 19.53 | 68 | 44 | 8.55 | 93.2 | 8.18
2009-08-08 | 20:01 | 19.4 | 68 | 44 | 8.55 | 93 | 8.17
2009-08-08 | 20:16 | 19.26 | 67 | 44 | 8.59 | 93.1 | 8.15
2009-08-08 | 20:31 | 19.14 | 67 | 44 | 8.61 | 93.1 | 8.14
2009-08-08 | 20:46 | 19.01 | 67 | 44 | 8.6 | 92.7 | 8.13
2009-08-08 | 21:01 | 18.88 | 67 | 44 | 8.6 | 92.5 | 8.12
2009-08-08 | 21:16 | 18.74 | 67 | 44 | 8.65 | 92.8 | 8.11
2009-08-08 | 21:31 | 18.61 | 67 | 44 | 8.69 | 92.9 | 8.11
2009-08-08 | 21:46 | 18.47 | 67 | 44 | 8.72 | 93 | 8.1
2009-08-08 | 22:01 | 18.34 | 67 | 44 | 8.75 | 93.1 | 8.1
2009-08-08 | 22:16 | 18.2 | 67 | 44 | 8.78 | 93.2 | 8.1
2009-08-08 | 22:31 | 18.03 | 67 | 44 | 8.82 | 93.2 | 8.1
2009-08-08 | 22:46 | 17.85 | 67 | 44 | 8.86 | 93.3 | 8.09
2009-08-08 | 23:01 | 17.66 | 67 | 43 | 8.9 | 93.3 | 8.09
2009-08-08 | 23:16 | 17.48 | 67 | 43 | 8.94 | 93.5 | 8.09
2009-08-08 | 23:31 | 17.3 | 67 | 43 | 8.95 | 93.2 | 8.08
2009-08-08 | 23:46 | 17.12 | 67 | 43 | 8.97 | 93.1 | 8.09
2009-08-09 | 0:01 | 16.95 | 67 | 43 | 9.03 | 93.4 | 8.08
2009-08-09 | 0:16 | 16.79 | 67 | 43 | 9.08 | 93.6 | 8.08
2009-08-09 | 0:31 | 16.63 | 67 | 43 | 9.09 | 93.3 | 8.08
2009-08-09 | 0:46 | 16.47 | 67 | 43 | 9.15 | 93.6 | 8.08
2009-08-09 | 1:01 | 16.31 | 67 | 43 | 9.16 | 93.4 | 8.08
2009-08-09 | 1:16 | 16.17 | 67 | 43 | 9.21 | 93.7 | 8.08
2009-08-09 | 1:31 | 16.03 | 67 | 43 | 9.25 | 93.8 | 8.08
2009-08-09 | 1:46 | 15.89 | 67 | 43 | 9.26 | 93.6 | 8.08
2009-08-09 | 2:01 | 15.76 | 67 | 43 | 9.3 | 93.8 | 8.08
2009-08-09 | 2:16 | 15.64 | 67 | 43 | 9.34 | 93.9 | 8.07
2009-08-09 | 2:31 | 15.52 | 67 | 43 | 9.37 | 94 | 8.08
2009-08-09 | 2:46 | 15.42 | 67 | 43 | 9.4 | 94 | 8.08
2009-08-09 | 3:01 | 15.31 | 67 | 43 | 9.44 | 94.2 | 8.08
2009-08-09 | 3:16 | 15.19 | 67 | 43 | 9.42 | 93.8 | 8.07
2009-08-09 | 3:31 | 15.06 | 67 | 44 | 9.43 | 93.7 | 8.07
2009-08-09 | 3:46 | 14.95 | 67 | 44 | 9.45 | 93.7 | 8.07
2009-08-09 | 4:01 | 14.83 | 67 | 44 | 9.48 | 93.7 | 8.07
2009-08-09 | 4:16 | 14.71 | 67 | 44 | 9.5 | 93.7 | 8.07
2009-08-09 | 4:31 | 14.59 | 67 | 44 | 9.56 | 94 | 8.07
2009-08-09 | 4:46 | 14.48 | 67 | 44 | 9.57 | 93.8 | 8.07
2009-08-09 | 5:01 | 14.37 | 67 | 44 | 9.62 | 94.2 | 8.07
2009-08-09 | 5:16 | 14.26 | 68 | 44 | 9.66 | 94.3 | 8.07
2009-08-09 | 5:31 | 14.14 | 68 | 44 | 9.69 | 94.3 | 8.07
2009-08-09 | 5:46 | 14.03 | 68 | 44 | 9.69 | 94.1 | 8.07
2009-08-09 | 6:01 | 13.92 | 68 | 44 | 9.75 | 94.4 | 8.07
2009-08-09 | 6:16 | 13.82 | 68 | 44 | 9.78 | 94.5 | 8.07
2009-08-09 | 6:31 | 13.72 | 68 | 44 | 9.78 | 94.4 | 8.08
2009-08-09 | 6:46 | 13.62 | 68 | 44 | 9.81 | 94.4 | 8.08
2009-08-09 | 7:01 | 13.53 | 68 | 45 | 9.88 | 94.9 | 8.09
2009-08-09 | 7:16 | 13.45 | 68 | 45 | 9.9 | 94.9 | 8.1
2009-08-09 | 7:31 | 13.38 | 69 | 45 | 9.91 | 94.9 | 8.1
2009-08-09 | 7:46 | 13.33 | 69 | 45 | 9.94 | 95 | 8.11
2009-08-09 | 8:01 | 13.3 | 69 | 45 | 9.96 | 95.1 | 8.12
2009-08-09 | 8:16 | 13.28 | 69 | 45 | 9.97 | 95.2 | 8.13
2009-08-09 | 8:31 | 13.31 | 69 | 45 | 10.02 | 95.8 | 8.13
2009-08-09 | 8:46 | 13.35 | 69 | 45 | 10.06 | 96.2 | 8.14
2009-08-09 | 9:01 | 13.41 | 69 | 45 | 10.03 | 96.1 | 8.15
2009-08-09 | 9:16 | 13.44 | 69 | 45 | 10.03 | 96.1 | 8.16
2009-08-09 | 9:31 | 13.49 | 69 | 45 | 10.03 | 96.2 | 8.16
2009-08-09 | 9:46 | 13.56 | 69 | 45 | 10.03 | 96.4 | 8.17
2009-08-09 | 10:01 | 13.66 | 69 | 45 | 10.06 | 96.9 | 8.17
2009-08-09 | 10:16 | 13.78 | 69 | 45 | 10.04 | 96.9 | 8.19
2009-08-09 | 10:31 | 13.92 | 69 | 45 | 10.06 | 97.4 | 8.19
2009-08-09 | 10:46 | 14.08 | 69 | 45 | 10.05 | 97.7 | 8.2
2009-08-09 | 11:01 | 14.23 | 69 | 45 | 10.04 | 97.9 | 8.21
2009-08-09 | 11:16 | 14.4 | 69 | 45 | 10.02 | 98.1 | 8.21
2009-08-09 | 11:31 | 14.69 | 69 | 45 | 9.97 | 98.3 | 8.22
2009-08-09 | 11:46 | 15.06 | 69 | 45 | 9.91 | 98.4 | 8.23
2009-08-09 | 12:01 | 15.45 | 69 | 45 | 9.8 | 98.1 | 8.23
2009-08-09 | 12:16 | 15.81 | 69 | 45 | 9.72 | 98.1 | 8.24
2009-08-09 | 12:31 | 16.27 | 69 | 45 | 9.61 | 97.9 | 8.26
2009-08-09 | 12:46 | 16.82 | 69 | 45 | 9.53 | 98.2 | 8.26
2009-08-09 | 13:01 | 17.36 | 69 | 45 | 9.43 | 98.3 | 8.27
2009-08-09 | 13:16 | 17.84 | 69 | 45 | 9.3 | 98 | 8.29
2009-08-09 | 13:31 | 18.26 | 68 | 45 | 9.23 | 98.1 | 8.29
2009-08-09 | 13:46 | 18.56 | 68 | 44 | 9.13 | 97.6 | 8.3
2009-08-09 | 14:01 | 18.99 | 68 | 44 | 9 | 97 | 8.31
2009-08-09 | 14:16 | 19.7 | 68 | 44 | 8.9 | 97.3 | 8.33
2009-08-09 | 14:31 | 20.42 | 68 | 44 | 8.75 | 97.1 | 8.33
2009-08-09 | 14:46 | 20.8 | 68 | 44 | 8.67 | 96.8 | 8.33
2009-08-09 | 15:01 | 21.19 | 68 | 44 | 8.57 | 96.5 | 8.34
2009-08-09 | 15:16 | 21.52 | 68 | 44 | 8.52 | 96.5 | 8.34
2009-08-09 | 15:31 | 21.86 | 68 | 44 | 8.45 | 96.4 | 8.35
2009-08-09 | 15:46 | 22.1 | 67 | 44 | 8.37 | 95.9 | 8.34
2009-08-09 | 16:01 | 22.22 | 67 | 44 | 8.35 | 95.9 | 8.35
2009-08-09 | 16:16 | 22.2 | 67 | 44 | 8.34 | 95.8 | 8.33
2009-08-09 | 16:31 | 22.13 | 67 | 44 | 8.34 | 95.6 | 8.33
2009-08-09 | 16:46 | 22.01 | 67 | 44 | 8.31 | 95 | 8.32
2009-08-09 | 17:01 | 21.84 | 67 | 43 | 8.3 | 94.6 | 8.29
2009-08-09 | 17:16 | 21.71 | 67 | 43 | 8.29 | 94.3 | 8.28
2009-08-09 | 17:31 | 21.47 | 67 | 43 | 8.31 | 94.1 | 8.25
2009-08-09 | 17:46 | 21.23 | 67 | 43 | 8.33 | 93.9 | 8.23
2009-08-09 | 18:01 | 21.11 | 67 | 43 | 8.36 | 94 | 8.22
2009-08-09 | 18:16 | 20.91 | 66 | 43 | 8.4 | 94 | 8.21
2009-08-09 | 18:31 | 20.73 | 66 | 43 | 8.42 | 94 | 8.19
2009-08-09 | 18:46 | 20.59 | 66 | 43 | 8.4 | 93.5 | 8.18
2009-08-09 | 19:01 | 20.45 | 66 | 43 | 8.41 | 93.3 | 8.17
2009-08-09 | 19:16 | 20.35 | 66 | 43 | 8.41 | 93.1 | 8.16
2009-08-09 | 19:31 | 20.22 | 66 | 43 | 8.44 | 93.2 | 8.14
2009-08-09 | 19:46 | 20.17 | 66 | 43 | 8.42 | 92.9 | 8.14
2009-08-09 | 20:01 | 20.1 | 66 | 43 | 8.42 | 92.8 | 8.13
2009-08-09 | 20:16 | 19.98 | 66 | 43 | 8.43 | 92.6 | 8.12
2009-08-09 | 20:31 | 19.86 | 66 | 43 | 8.46 | 92.7 | 8.1
2009-08-09 | 20:46 | 19.75 | 66 | 43 | 8.48 | 92.8 | 8.1
2009-08-09 | 21:01 | 19.62 | 66 | 43 | 8.48 | 92.5 | 8.09
2009-08-09 | 21:16 | 19.48 | 66 | 43 | 8.52 | 92.7 | 8.08
2009-08-09 | 21:31 | 19.35 | 66 | 43 | 8.53 | 92.6 | 8.08
2009-08-09 | 21:46 | 19.2 | 65 | 42 | 8.54 | 92.5 | 8.08
2009-08-09 | 22:01 | 19.04 | 65 | 42 | 8.6 | 92.8 | 8.07
2009-08-09 | 22:16 | 18.87 | 65 | 42 | 8.62 | 92.6 | 8.07
2009-08-09 | 22:31 | 18.69 | 65 | 42 | 8.68 | 92.9 | 8.06
2009-08-09 | 22:46 | 18.51 | 65 | 42 | 8.69 | 92.7 | 8.06
2009-08-09 | 23:01 | 18.33 | 65 | 42 | 8.75 | 93.1 | 8.06
2009-08-09 | 23:16 | 18.13 | 65 | 42 | 8.8 | 93.2 | 8.06
2009-08-09 | 23:31 | 17.93 | 65 | 42 | 8.85 | 93.4 | 8.06
2009-08-09 | 23:46 | 17.73 | 65 | 42 | 8.89 | 93.4 | 8.06
2009-08-10 | 0:01 | 17.52 | 65 | 42 | 8.94 | 93.5 | 8.06
2009-08-10 | 0:16 | 17.32 | 65 | 42 | 8.98 | 93.5 | 8.05
2009-08-10 | 0:31 | 17.14 | 65 | 42 | 8.99 | 93.3 | 8.05
2009-08-10 | 0:46 | 16.97 | 65 | 42 | 9.04 | 93.5 | 8.05
2009-08-10 | 1:01 | 16.81 | 65 | 42 | 9.09 | 93.7 | 8.05
2009-08-10 | 1:16 | 16.66 | 65 | 42 | 9.12 | 93.7 | 8.05
2009-08-10 | 1:31 | 16.51 | 65 | 42 | 9.12 | 93.4 | 8.05
2009-08-10 | 1:46 | 16.36 | 65 | 42 | 9.15 | 93.4 | 8.05
2009-08-10 | 2:01 | 16.23 | 65 | 42 | 9.2 | 93.7 | 8.05
2009-08-10 | 2:16 | 16.09 | 65 | 42 | 9.24 | 93.8 | 8.05
2009-08-10 | 2:31 | 15.95 | 65 | 42 | 9.28 | 93.9 | 8.05
2009-08-10 | 2:46 | 15.81 | 65 | 42 | 9.28 | 93.7 | 8.05
2009-08-10 | 3:01 | 15.67 | 65 | 42 | 9.34 | 93.9 | 8.05
2009-08-10 | 3:16 | 15.52 | 65 | 42 | 9.34 | 93.7 | 8.05
2009-08-10 | 3:31 | 15.39 | 65 | 42 | 9.4 | 94 | 8.05
2009-08-10 | 3:46 | 15.23 | 65 | 42 | 9.44 | 94.1 | 8.04
2009-08-10 | 4:01 | 15.09 | 65 | 42 | 9.48 | 94.2 | 8.04
2009-08-10 | 4:16 | 14.93 | 65 | 42 | 9.52 | 94.3 | 8.04
2009-08-10 | 4:31 | 14.78 | 65 | 42 | 9.56 | 94.4 | 8.05
2009-08-10 | 4:46 | 14.65 | 65 | 42 | 9.56 | 94.1 | 8.04
2009-08-10 | 5:01 | 14.52 | 65 | 42 | 9.58 | 94 | 8.04
2009-08-10 | 5:16 | 14.39 | 65 | 43 | 9.63 | 94.3 | 8.04
2009-08-10 | 5:31 | 14.27 | 65 | 43 | 9.67 | 94.4 | 8.04
2009-08-10 | 5:46 | 14.14 | 66 | 43 | 9.69 | 94.3 | 8.04
2009-08-10 | 6:01 | 14.01 | 66 | 43 | 9.75 | 94.7 | 8.04
2009-08-10 | 6:16 | 13.87 | 66 | 43 | 9.79 | 94.8 | 8.04
2009-08-10 | 6:31 | 13.74 | 66 | 43 | 9.83 | 94.9 | 8.05
2009-08-10 | 6:46 | 13.62 | 66 | 43 | 9.89 | 95.1 | 8.05
2009-08-10 | 7:01 | 13.48 | 66 | 43 | 9.93 | 95.3 | 8.05
2009-08-10 | 7:16 | 13.37 | 66 | 43 | 9.93 | 95.1 | 8.06
2009-08-10 | 7:31 | 13.26 | 66 | 43 | 9.99 | 95.4 | 8.07
2009-08-10 | 7:46 | 13.16 | 66 | 43 | 10.04 | 95.6 | 8.07
2009-08-10 | 8:01 | 13.1 | 67 | 43 | 10.04 | 95.5 | 8.07
2009-08-10 | 8:16 | 13.06 | 67 | 43 | 10.05 | 95.5 | 8.08
2009-08-10 | 8:31 | 13.06 | 67 | 43 | 10.06 | 95.6 | 8.08
2009-08-10 | 8:46 | 13.07 | 67 | 43 | 10.08 | 95.8 | 8.09
2009-08-10 | 9:01 | 13.1 | 67 | 43 | 10.09 | 96 | 8.1
2009-08-10 | 9:16 | 13.11 | 67 | 43 | 10.14 | 96.5 | 8.1
2009-08-10 | 9:31 | 13.13 | 67 | 44 | 10.13 | 96.4 | 8.11
2009-08-10 | 9:46 | 13.17 | 67 | 43 | 10.17 | 96.9 | 8.12
2009-08-10 | 10:01 | 13.25 | 67 | 43 | 10.18 | 97.2 | 8.13
2009-08-10 | 10:16 | 13.35 | 67 | 43 | 10.18 | 97.4 | 8.13
2009-08-10 | 10:31 | 13.48 | 67 | 43 | 10.17 | 97.6 | 8.13
2009-08-10 | 10:46 | 13.61 | 66 | 43 | 10.17 | 97.8 | 8.14
2009-08-10 | 11:01 | 13.78 | 66 | 43 | 10.15 | 98 | 8.15
2009-08-10 | 11:16 | 14.03 | 66 | 43 | 10.08 | 97.9 | 8.16
2009-08-10 | 11:31 | 14.33 | 66 | 43 | 10.01 | 97.8 | 8.16
2009-08-10 | 11:46 | 14.71 | 66 | 43 | 9.92 | 97.8 | 8.17
2009-08-10 | 12:01 | 15.13 | 66 | 43 | 9.87 | 98.1 | 8.17
2009-08-10 | 12:16 | 15.59 | 65 | 43 | 9.77 | 98.2 | 8.18
2009-08-10 | 12:31 | 16.09 | 65 | 42 | 9.68 | 98.3 | 8.19
2009-08-10 | 12:46 | 16.62 | 65 | 42 | 9.53 | 97.8 | 8.2
2009-08-10 | 13:01 | 17.16 | 65 | 42 | 9.43 | 97.9 | 8.2
2009-08-10 | 13:16 | 17.73 | 65 | 42 | 9.3 | 97.7 | 8.22
2009-08-10 | 13:31 | 18.28 | 64 | 42 | 9.2 | 97.8 | 8.21
2009-08-10 | 13:46 | 18.83 | 64 | 42 | 9.11 | 97.9 | 8.22
2009-08-10 | 14:01 | 19.38 | 64 | 42 | 9.01 | 97.9 | 8.25
2009-08-10 | 14:16 | 19.9 | 64 | 42 | 8.91 | 97.8 | 8.26
2009-08-10 | 14:31 | 20.35 | 64 | 41 | 8.78 | 97.2 | 8.26
2009-08-10 | 14:46 | 20.79 | 64 | 41 | 8.7 | 97.2 | 8.26
2009-08-10 | 15:01 | 21.2 | 64 | 41 | 8.59 | 96.8 | 8.27
2009-08-10 | 15:16 | 21.47 | 63 | 41 | 8.51 | 96.4 | 8.28
2009-08-10 | 15:31 | 21.76 | 63 | 41 | 8.46 | 96.4 | 8.28
2009-08-10 | 15:46 | 21.95 | 63 | 41 | 8.4 | 95.9 | 8.28
2009-08-10 | 16:01 | 22.06 | 63 | 41 | 8.36 | 95.7 | 8.29
2009-08-10 | 16:16 | 22.1 | 63 | 41 | 8.33 | 95.4 | 8.28
2009-08-10 | 16:31 | 22.01 | 63 | 41 | 8.31 | 95.1 | 8.27
2009-08-10 | 16:46 | 21.88 | 63 | 41 | 8.34 | 95.1 | 8.25
2009-08-10 | 17:01 | 21.78 | 63 | 41 | 8.31 | 94.7 | 8.24
2009-08-10 | 17:16 | 21.57 | 63 | 41 | 8.34 | 94.6 | 8.22
2009-08-10 | 17:31 | 21.35 | 63 | 41 | 8.36 | 94.5 | 8.2
2009-08-10 | 17:46 | 21.15 | 63 | 41 | 8.36 | 94.1 | 8.18
2009-08-10 | 18:01 | 20.91 | 63 | 41 | 8.39 | 94 | 8.16
2009-08-10 | 18:16 | 20.74 | 63 | 41 | 8.42 | 94 | 8.16
2009-08-10 | 18:31 | 20.58 | 63 | 41 | 8.43 | 93.8 | 8.15
2009-08-10 | 18:46 | 20.44 | 62 | 41 | 8.45 | 93.8 | 8.13
2009-08-10 | 19:01 | 20.3 | 62 | 41 | 8.46 | 93.7 | 8.12
2009-08-10 | 19:16 | 20.21 | 62 | 41 | 8.44 | 93.2 | 8.11
2009-08-10 | 19:31 | 20.18 | 62 | 41 | 8.43 | 93 | 8.1
2009-08-10 | 19:46 | 20.09 | 62 | 40 | 8.46 | 93.2 | 8.08
2009-08-10 | 20:01 | 20.01 | 62 | 40 | 8.44 | 92.9 | 8.08
2009-08-10 | 20:16 | 19.94 | 62 | 40 | 8.43 | 92.7 | 8.07
2009-08-10 | 20:31 | 19.82 | 62 | 40 | 8.48 | 92.9 | 8.06
2009-08-10 | 20:46 | 19.72 | 62 | 40 | 8.49 | 92.9 | 8.05
2009-08-10 | 21:01 | 19.59 | 62 | 40 | 8.49 | 92.6 | 8.04
2009-08-10 | 21:16 | 19.45 | 62 | 40 | 8.53 | 92.8 | 8.04
2009-08-10 | 21:31 | 19.3 | 62 | 40 | 8.57 | 92.9 | 8.04
2009-08-10 | 21:46 | 19.13 | 62 | 40 | 8.61 | 93.1 | 8.03
2009-08-10 | 22:01 | 18.96 | 62 | 40 | 8.61 | 92.8 | 8.02
2009-08-10 | 22:16 | 18.77 | 62 | 40 | 8.67 | 93 | 8.01
2009-08-10 | 22:31 | 18.58 | 62 | 40 | 8.72 | 93.2 | 8.02
2009-08-10 | 22:46 | 18.39 | 62 | 40 | 8.76 | 93.3 | 8.02
2009-08-10 | 23:01 | 18.22 | 62 | 40 | 8.77 | 93.1 | 8.02
2009-08-10 | 23:16 | 18.03 | 62 | 40 | 8.82 | 93.2 | 8.02
2009-08-10 | 23:31 | 17.84 | 62 | 40 | 8.87 | 93.4 | 8.02
2009-08-10 | 23:46 | 17.65 | 62 | 40 | 8.91 | 93.4 | 8.01
2009-08-11 | 0:01 | 17.48 | 62 | 40 | 8.92 | 93.2 | 8.02
2009-08-11 | 0:16 | 17.31 | 62 | 40 | 8.97 | 93.5 | 8.01
2009-08-11 | 0:31 | 17.16 | 62 | 40 | 8.98 | 93.2 | 8.01
2009-08-11 | 0:46 | 17 | 62 | 40 | 9.04 | 93.5 | 8.01
2009-08-11 | 1:01 | 16.84 | 62 | 40 | 9.08 | 93.7 | 8.01
2009-08-11 | 1:16 | 16.69 | 62 | 40 | 9.12 | 93.8 | 8.02
2009-08-11 | 1:31 | 16.54 | 62 | 40 | 9.16 | 93.9 | 8.01
2009-08-11 | 1:46 | 16.41 | 62 | 40 | 9.2 | 94 | 8.02
2009-08-11 | 2:01 | 16.28 | 62 | 40 | 9.18 | 93.6 | 8.02
2009-08-11 | 2:16 | 16.15 | 62 | 40 | 9.23 | 93.8 | 8.02
2009-08-11 | 2:31 | 16 | 62 | 41 | 9.27 | 94 | 8.02
2009-08-11 | 2:46 | 15.87 | 62 | 41 | 9.27 | 93.7 | 8.02
2009-08-11 | 3:01 | 15.73 | 62 | 41 | 9.32 | 93.9 | 8.02
2009-08-11 | 3:16 | 15.58 | 62 | 41 | 9.34 | 93.8 | 8.01
2009-08-11 | 3:31 | 15.42 | 63 | 41 | 9.36 | 93.7 | 8.02
2009-08-11 | 3:46 | 15.27 | 63 | 41 | 9.42 | 94 | 8.02
2009-08-11 | 4:01 | 15.12 | 63 | 41 | 9.47 | 94.2 | 8.01
2009-08-11 | 4:16 | 14.97 | 63 | 41 | 9.48 | 93.9 | 8.02
2009-08-11 | 4:31 | 14.83 | 63 | 41 | 9.53 | 94.1 | 8.01
2009-08-11 | 4:46 | 14.7 | 63 | 41 | 9.54 | 94 | 8.02
2009-08-11 | 5:01 | 14.56 | 63 | 41 | 9.56 | 93.9 | 8.01
2009-08-11 | 5:16 | 14.43 | 63 | 41 | 9.58 | 93.9 | 8.02
2009-08-11 | 5:31 | 14.31 | 64 | 41 | 9.61 | 93.9 | 8.02
2009-08-11 | 5:46 | 14.19 | 64 | 41 | 9.67 | 94.2 | 8.02
2009-08-11 | 6:01 | 14.07 | 64 | 42 | 9.72 | 94.4 | 8.02
2009-08-11 | 6:16 | 13.95 | 64 | 42 | 9.76 | 94.6 | 8.02
2009-08-11 | 6:31 | 13.84 | 65 | 42 | 9.81 | 94.9 | 8.03
2009-08-11 | 6:46 | 13.73 | 65 | 42 | 9.85 | 95 | 8.03
2009-08-11 | 7:01 | 13.64 | 65 | 42 | 9.89 | 95.3 | 8.04
2009-08-11 | 7:16 | 13.55 | 66 | 43 | 9.94 | 95.5 | 8.05
2009-08-11 | 7:31 | 13.48 | 66 | 43 | 9.94 | 95.4 | 8.06
2009-08-11 | 7:46 | 13.42 | 66 | 43 | 9.99 | 95.7 | 8.07
2009-08-11 | 8:01 | 13.37 | 67 | 43 | 10.04 | 96 | 8.08
2009-08-11 | 8:16 | 13.34 | 67 | 43 | 10.06 | 96.2 | 8.09
2009-08-11 | 8:31 | 13.35 | 67 | 44 | 10.08 | 96.4 | 8.11
2009-08-11 | 8:46 | 13.38 | 67 | 44 | 10.1 | 96.7 | 8.11
2009-08-11 | 9:01 | 13.44 | 68 | 44 | 10.08 | 96.6 | 8.12
2009-08-11 | 9:16 | 13.48 | 68 | 44 | 10.06 | 96.5 | 8.14
2009-08-11 | 9:31 | 13.54 | 68 | 44 | 10.1 | 97 | 8.14
2009-08-11 | 9:46 | 13.62 | 68 | 44 | 10.07 | 97 | 8.15
2009-08-11 | 10:01 | 13.71 | 68 | 44 | 10.09 | 97.3 | 8.16
2009-08-11 | 10:16 | 13.83 | 69 | 45 | 10.07 | 97.3 | 8.16
2009-08-11 | 10:31 | 13.99 | 69 | 45 | 10.09 | 97.8 | 8.18
2009-08-11 | 10:46 | 14.14 | 69 | 45 | 10.09 | 98.2 | 8.19
2009-08-11 | 11:01 | 14.35 | 69 | 45 | 10.08 | 98.6 | 8.2
2009-08-11 | 11:16 | 14.59 | 69 | 45 | 10.02 | 98.5 | 8.21
2009-08-11 | 11:31 | 14.87 | 69 | 45 | 9.95 | 98.4 | 8.22
2009-08-11 | 11:46 | 15.24 | 69 | 45 | 9.88 | 98.4 | 8.24
2009-08-11 | 12:01 | 15.6 | 69 | 45 | 9.83 | 98.8 | 8.25
2009-08-11 | 12:16 | 15.94 | 69 | 45 | 9.77 | 98.8 | 8.27
2009-08-11 | 12:31 | 16.3 | 69 | 45 | 9.69 | 98.8 | 8.29
2009-08-11 | 12:46 | 16.74 | 69 | 45 | 9.61 | 98.9 | 8.29
2009-08-11 | 13:01 | 17.25 | 68 | 44 | 9.52 | 99 | 8.32
2009-08-11 | 13:16 | 17.63 | 68 | 44 | 9.43 | 98.9 | 8.32
2009-08-11 | 13:31 | 18.1 | 68 | 44 | 9.33 | 98.8 | 8.33
2009-08-11 | 13:46 | 18.53 | 68 | 44 | 9.24 | 98.7 | 8.34
2009-08-11 | 14:01 | 18.86 | 68 | 44 | 9.17 | 98.5 | 8.35
2009-08-11 | 14:16 | 19.3 | 67 | 44 | 9.07 | 98.4 | 8.35
2009-08-11 | 14:31 | 19.74 | 67 | 44 | 8.94 | 97.8 | 8.35
2009-08-11 | 14:46 | 20.01 | 67 | 43 | 8.86 | 97.5 | 8.36
2009-08-11 | 15:01 | 20.3 | 67 | 43 | 8.81 | 97.4 | 8.36
2009-08-11 | 15:16 | 20.54 | 66 | 43 | 8.74 | 97.2 | 8.36
2009-08-11 | 15:31 | 20.69 | 66 | 43 | 8.69 | 96.9 | 8.35
2009-08-11 | 15:46 | 20.73 | 66 | 43 | 8.68 | 96.9 | 8.35
2009-08-11 | 16:01 | 20.7 | 66 | 43 | 8.68 | 96.8 | 8.35
2009-08-11 | 16:16 | 20.63 | 66 | 43 | 8.65 | 96.4 | 8.33
2009-08-11 | 16:31 | 20.52 | 66 | 43 | 8.67 | 96.3 | 8.32
2009-08-11 | 16:46 | 20.43 | 66 | 43 | 8.68 | 96.3 | 8.31
2009-08-11 | 17:01 | 20.35 | 65 | 42 | 8.64 | 95.7 | 8.3
2009-08-11 | 17:16 | 20.26 | 65 | 42 | 8.66 | 95.8 | 8.29
2009-08-11 | 17:31 | 20.22 | 65 | 42 | 8.65 | 95.6 | 8.27
2009-08-11 | 17:46 | 20.16 | 65 | 42 | 8.64 | 95.3 | 8.25
2009-08-11 | 18:01 | 20.12 | 65 | 42 | 8.6 | 94.8 | 8.23
2009-08-11 | 18:16 | 20.11 | 65 | 42 | 8.59 | 94.7 | 8.2
2009-08-11 | 18:31 | 20.03 | 65 | 42 | 8.56 | 94.2 | 8.18
2009-08-11 | 18:46 | 19.98 | 65 | 42 | 8.53 | 93.8 | 8.17
2009-08-11 | 19:01 | 19.93 | 65 | 42 | 8.52 | 93.6 | 8.15
2009-08-11 | 19:16 | 19.86 | 65 | 42 | 8.55 | 93.8 | 8.13
2009-08-11 | 19:31 | 19.79 | 65 | 42 | 8.54 | 93.6 | 8.12
2009-08-11 | 19:46 | 19.71 | 65 | 42 | 8.56 | 93.6 | 8.11
2009-08-11 | 20:01 | 19.63 | 65 | 42 | 8.53 | 93.2 | 8.09
2009-08-11 | 20:16 | 19.53 | 65 | 42 | 8.56 | 93.2 | 8.08
2009-08-11 | 20:31 | 19.42 | 65 | 42 | 8.55 | 93 | 8.07
2009-08-11 | 20:46 | 19.34 | 65 | 42 | 8.58 | 93.1 | 8.06
2009-08-11 | 21:01 | 19.21 | 65 | 42 | 8.58 | 92.9 | 8.05
2009-08-11 | 21:16 | 19.08 | 65 | 42 | 8.63 | 93.2 | 8.04
2009-08-11 | 21:31 | 18.95 | 65 | 42 | 8.66 | 93.3 | 8.05
2009-08-11 | 21:46 | 18.83 | 65 | 42 | 8.69 | 93.4 | 8.04
2009-08-11 | 22:01 | 18.71 | 64 | 42 | 8.73 | 93.5 | 8.04
2009-08-11 | 22:16 | 18.61 | 64 | 42 | 8.75 | 93.5 | 8.04
2009-08-11 | 22:31 | 18.52 | 64 | 42 | 8.74 | 93.3 | 8.03
2009-08-11 | 22:46 | 18.42 | 64 | 42 | 8.75 | 93.3 | 8.04
2009-08-11 | 23:01 | 18.34 | 64 | 42 | 8.76 | 93.1 | 8.04
2009-08-11 | 23:16 | 18.24 | 64 | 42 | 8.81 | 93.5 | 8.04
2009-08-11 | 23:31 | 18.16 | 64 | 42 | 8.8 | 93.3 | 8.04
2009-08-11 | 23:46 | 18.09 | 64 | 42 | 8.83 | 93.5 | 8.04
2009-08-12 | 0:01 | 18.03 | 64 | 42 | 8.86 | 93.6 | 8.04
2009-08-12 | 0:16 | 17.97 | 64 | 42 | 8.84 | 93.3 | 8.03
2009-08-12 | 0:31 | 17.91 | 64 | 42 | 8.88 | 93.6 | 8.04
2009-08-12 | 0:46 | 17.86 | 64 | 42 | 8.86 | 93.3 | 8.04
2009-08-12 | 1:01 | 17.81 | 64 | 42 | 8.89 | 93.6 | 8.04
2009-08-12 | 1:16 | 17.77 | 64 | 42 | 8.89 | 93.4 | 8.04
2009-08-12 | 1:31 | 17.72 | 64 | 42 | 8.91 | 93.6 | 8.03
2009-08-12 | 1:46 | 17.66 | 64 | 42 | 8.9 | 93.4 | 8.03
2009-08-12 | 2:01 | 17.61 | 64 | 42 | 8.91 | 93.3 | 8.03
2009-08-12 | 2:16 | 17.57 | 64 | 42 | 8.91 | 93.3 | 8.04
2009-08-12 | 2:31 | 17.53 | 64 | 42 | 8.91 | 93.2 | 8.03
2009-08-12 | 2:46 | 17.49 | 64 | 42 | 8.95 | 93.6 | 8.04
2009-08-12 | 3:01 | 17.44 | 64 | 42 | 8.97 | 93.7 | 8.03
2009-08-12 | 3:16 | 17.38 | 65 | 42 | 8.99 | 93.7 | 8.04
2009-08-12 | 3:31 | 17.34 | 65 | 42 | 9.01 | 93.9 | 8.04
2009-08-12 | 3:46 | 17.31 | 65 | 42 | 9.01 | 93.8 | 8.04
2009-08-12 | 4:01 | 17.27 | 65 | 42 | 8.99 | 93.6 | 8.04
2009-08-12 | 4:16 | 17.23 | 65 | 42 | 8.99 | 93.5 | 8.05
2009-08-12 | 4:31 | 17.18 | 65 | 42 | 9.02 | 93.7 | 8.04
2009-08-12 | 4:46 | 17.13 | 65 | 42 | 9.01 | 93.5 | 8.05
2009-08-12 | 5:01 | 17.09 | 65 | 42 | 9.04 | 93.8 | 8.05
2009-08-12 | 5:16 | 17.06 | 66 | 43 | 9.07 | 94 | 8.05
2009-08-12 | 5:31 | 17.02 | 66 | 43 | 9.07 | 93.9 | 8.04
2009-08-12 | 5:46 | 16.97 | 66 | 43 | 9.09 | 94 | 8.05
2009-08-12 | 6:01 | 16.92 | 66 | 43 | 9.07 | 93.7 | 8.05
2009-08-12 | 6:16 | 16.88 | 67 | 43 | 9.11 | 94 | 8.06
2009-08-12 | 6:31 | 16.83 | 67 | 43 | 9.13 | 94.2 | 8.06
2009-08-12 | 6:46 | 16.77 | 67 | 43 | 9.16 | 94.3 | 8.06
2009-08-12 | 7:01 | 16.73 | 67 | 44 | 9.16 | 94.2 | 8.07
2009-08-12 | 7:16 | 16.68 | 67 | 44 | 9.21 | 94.7 | 8.08
2009-08-12 | 7:31 | 16.64 | 68 | 44 | 9.24 | 94.9 | 8.09
2009-08-12 | 7:46 | 16.61 | 68 | 44 | 9.27 | 95.2 | 8.11
2009-08-12 | 8:01 | 16.58 | 68 | 44 | 9.31 | 95.5 | 8.12
2009-08-12 | 8:16 | 16.58 | 68 | 44 | 9.32 | 95.6 | 8.14
2009-08-12 | 8:31 | 16.56 | 68 | 44 | 9.37 | 96 | 8.16
2009-08-12 | 8:46 | 16.55 | 69 | 45 | 9.4 | 96.3 | 8.17
2009-08-12 | 9:01 | 16.55 | 69 | 45 | 9.41 | 96.5 | 8.18
2009-08-12 | 9:16 | 16.54 | 69 | 45 | 9.39 | 96.3 | 8.2
2009-08-12 | 9:31 | 16.52 | 69 | 45 | 9.41 | 96.4 | 8.2
2009-08-12 | 9:46 | 16.5 | 69 | 45 | 9.4 | 96.3 | 8.2
2009-08-12 | 10:01 | 16.51 | 69 | 45 | 9.43 | 96.6 | 8.19
2009-08-12 | 10:16 | 16.53 | 69 | 45 | 9.46 | 96.9 | 8.21
2009-08-12 | 10:31 | 16.54 | 69 | 45 | 9.48 | 97.2 | 8.22
2009-08-12 | 10:46 | 16.58 | 68 | 44 | 9.45 | 96.9 | 8.23
2009-08-12 | 11:01 | 16.62 | 68 | 44 | 9.45 | 97 | 8.24
2009-08-12 | 11:16 | 16.81 | 68 | 44 | 9.45 | 97.4 | 8.27
2009-08-12 | 11:31 | 17 | 68 | 44 | 9.48 | 98.1 | 8.28
2009-08-12 | 11:46 | 17.07 | 67 | 44 | 9.47 | 98.1 | 8.29
2009-08-12 | 12:01 | 17.36 | 67 | 43 | 9.44 | 98.4 | 8.3
2009-08-12 | 12:16 | 17.67 | 67 | 43 | 9.38 | 98.4 | 8.29
2009-08-12 | 12:31 | 17.91 | 66 | 43 | 9.32 | 98.3 | 8.3
2009-08-12 | 12:46 | 18.36 | 66 | 43 | 9.26 | 98.6 | 8.3
2009-08-12 | 13:01 | 18.76 | 66 | 43 | 9.18 | 98.5 | 8.3
2009-08-12 | 13:16 | 19.11 | 66 | 43 | 9.11 | 98.4 | 8.32
2009-08-12 | 13:31 | 19.5 | 65 | 42 | 9 | 98 | 8.31
2009-08-12 | 13:46 | 19.87 | 65 | 42 | 8.94 | 98.1 | 8.31
2009-08-12 | 14:01 | 20.3 | 65 | 42 | 8.89 | 98.4 | 8.32
2009-08-12 | 14:16 | 20.72 | 65 | 42 | 8.81 | 98.3 | 8.33
2009-08-12 | 14:31 | 21.17 | 65 | 42 | 8.71 | 98 | 8.35
2009-08-12 | 14:46 | 21.58 | 64 | 42 | 8.61 | 97.7 | 8.36
2009-08-12 | 15:01 | 21.86 | 64 | 42 | 8.58 | 97.8 | 8.35
2009-08-12 | 15:16 | 21.92 | 64 | 42 | 8.5 | 97 | 8.37
2009-08-12 | 15:31 | 21.85 | 64 | 42 | 8.47 | 96.6 | 8.37
2009-08-12 | 15:46 | 21.78 | 64 | 42 | 8.45 | 96.3 | 8.36
2009-08-12 | 16:01 | 21.62 | 64 | 41 | 8.49 | 96.3 | 8.35
2009-08-12 | 16:16 | 21.43 | 64 | 41 | 8.49 | 96.1 | 8.34
2009-08-12 | 16:31 | 21.24 | 63 | 41 | 8.53 | 96.1 | 8.33
2009-08-12 | 16:46 | 21.06 | 63 | 41 | 8.52 | 95.6 | 8.32
2009-08-12 | 17:01 | 20.89 | 63 | 41 | 8.54 | 95.6 | 8.29
2009-08-12 | 17:16 | 20.74 | 63 | 41 | 8.49 | 94.7 | 8.25
2009-08-12 | 17:31 | 20.66 | 63 | 41 | 8.47 | 94.4 | 8.22
2009-08-12 | 17:46 | 20.56 | 63 | 41 | 8.47 | 94.2 | 8.21
2009-08-12 | 18:01 | 20.48 | 63 | 41 | 8.52 | 94.6 | 8.2
2009-08-12 | 18:16 | 20.38 | 63 | 41 | 8.5 | 94.2 | 8.18
2009-08-12 | 18:31 | 20.25 | 63 | 41 | 8.5 | 93.9 | 8.17
2009-08-12 | 18:46 | 20.15 | 63 | 41 | 8.5 | 93.7 | 8.15
2009-08-12 | 19:01 | 20.02 | 63 | 41 | 8.53 | 93.8 | 8.13
2009-08-12 | 19:16 | 19.94 | 63 | 41 | 8.51 | 93.5 | 8.11
2009-08-12 | 19:31 | 19.79 | 63 | 41 | 8.55 | 93.7 | 8.1
2009-08-12 | 19:46 | 19.66 | 62 | 41 | 8.57 | 93.7 | 8.09
2009-08-12 | 20:01 | 19.51 | 62 | 41 | 8.6 | 93.6 | 8.08
2009-08-12 | 20:16 | 19.4 | 62 | 41 | 8.61 | 93.6 | 8.07
2009-08-12 | 20:31 | 19.24 | 62 | 40 | 8.65 | 93.7 | 8.06
2009-08-12 | 20:46 | 19.07 | 62 | 40 | 8.67 | 93.6 | 8.05
2009-08-12 | 21:01 | 18.91 | 62 | 40 | 8.68 | 93.4 | 8.05
2009-08-12 | 21:16 | 18.77 | 62 | 40 | 8.69 | 93.3 | 8.04
2009-08-12 | 21:31 | 18.64 | 62 | 40 | 8.75 | 93.6 | 8.04
2009-08-12 | 21:46 | 18.5 | 62 | 40 | 8.79 | 93.8 | 8.04
2009-08-12 | 22:01 | 18.37 | 62 | 40 | 8.78 | 93.5 | 8.03
2009-08-12 | 22:16 | 18.25 | 62 | 40 | 8.83 | 93.7 | 8.03
2009-08-12 | 22:31 | 18.12 | 62 | 40 | 8.86 | 93.8 | 8.03
2009-08-12 | 22:46 | 17.99 | 62 | 40 | 8.89 | 93.9 | 8.03
2009-08-12 | 23:01 | 17.86 | 62 | 40 | 8.92 | 94 | 8.03
2009-08-12 | 23:16 | 17.74 | 62 | 40 | 8.95 | 94.1 | 8.04
2009-08-12 | 23:31 | 17.63 | 62 | 40 | 8.98 | 94.1 | 8.03
2009-08-12 | 23:46 | 17.53 | 62 | 40 | 8.97 | 93.8 | 8.03
2009-08-13 | 0:01 | 17.44 | 62 | 40 | 9.01 | 94.1 | 8.03
2009-08-13 | 0:16 | 17.36 | 62 | 40 | 9 | 93.8 | 8.03
2009-08-13 | 0:31 | 17.26 | 62 | 40 | 9.04 | 94 | 8.03
2009-08-13 | 0:46 | 17.18 | 61 | 40 | 9.06 | 94.1 | 8.03
2009-08-13 | 1:01 | 17.11 | 61 | 40 | 9.09 | 94.2 | 8.03
2009-08-13 | 1:16 | 17.06 | 61 | 40 | 9.06 | 93.9 | 8.02
2009-08-13 | 1:31 | 16.99 | 61 | 40 | 9.08 | 93.9 | 8.02
2009-08-13 | 1:46 | 16.92 | 62 | 40 | 9.08 | 93.8 | 8.03
2009-08-13 | 2:01 | 16.85 | 62 | 40 | 9.12 | 94 | 8.02
2009-08-13 | 2:16 | 16.8 | 62 | 40 | 9.14 | 94.2 | 8.02
2009-08-13 | 2:31 | 16.75 | 62 | 40 | 9.12 | 93.8 | 8.03
2009-08-13 | 2:46 | 16.69 | 62 | 40 | 9.12 | 93.8 | 8.03
2009-08-13 | 3:01 | 16.63 | 62 | 40 | 9.16 | 94 | 8.03
2009-08-13 | 3:16 | 16.59 | 62 | 40 | 9.19 | 94.2 | 8.03
2009-08-13 | 3:31 | 16.53 | 62 | 40 | 9.2 | 94.3 | 8.03
2009-08-13 | 3:46 | 16.46 | 62 | 40 | 9.18 | 94 | 8.03
2009-08-13 | 4:01 | 16.41 | 62 | 41 | 9.22 | 94.2 | 8.03
2009-08-13 | 4:16 | 16.35 | 63 | 41 | 9.2 | 93.9 | 8.03
2009-08-13 | 4:31 | 16.29 | 63 | 41 | 9.25 | 94.3 | 8.03
2009-08-13 | 4:46 | 16.22 | 63 | 41 | 9.23 | 94 | 8.04
2009-08-13 | 5:01 | 16.13 | 63 | 41 | 9.27 | 94.2 | 8.04
2009-08-13 | 5:16 | 16.06 | 63 | 41 | 9.3 | 94.3 | 8.04
2009-08-13 | 5:31 | 15.98 | 63 | 41 | 9.33 | 94.4 | 8.05
2009-08-13 | 5:46 | 15.91 | 63 | 41 | 9.35 | 94.6 | 8.04
2009-08-13 | 6:01 | 15.83 | 64 | 41 | 9.37 | 94.6 | 8.04
2009-08-13 | 6:16 | 15.76 | 64 | 41 | 9.39 | 94.7 | 8.04
2009-08-13 | 6:31 | 15.7 | 64 | 41 | 9.38 | 94.4 | 8.04
2009-08-13 | 6:46 | 15.64 | 64 | 41 | 9.43 | 94.8 | 8.05
2009-08-13 | 7:01 | 15.59 | 64 | 42 | 9.46 | 95.1 | 8.06
2009-08-13 | 7:16 | 15.56 | 64 | 42 | 9.47 | 95.1 | 8.07
2009-08-13 | 7:31 | 15.52 | 64 | 42 | 9.49 | 95.1 | 8.08
2009-08-13 | 7:46 | 15.49 | 64 | 42 | 9.5 | 95.2 | 8.09
2009-08-13 | 8:01 | 15.48 | 64 | 42 | 9.53 | 95.5 | 8.11
2009-08-13 | 8:16 | 15.46 | 64 | 42 | 9.55 | 95.6 | 8.11
2009-08-13 | 8:31 | 15.44 | 64 | 42 | 9.58 | 96 | 8.12
2009-08-13 | 8:46 | 15.41 | 64 | 42 | 9.58 | 95.8 | 8.13
2009-08-13 | 9:01 | 15.38 | 64 | 42 | 9.59 | 95.9 | 8.14
2009-08-13 | 9:16 | 15.38 | 64 | 42 | 9.64 | 96.4 | 8.15
2009-08-13 | 9:31 | 15.4 | 64 | 42 | 9.64 | 96.4 | 8.16
2009-08-13 | 9:46 | 15.47 | 64 | 42 | 9.66 | 96.7 | 8.17
2009-08-13 | 10:01 | 15.54 | 64 | 42 | 9.66 | 96.9 | 8.19
2009-08-13 | 10:16 | 15.5 | 64 | 42 | 9.7 | 97.2 | 8.2
2009-08-13 | 10:31 | 15.55 | 64 | 41 | 9.71 | 97.4 | 8.2
2009-08-13 | 10:46 | 15.68 | 63 | 41 | 9.71 | 97.8 | 8.22
2009-08-13 | 11:01 | 15.72 | 63 | 41 | 9.71 | 97.8 | 8.23
2009-08-13 | 11:16 | 15.79 | 63 | 41 | 9.7 | 97.8 | 8.23
2009-08-13 | 11:31 | 15.84 | 63 | 41 | 9.68 | 97.8 | 8.23
2009-08-13 | 11:46 | 15.98 | 63 | 41 | 9.68 | 98 | 8.23
2009-08-13 | 12:01 | 16.1 | 62 | 40 | 9.66 | 98.1 | 8.23
2009-08-13 | 12:16 | 16.4 | 62 | 40 | 9.59 | 97.9 | 8.25
2009-08-13 | 12:31 | 16.71 | 62 | 40 | 9.51 | 97.8 | 8.26
2009-08-13 | 12:46 | 17 | 62 | 40 | 9.49 | 98.2 | 8.25
2009-08-13 | 13:01 | 17.49 | 62 | 40 | 9.41 | 98.4 | 8.27
2009-08-13 | 13:16 | 17.87 | 61 | 40 | 9.32 | 98.2 | 8.28
2009-08-13 | 13:31 | 18.25 | 61 | 40 | 9.24 | 98.1 | 8.29
2009-08-13 | 13:46 | 18.64 | 61 | 40 | 9.19 | 98.4 | 8.3
2009-08-13 | 14:01 | 18.99 | 61 | 40 | 9.1 | 98.1 | 8.32
2009-08-13 | 14:16 | 19.32 | 61 | 40 | 9.05 | 98.2 | 8.32
2009-08-13 | 14:31 | 19.59 | 61 | 39 | 8.97 | 97.8 | 8.32
2009-08-13 | 14:46 | 19.78 | 61 | 39 | 8.96 | 98.1 | 8.32
2009-08-13 | 15:01 | 19.89 | 61 | 39 | 8.88 | 97.5 | 8.33
2009-08-13 | 15:16 | 20.01 | 61 | 39 | 8.87 | 97.6 | 8.33
2009-08-13 | 15:31 | 20.02 | 61 | 39 | 8.85 | 97.4 | 8.32
2009-08-13 | 15:46 | 19.88 | 60 | 39 | 8.86 | 97.3 | 8.32
2009-08-13 | 16:01 | 19.79 | 60 | 39 | 8.85 | 97 | 8.32
2009-08-13 | 16:16 | 19.57 | 60 | 39 | 8.86 | 96.6 | 8.32
2009-08-13 | 16:31 | 19.43 | 60 | 39 | 8.86 | 96.3 | 8.3
2009-08-13 | 16:46 | 19.37 | 60 | 39 | 8.88 | 96.5 | 8.29
2009-08-13 | 17:01 | 19.3 | 60 | 39 | 8.89 | 96.5 | 8.29
2009-08-13 | 17:16 | 19.28 | 60 | 39 | 8.86 | 96 | 8.27
2009-08-13 | 17:31 | 19.19 | 60 | 39 | 8.85 | 95.8 | 8.26
2009-08-13 | 17:46 | 19.14 | 60 | 39 | 8.88 | 96 | 8.25
2009-08-13 | 18:01 | 19.02 | 60 | 39 | 8.89 | 95.9 | 8.24
2009-08-13 | 18:16 | 18.94 | 60 | 39 | 8.89 | 95.7 | 8.23
2009-08-13 | 18:31 | 18.84 | 60 | 39 | 8.9 | 95.6 | 8.21
2009-08-13 | 18:46 | 18.7 | 60 | 39 | 8.9 | 95.4 | 8.19
2009-08-13 | 19:01 | 18.57 | 60 | 39 | 8.88 | 94.9 | 8.16
2009-08-13 | 19:16 | 18.38 | 60 | 39 | 8.88 | 94.5 | 8.14
2009-08-13 | 19:31 | 18.23 | 60 | 39 | 8.91 | 94.6 | 8.12
2009-08-13 | 19:46 | 18.06 | 60 | 39 | 8.91 | 94.2 | 8.1
2009-08-13 | 20:01 | 17.86 | 59 | 39 | 8.92 | 94 | 8.08
2009-08-13 | 20:16 | 17.67 | 59 | 39 | 8.94 | 93.8 | 8.06
2009-08-13 | 20:31 | 17.49 | 59 | 39 | 9 | 94 | 8.05
2009-08-13 | 20:46 | 17.33 | 59 | 39 | 9 | 93.8 | 8.04
2009-08-13 | 21:01 | 17.19 | 59 | 38 | 9.05 | 94 | 8.03
2009-08-13 | 21:16 | 17.06 | 59 | 38 | 9.08 | 94.1 | 8.02
2009-08-13 | 21:31 | 16.9 | 59 | 38 | 9.13 | 94.3 | 8.02
2009-08-13 | 21:46 | 16.74 | 59 | 38 | 9.16 | 94.3 | 8.01
2009-08-13 | 22:01 | 16.6 | 59 | 38 | 9.16 | 93.9 | 8.01
2009-08-13 | 22:16 | 16.44 | 59 | 38 | 9.21 | 94.2 | 8
2009-08-13 | 22:31 | 16.28 | 59 | 38 | 9.27 | 94.5 | 8.01
2009-08-13 | 22:46 | 16.13 | 59 | 38 | 9.3 | 94.4 | 8
2009-08-13 | 23:01 | 15.98 | 59 | 38 | 9.33 | 94.5 | 8
2009-08-13 | 23:16 | 15.82 | 59 | 38 | 9.33 | 94.2 | 8
2009-08-13 | 23:31 | 15.66 | 59 | 38 | 9.36 | 94.1 | 8
2009-08-13 | 23:46 | 15.51 | 59 | 38 | 9.41 | 94.4 | 8
2009-08-14 | 0:01 | 15.35 | 59 | 38 | 9.47 | 94.6 | 8
2009-08-14 | 0:16 | 15.22 | 59 | 38 | 9.5 | 94.7 | 8
2009-08-14 | 0:31 | 15.1 | 59 | 38 | 9.5 | 94.4 | 8
2009-08-14 | 0:46 | 15 | 59 | 38 | 9.56 | 94.8 | 8
2009-08-14 | 1:01 | 14.88 | 59 | 38 | 9.6 | 94.9 | 8
2009-08-14 | 1:16 | 14.76 | 59 | 38 | 9.62 | 94.9 | 8
2009-08-14 | 1:31 | 14.67 | 59 | 38 | 9.61 | 94.6 | 8
2009-08-14 | 1:46 | 14.56 | 59 | 39 | 9.66 | 94.9 | 8
2009-08-14 | 2:01 | 14.48 | 59 | 39 | 9.69 | 95 | 8
2009-08-14 | 2:16 | 14.4 | 59 | 39 | 9.68 | 94.7 | 8.01
2009-08-14 | 2:31 | 14.33 | 60 | 39 | 9.69 | 94.7 | 8.01
2009-08-14 | 2:46 | 14.25 | 60 | 39 | 9.7 | 94.7 | 8
2009-08-14 | 3:01 | 14.17 | 60 | 39 | 9.72 | 94.6 | 8
2009-08-14 | 3:16 | 14.08 | 60 | 39 | 9.77 | 95 | 8
2009-08-14 | 3:31 | 13.99 | 60 | 39 | 9.77 | 94.8 | 8
2009-08-14 | 3:46 | 13.9 | 60 | 39 | 9.78 | 94.7 | 8.01
2009-08-14 | 4:01 | 13.8 | 60 | 39 | 9.84 | 95.1 | 8.01
2009-08-14 | 4:16 | 13.69 | 61 | 39 | 9.84 | 94.9 | 8.01
2009-08-14 | 4:31 | 13.59 | 61 | 40 | 9.89 | 95.2 | 8.01
2009-08-14 | 4:46 | 13.49 | 61 | 40 | 9.9 | 95 | 8.01
2009-08-14 | 5:01 | 13.41 | 61 | 40 | 9.95 | 95.3 | 8.01
2009-08-14 | 5:16 | 13.33 | 61 | 40 | 9.94 | 95.1 | 8.01
2009-08-14 | 5:31 | 13.23 | 62 | 40 | 9.99 | 95.3 | 8.02
2009-08-14 | 5:46 | 13.12 | 62 | 40 | 10 | 95.2 | 8.02
2009-08-14 | 6:01 | 13.02 | 62 | 40 | 10.05 | 95.4 | 8.01
2009-08-14 | 6:16 | 12.91 | 62 | 40 | 10.09 | 95.5 | 8.02
2009-08-14 | 6:31 | 12.8 | 62 | 41 | 10.13 | 95.7 | 8.02
2009-08-14 | 6:46 | 12.7 | 63 | 41 | 10.17 | 95.9 | 8.03
2009-08-14 | 7:01 | 12.59 | 63 | 41 | 10.18 | 95.8 | 8.04
2009-08-14 | 7:16 | 12.48 | 63 | 41 | 10.25 | 96.2 | 8.04
2009-08-14 | 7:31 | 12.38 | 63 | 41 | 10.26 | 96.1 | 8.05
2009-08-14 | 7:46 | 12.31 | 63 | 41 | 10.33 | 96.5 | 8.06
2009-08-14 | 8:01 | 12.25 | 63 | 41 | 10.34 | 96.5 | 8.07
2009-08-14 | 8:16 | 12.22 | 63 | 41 | 10.36 | 96.6 | 8.08
2009-08-14 | 8:31 | 12.21 | 64 | 41 | 10.42 | 97.1 | 8.09
2009-08-14 | 8:46 | 12.25 | 64 | 41 | 10.45 | 97.6 | 8.11
2009-08-14 | 9:01 | 12.26 | 64 | 41 | 10.48 | 97.8 | 8.11
2009-08-14 | 9:16 | 12.25 | 64 | 42 | 10.51 | 98.1 | 8.12
2009-08-14 | 9:31 | 12.27 | 64 | 42 | 10.52 | 98.2 | 8.13
2009-08-14 | 9:46 | 12.36 | 64 | 42 | 10.5 | 98.2 | 8.15
2009-08-14 | 10:01 | 12.45 | 64 | 42 | 10.51 | 98.5 | 8.16
2009-08-14 | 10:16 | 12.57 | 64 | 41 | 10.48 | 98.5 | 8.17
2009-08-14 | 10:31 | 12.78 | 64 | 41 | 10.49 | 99.1 | 8.18
2009-08-14 | 10:46 | 12.87 | 64 | 41 | 10.44 | 98.9 | 8.18
2009-08-14 | 11:01 | 13.07 | 63 | 41 | 10.41 | 99 | 8.2
2009-08-14 | 11:16 | 13.09 | 63 | 41 | 10.43 | 99.1 | 8.2
2009-08-14 | 11:31 | 13.13 | 63 | 41 | 10.37 | 98.7 | 8.2
2009-08-14 | 11:46 | 13.27 | 63 | 41 | 10.31 | 98.5 | 8.22
2009-08-14 | 12:01 | 13.59 | 63 | 41 | 10.31 | 99.2 | 8.22
2009-08-14 | 12:16 | 13.78 | 63 | 41 | 10.28 | 99.3 | 8.22
2009-08-14 | 12:31 | 13.87 | 63 | 41 | 10.24 | 99.1 | 8.22
2009-08-14 | 12:46 | 14.06 | 63 | 41 | 10.2 | 99.1 | 8.23
2009-08-14 | 13:01 | 14.39 | 63 | 41 | 10.1 | 98.9 | 8.24
2009-08-14 | 13:16 | 14.72 | 63 | 41 | 10.05 | 99.1 | 8.25
2009-08-14 | 13:31 | 14.91 | 62 | 41 | 9.98 | 98.8 | 8.25
2009-08-14 | 13:46 | 15.28 | 62 | 41 | 9.91 | 98.9 | 8.25
2009-08-14 | 14:01 | 15.87 | 62 | 41 | 9.83 | 99.3 | 8.27
2009-08-14 | 14:16 | 16 | 62 | 41 | 9.75 | 98.8 | 8.26
2009-08-14 | 14:31 | 15.91 | 62 | 41 | 9.7 | 98.1 | 8.25
2009-08-14 | 14:46 | 15.99 | 62 | 41 | 9.64 | 97.7 | 8.25
2009-08-14 | 15:01 | 16 | 62 | 41 | 9.65 | 97.8 | 8.24
2009-08-14 | 15:16 | 15.95 | 62 | 40 | 9.66 | 97.8 | 8.23
2009-08-14 | 15:31 | 16.01 | 62 | 40 | 9.64 | 97.7 | 8.24
2009-08-14 | 15:46 | 16.09 | 62 | 40 | 9.64 | 97.8 | 8.23
2009-08-14 | 16:01 | 16.27 | 62 | 40 | 9.64 | 98.2 | 8.24
2009-08-14 | 16:16 | 16.48 | 62 | 40 | 9.58 | 98 | 8.25
2009-08-14 | 16:31 | 16.62 | 62 | 40 | 9.55 | 98 | 8.25
2009-08-14 | 16:46 | 16.7 | 62 | 40 | 9.49 | 97.5 | 8.25
2009-08-14 | 17:01 | 16.61 | 62 | 40 | 9.45 | 97 | 8.24
2009-08-14 | 17:16 | 16.42 | 62 | 40 | 9.45 | 96.6 | 8.22
2009-08-14 | 17:31 | 16.25 | 62 | 40 | 9.47 | 96.5 | 8.21
2009-08-14 | 17:46 | 16.09 | 62 | 40 | 9.52 | 96.7 | 8.19
2009-08-14 | 18:01 | 16 | 62 | 40 | 9.53 | 96.5 | 8.17
2009-08-14 | 18:16 | 15.91 | 62 | 40 | 9.53 | 96.4 | 8.16
2009-08-14 | 18:31 | 15.82 | 62 | 40 | 9.5 | 95.9 | 8.14
2009-08-14 | 18:46 | 15.76 | 62 | 40 | 9.49 | 95.7 | 8.13
2009-08-14 | 19:01 | 15.71 | 62 | 40 | 9.51 | 95.7 | 8.12
2009-08-14 | 19:16 | 15.71 | 62 | 40 | 9.52 | 95.8 | 8.11
2009-08-14 | 19:31 | 15.71 | 62 | 40 | 9.47 | 95.4 | 8.1
2009-08-14 | 19:46 | 15.66 | 62 | 40 | 9.46 | 95.2 | 8.1
2009-08-14 | 20:01 | 15.58 | 62 | 40 | 9.46 | 95 | 8.09
2009-08-14 | 20:16 | 15.48 | 62 | 40 | 9.46 | 94.8 | 8.07
2009-08-14 | 20:31 | 15.37 | 62 | 40 | 9.47 | 94.7 | 8.06
2009-08-14 | 20:46 | 15.26 | 62 | 40 | 9.52 | 94.9 | 8.05
2009-08-14 | 21:01 | 15.13 | 62 | 40 | 9.56 | 95.1 | 8.04
2009-08-14 | 21:16 | 15.01 | 62 | 40 | 9.59 | 95.2 | 8.04
2009-08-14 | 21:31 | 14.9 | 62 | 40 | 9.62 | 95.2 | 8.03
2009-08-14 | 21:46 | 14.79 | 62 | 40 | 9.61 | 94.9 | 8.03
2009-08-14 | 22:01 | 14.66 | 62 | 40 | 9.66 | 95.1 | 8.03
2009-08-14 | 22:16 | 14.52 | 62 | 40 | 9.71 | 95.3 | 8.03
2009-08-14 | 22:31 | 14.38 | 62 | 40 | 9.74 | 95.3 | 8.02
2009-08-14 | 22:46 | 14.23 | 62 | 40 | 9.75 | 95.1 | 8.02
2009-08-14 | 23:01 | 14.11 | 62 | 40 | 9.77 | 95 | 8.03
2009-08-14 | 23:16 | 13.98 | 62 | 40 | 9.82 | 95.3 | 8.02
2009-08-14 | 23:31 | 13.86 | 62 | 40 | 9.87 | 95.5 | 8.02
2009-08-14 | 23:46 | 13.73 | 62 | 40 | 9.86 | 95.1 | 8.01
2009-08-15 | 0:01 | 13.61 | 62 | 40 | 9.92 | 95.4 | 8.02
2009-08-15 | 0:16 | 13.48 | 62 | 40 | 9.92 | 95.2 | 8.01
2009-08-15 | 0:31 | 13.36 | 62 | 40 | 9.98 | 95.5 | 8.01
2009-08-15 | 0:46 | 13.23 | 62 | 40 | 9.99 | 95.3 | 8.02
2009-08-15 | 1:01 | 13.09 | 62 | 40 | 10.01 | 95.2 | 8.01
2009-08-15 | 1:16 | 12.96 | 62 | 40 | 10.04 | 95.2 | 8.01
2009-08-15 | 1:31 | 12.82 | 62 | 40 | 10.1 | 95.5 | 8.01
2009-08-15 | 1:46 | 12.7 | 62 | 40 | 10.11 | 95.3 | 8.01
2009-08-15 | 2:01 | 12.57 | 62 | 40 | 10.18 | 95.7 | 8.02
2009-08-15 | 2:16 | 12.43 | 62 | 40 | 10.22 | 95.7 | 8.01
2009-08-15 | 2:31 | 12.3 | 62 | 40 | 10.23 | 95.5 | 8.01
2009-08-15 | 2:46 | 12.17 | 62 | 40 | 10.28 | 95.7 | 8.01
2009-08-15 | 3:01 | 12.04 | 62 | 40 | 10.33 | 95.9 | 8.01
2009-08-15 | 3:16 | 11.92 | 62 | 40 | 10.37 | 96 | 8.01
2009-08-15 | 3:31 | 11.8 | 62 | 41 | 10.36 | 95.7 | 8.01
2009-08-15 | 3:46 | 11.68 | 63 | 41 | 10.42 | 95.9 | 8.01
2009-08-15 | 4:01 | 11.56 | 63 | 41 | 10.42 | 95.8 | 8.02
2009-08-15 | 4:16 | 11.45 | 63 | 41 | 10.45 | 95.7 | 8.02
2009-08-15 | 4:31 | 11.34 | 63 | 41 | 10.51 | 96 | 8.01
2009-08-15 | 4:46 | 11.23 | 63 | 41 | 10.55 | 96.2 | 8.01
2009-08-15 | 5:01 | 11.13 | 63 | 41 | 10.54 | 95.9 | 8.02
2009-08-15 | 5:16 | 11.02 | 64 | 41 | 10.56 | 95.8 | 8.02
2009-08-15 | 5:31 | 10.92 | 64 | 41 | 10.62 | 96.2 | 8.02
2009-08-15 | 5:46 | 10.81 | 64 | 42 | 10.63 | 96 | 8.02
2009-08-15 | 6:01 | 10.72 | 64 | 42 | 10.68 | 96.2 | 8.02
2009-08-15 | 6:16 | 10.62 | 64 | 42 | 10.69 | 96.1 | 8.02
2009-08-15 | 6:31 | 10.51 | 65 | 42 | 10.73 | 96.2 | 8.03
2009-08-15 | 6:46 | 10.41 | 65 | 42 | 10.79 | 96.5 | 8.03
2009-08-15 | 7:01 | 10.32 | 65 | 42 | 10.85 | 96.8 | 8.03
2009-08-15 | 7:16 | 10.25 | 65 | 42 | 10.85 | 96.7 | 8.03
2009-08-15 | 7:31 | 10.17 | 65 | 43 | 10.87 | 96.7 | 8.05
2009-08-15 | 7:46 | 10.1 | 66 | 43 | 10.94 | 97.2 | 8.05
2009-08-15 | 8:01 | 10.06 | 66 | 43 | 11 | 97.5 | 8.06
2009-08-15 | 8:16 | 10.04 | 66 | 43 | 10.98 | 97.4 | 8.06
2009-08-15 | 8:31 | 10.05 | 66 | 43 | 10.97 | 97.3 | 8.08
2009-08-15 | 8:46 | 10.08 | 66 | 43 | 10.98 | 97.5 | 8.08
2009-08-15 | 9:01 | 10.14 | 66 | 43 | 10.99 | 97.7 | 8.09
2009-08-15 | 9:16 | 10.17 | 66 | 43 | 11.04 | 98.2 | 8.1
2009-08-15 | 9:31 | 10.24 | 66 | 43 | 11.02 | 98.2 | 8.11
2009-08-15 | 9:46 | 10.32 | 66 | 43 | 11.05 | 98.6 | 8.11
2009-08-15 | 10:01 | 10.42 | 66 | 43 | 11.01 | 98.5 | 8.13
2009-08-15 | 10:16 | 10.54 | 66 | 43 | 11.04 | 99.1 | 8.13
2009-08-15 | 10:31 | 10.68 | 67 | 43 | 11.03 | 99.3 | 8.13
2009-08-15 | 10:46 | 10.82 | 67 | 43 | 10.98 | 99.2 | 8.14
2009-08-15 | 11:01 | 10.99 | 67 | 43 | 10.98 | 99.6 | 8.14
2009-08-15 | 11:16 | 11.2 | 67 | 43 | 10.96 | 99.9 | 8.15
2009-08-15 | 11:31 | 11.44 | 67 | 43 | 10.92 | 100 | 8.16
2009-08-15 | 11:46 | 11.75 | 67 | 43 | 10.81 | 99.7 | 8.17
2009-08-15 | 12:01 | 12.13 | 67 | 43 | 10.74 | 100 | 8.16
2009-08-15 | 12:16 | 12.58 | 67 | 43 | 10.59 | 99.6 | 8.17
2009-08-15 | 12:31 | 13.05 | 67 | 43 | 10.49 | 99.7 | 8.18
2009-08-15 | 12:46 | 13.54 | 67 | 43 | 10.35 | 99.4 | 8.2
2009-08-15 | 13:01 | 14.05 | 67 | 43 | 10.24 | 99.5 | 8.2
2009-08-15 | 13:16 | 14.61 | 67 | 43 | 10.11 | 99.5 | 8.2
2009-08-15 | 13:31 | 15.14 | 67 | 43 | 9.99 | 99.4 | 8.22
2009-08-15 | 13:46 | 15.71 | 67 | 43 | 9.87 | 99.3 | 8.24
2009-08-15 | 14:01 | 16.24 | 67 | 43 | 9.71 | 98.9 | 8.24
2009-08-15 | 14:16 | 16.8 | 67 | 43 | 9.6 | 98.9 | 8.25
2009-08-15 | 14:31 | 17.32 | 66 | 43 | 9.45 | 98.4 | 8.26
2009-08-15 | 14:46 | 17.78 | 66 | 43 | 9.36 | 98.4 | 8.27
2009-08-15 | 15:01 | 18.18 | 66 | 43 | 9.27 | 98.3 | 8.27
2009-08-15 | 15:16 | 18.54 | 66 | 43 | 9.18 | 98.1 | 8.28
2009-08-15 | 15:31 | 18.85 | 66 | 43 | 9.07 | 97.5 | 8.29
2009-08-15 | 15:46 | 19.1 | 66 | 43 | 9.02 | 97.5 | 8.28
2009-08-15 | 16:01 | 19.25 | 66 | 43 | 9 | 97.5 | 8.29
2009-08-15 | 16:16 | 19.34 | 66 | 43 | 8.94 | 97 | 8.3
2009-08-15 | 16:31 | 19.33 | 66 | 43 | 8.91 | 96.7 | 8.29
2009-08-15 | 16:46 | 19.24 | 66 | 43 | 8.94 | 96.8 | 8.29
2009-08-15 | 17:01 | 19.11 | 66 | 43 | 8.92 | 96.4 | 8.28
2009-08-15 | 17:16 | 18.93 | 66 | 43 | 8.95 | 96.4 | 8.27
2009-08-15 | 17:31 | 18.75 | 66 | 43 | 8.98 | 96.3 | 8.25
2009-08-15 | 17:46 | 18.59 | 66 | 43 | 8.97 | 95.9 | 8.23
2009-08-15 | 18:01 | 18.44 | 66 | 43 | 8.97 | 95.6 | 8.22
2009-08-15 | 18:16 | 18.32 | 66 | 43 | 8.95 | 95.2 | 8.22
2009-08-15 | 18:31 | 18.19 | 66 | 43 | 8.98 | 95.3 | 8.18
2009-08-15 | 18:46 | 18.08 | 66 | 43 | 9 | 95.2 | 8.16
2009-08-15 | 19:01 | 17.98 | 66 | 43 | 9 | 95.1 | 8.16
2009-08-15 | 19:16 | 17.89 | 66 | 43 | 9.01 | 95 | 8.14
2009-08-15 | 19:31 | 17.84 | 65 | 43 | 9.01 | 94.9 | 8.13
2009-08-15 | 19:46 | 17.8 | 66 | 43 | 8.98 | 94.4 | 8.12
2009-08-15 | 20:01 | 17.74 | 65 | 43 | 8.97 | 94.2 | 8.11
2009-08-15 | 20:16 | 17.67 | 65 | 43 | 8.96 | 94 | 8.1
2009-08-15 | 20:31 | 17.57 | 65 | 43 | 8.96 | 93.8 | 8.08
2009-08-15 | 20:46 | 17.45 | 65 | 43 | 9 | 94 | 8.08
2009-08-15 | 21:01 | 17.33 | 66 | 43 | 9.04 | 94.1 | 8.07
2009-08-15 | 21:16 | 17.16 | 65 | 43 | 9.05 | 93.9 | 8.06
2009-08-15 | 21:31 | 16.99 | 66 | 43 | 9.1 | 94.1 | 8.06
2009-08-15 | 21:46 | 16.81 | 65 | 43 | 9.12 | 94 | 8.05
2009-08-15 | 22:01 | 16.62 | 65 | 43 | 9.18 | 94.2 | 8.06
2009-08-15 | 22:16 | 16.43 | 65 | 43 | 9.23 | 94.4 | 8.05
2009-08-15 | 22:31 | 16.23 | 65 | 42 | 9.29 | 94.6 | 8.05
2009-08-15 | 22:46 | 16.03 | 65 | 42 | 9.33 | 94.6 | 8.05
2009-08-15 | 23:01 | 15.83 | 65 | 42 | 9.35 | 94.4 | 8.04
2009-08-15 | 23:16 | 15.64 | 65 | 42 | 9.42 | 94.7 | 8.04
2009-08-15 | 23:31 | 15.46 | 65 | 42 | 9.46 | 94.8 | 8.05
2009-08-15 | 23:46 | 15.27 | 65 | 42 | 9.48 | 94.6 | 8.05
2009-08-16 | 0:01 | 15.08 | 65 | 42 | 9.55 | 94.9 | 8.04
2009-08-16 | 0:16 | 14.9 | 65 | 42 | 9.6 | 95 | 8.04
2009-08-16 | 0:31 | 14.72 | 65 | 42 | 9.66 | 95.2 | 8.05
2009-08-16 | 0:46 | 14.54 | 65 | 42 | 9.66 | 94.9 | 8.04
2009-08-16 | 1:01 | 14.38 | 65 | 42 | 9.68 | 94.7 | 8.05
2009-08-16 | 1:16 | 14.21 | 65 | 42 | 9.76 | 95.1 | 8.04
2009-08-16 | 1:31 | 14.05 | 65 | 42 | 9.77 | 95 | 8.05
2009-08-16 | 1:46 | 13.9 | 66 | 43 | 9.84 | 95.3 | 8.04
2009-08-16 | 2:01 | 13.75 | 65 | 42 | 9.89 | 95.4 | 8.04
2009-08-16 | 2:16 | 13.6 | 65 | 43 | 9.89 | 95.2 | 8.04
2009-08-16 | 2:31 | 13.45 | 65 | 43 | 9.95 | 95.4 | 8.04
2009-08-16 | 2:46 | 13.32 | 66 | 43 | 9.99 | 95.5 | 8.04
2009-08-16 | 3:01 | 13.19 | 66 | 43 | 10 | 95.3 | 8.04
2009-08-16 | 3:16 | 13.06 | 66 | 43 | 10.02 | 95.2 | 8.05
2009-08-16 | 3:31 | 12.94 | 66 | 43 | 10.07 | 95.5 | 8.04
2009-08-16 | 3:46 | 12.82 | 66 | 43 | 10.08 | 95.3 | 8.04
2009-08-16 | 4:01 | 12.7 | 66 | 43 | 10.15 | 95.7 | 8.04
2009-08-16 | 4:16 | 12.58 | 66 | 43 | 10.18 | 95.8 | 8.04
2009-08-16 | 4:31 | 12.46 | 66 | 43 | 10.23 | 95.9 | 8.04
2009-08-16 | 4:46 | 12.33 | 66 | 43 | 10.26 | 96 | 8.04
2009-08-16 | 5:01 | 12.21 | 66 | 43 | 10.26 | 95.6 | 8.04
2009-08-16 | 5:16 | 12.1 | 66 | 43 | 10.32 | 95.9 | 8.04
2009-08-16 | 5:31 | 11.98 | 66 | 43 | 10.35 | 96 | 8.04
2009-08-16 | 5:46 | 11.85 | 67 | 43 | 10.36 | 95.8 | 8.04
2009-08-16 | 6:01 | 11.73 | 67 | 43 | 10.37 | 95.7 | 8.04
2009-08-16 | 6:16 | 11.62 | 67 | 44 | 10.41 | 95.7 | 8.04
2009-08-16 | 6:31 | 11.51 | 67 | 44 | 10.47 | 96.1 | 8.04
2009-08-16 | 6:46 | 11.41 | 67 | 44 | 10.49 | 96 | 8.05
2009-08-16 | 7:01 | 11.31 | 68 | 44 | 10.52 | 96.1 | 8.06
2009-08-16 | 7:16 | 11.22 | 68 | 44 | 10.6 | 96.6 | 8.06
2009-08-16 | 7:31 | 11.14 | 68 | 44 | 10.65 | 96.9 | 8.07
2009-08-16 | 7:46 | 11.08 | 68 | 44 | 10.64 | 96.7 | 8.07
2009-08-16 | 8:01 | 11.05 | 68 | 44 | 10.69 | 97.1 | 8.08
2009-08-16 | 8:16 | 11.03 | 68 | 44 | 10.73 | 97.4 | 8.09
2009-08-16 | 8:31 | 11.06 | 68 | 44 | 10.74 | 97.6 | 8.09
2009-08-16 | 8:46 | 11.12 | 68 | 44 | 10.76 | 97.8 | 8.1
2009-08-16 | 9:01 | 11.21 | 69 | 45 | 10.76 | 98.1 | 8.11
2009-08-16 | 9:16 | 11.26 | 69 | 45 | 10.77 | 98.3 | 8.12
2009-08-16 | 9:31 | 11.33 | 69 | 45 | 10.77 | 98.5 | 8.12
2009-08-16 | 9:46 | 11.42 | 69 | 45 | 10.73 | 98.2 | 8.13
2009-08-16 | 10:01 | 11.52 | 69 | 45 | 10.75 | 98.7 | 8.14
2009-08-16 | 10:16 | 11.65 | 69 | 45 | 10.7 | 98.5 | 8.15
2009-08-16 | 10:31 | 11.79 | 69 | 45 | 10.72 | 99 | 8.16
2009-08-16 | 10:46 | 11.93 | 69 | 45 | 10.72 | 99.3 | 8.16
2009-08-16 | 11:01 | 12.09 | 69 | 45 | 10.66 | 99.1 | 8.16
2009-08-16 | 11:16 | 12.28 | 70 | 45 | 10.61 | 99.1 | 8.18
2009-08-16 | 11:31 | 12.53 | 70 | 45 | 10.55 | 99.1 | 8.18
2009-08-16 | 11:46 | 12.83 | 69 | 45 | 10.48 | 99.1 | 8.19
2009-08-16 | 12:01 | 13.21 | 69 | 45 | 10.44 | 99.5 | 8.2
2009-08-16 | 12:16 | 13.64 | 69 | 45 | 10.35 | 99.6 | 8.2
2009-08-16 | 12:31 | 14.12 | 70 | 46 | 10.19 | 99.1 | 8.21
2009-08-16 | 12:46 | 14.61 | 70 | 46 | 10.06 | 98.9 | 8.22
2009-08-16 | 13:01 | 15.15 | 69 | 45 | 9.93 | 98.8 | 8.23
2009-08-16 | 13:16 | 15.68 | 69 | 45 | 9.83 | 99 | 8.23
2009-08-16 | 13:31 | 16.23 | 69 | 45 | 9.72 | 99 | 8.26
2009-08-16 | 13:46 | 16.8 | 69 | 45 | 9.58 | 98.7 | 8.26
2009-08-16 | 14:01 | 17.37 | 69 | 45 | 9.47 | 98.7 | 8.27
2009-08-16 | 14:16 | 17.93 | 69 | 45 | 9.32 | 98.3 | 8.28
2009-08-16 | 14:31 | 18.46 | 69 | 45 | 9.23 | 98.4 | 8.3
2009-08-16 | 14:46 | 18.94 | 69 | 45 | 9.13 | 98.3 | 8.3
2009-08-16 | 15:01 | 19.33 | 69 | 45 | 9.03 | 98 | 8.31
2009-08-16 | 15:16 | 19.68 | 69 | 45 | 8.92 | 97.5 | 8.32
2009-08-16 | 15:31 | 19.97 | 69 | 45 | 8.87 | 97.5 | 8.32
2009-08-16 | 15:46 | 20.23 | 69 | 45 | 8.8 | 97.3 | 8.32
2009-08-16 | 16:01 | 20.4 | 68 | 45 | 8.76 | 97.1 | 8.33
2009-08-16 | 16:16 | 20.49 | 68 | 44 | 8.7 | 96.6 | 8.31
2009-08-16 | 16:31 | 20.48 | 68 | 44 | 8.7 | 96.6 | 8.31
2009-08-16 | 16:46 | 20.38 | 68 | 44 | 8.69 | 96.3 | 8.3
2009-08-16 | 17:01 | 20.26 | 68 | 44 | 8.67 | 95.8 | 8.28
2009-08-16 | 17:16 | 20.04 | 68 | 44 | 8.69 | 95.6 | 8.26
2009-08-16 | 17:31 | 19.82 | 68 | 44 | 8.71 | 95.5 | 8.24
2009-08-16 | 17:46 | 19.62 | 68 | 44 | 8.7 | 94.9 | 8.23
2009-08-16 | 18:01 | 19.44 | 68 | 44 | 8.71 | 94.8 | 8.21
2009-08-16 | 18:16 | 19.25 | 68 | 44 | 8.72 | 94.5 | 8.19
2009-08-16 | 18:31 | 19.06 | 67 | 44 | 8.74 | 94.4 | 8.18
2009-08-16 | 18:46 | 18.9 | 67 | 44 | 8.79 | 94.6 | 8.16
2009-08-16 | 19:01 | 18.75 | 67 | 44 | 8.83 | 94.7 | 8.15
2009-08-16 | 19:16 | 18.65 | 67 | 44 | 8.84 | 94.6 | 8.14
2009-08-16 | 19:31 | 18.6 | 67 | 44 | 8.82 | 94.3 | 8.13
2009-08-16 | 19:46 | 18.52 | 67 | 44 | 8.83 | 94.3 | 8.12
2009-08-16 | 20:01 | 18.45 | 67 | 44 | 8.84 | 94.2 | 8.11
2009-08-16 | 20:16 | 18.39 | 67 | 44 | 8.84 | 94.2 | 8.1
2009-08-16 | 20:31 | 18.34 | 67 | 44 | 8.83 | 93.9 | 8.09
2009-08-16 | 20:46 | 18.25 | 67 | 44 | 8.82 | 93.7 | 8.09
2009-08-16 | 21:01 | 18.14 | 67 | 44 | 8.87 | 93.9 | 8.08
2009-08-16 | 21:16 | 18.01 | 67 | 44 | 8.9 | 94 | 8.07
2009-08-16 | 21:31 | 17.86 | 67 | 44 | 8.93 | 94.1 | 8.06
2009-08-16 | 21:46 | 17.71 | 67 | 44 | 8.98 | 94.3 | 8.07
2009-08-16 | 22:01 | 17.53 | 67 | 44 | 9.02 | 94.4 | 8.06
2009-08-16 | 22:16 | 17.33 | 67 | 44 | 9.05 | 94.3 | 8.04
2009-08-16 | 22:31 | 17.11 | 67 | 44 | 9.12 | 94.6 | 8.05
2009-08-16 | 22:46 | 16.88 | 67 | 44 | 9.13 | 94.2 | 8.06
2009-08-16 | 23:01 | 16.64 | 67 | 44 | 9.2 | 94.5 | 8.05
2009-08-16 | 23:16 | 16.41 | 67 | 44 | 9.27 | 94.7 | 8.06
2009-08-16 | 23:31 | 16.18 | 67 | 43 | 9.33 | 94.9 | 8.05
2009-08-16 | 23:46 | 15.97 | 67 | 43 | 9.38 | 95 | 8.05
2009-08-17 | 0:01 | 15.75 | 67 | 43 | 9.43 | 95.1 | 8.05
2009-08-17 | 0:16 | 15.55 | 67 | 43 | 9.5 | 95.3 | 8.05
2009-08-17 | 0:31 | 15.33 | 67 | 43 | 9.54 | 95.2 | 8.05
2009-08-17 | 0:46 | 15.14 | 67 | 43 | 9.58 | 95.3 | 8.04
2009-08-17 | 1:01 | 14.95 | 67 | 43 | 9.63 | 95.4 | 8.04
2009-08-17 | 1:16 | 14.75 | 67 | 43 | 9.65 | 95.2 | 8.04
2009-08-17 | 1:31 | 14.57 | 67 | 43 | 9.71 | 95.4 | 8.05
2009-08-17 | 1:46 | 14.41 | 66 | 43 | 9.74 | 95.3 | 8.04
2009-08-17 | 2:01 | 14.23 | 67 | 43 | 9.75 | 95.1 | 8.04
2009-08-17 | 2:16 | 14.07 | 66 | 43 | 9.8 | 95.2 | 8.04
2009-08-17 | 2:31 | 13.93 | 67 | 43 | 9.83 | 95.2 | 8.04
2009-08-17 | 2:46 | 13.78 | 67 | 43 | 9.86 | 95.2 | 8.05
2009-08-17 | 3:01 | 13.65 | 67 | 43 | 9.89 | 95.2 | 8.04
2009-08-17 | 3:16 | 13.51 | 67 | 43 | 9.96 | 95.6 | 8.04
2009-08-17 | 3:31 | 13.37 | 67 | 43 | 10.02 | 95.9 | 8.04
2009-08-17 | 3:46 | 13.25 | 67 | 43 | 10.05 | 95.9 | 8.04
2009-08-17 | 4:01 | 13.12 | 69 | 45 | 10.06 | 95.8 | 8.04
2009-08-17 | 4:16 | 13 | 70 | 45 | 10.11 | 96 | 8.04
2009-08-17 | 4:31 | 12.89 | 70 | 45 | 10.14 | 96 | 8.04
2009-08-17 | 4:46 | 12.77 | 70 | 45 | 10.19 | 96.2 | 8.04
2009-08-17 | 5:01 | 12.65 | 68 | 44 | 10.22 | 96.2 | 8.04
2009-08-17 | 5:16 | 12.54 | 67 | 44 | 10.21 | 95.9 | 8.04
2009-08-17 | 5:31 | 12.43 | 68 | 44 | 10.23 | 95.9 | 8.05
2009-08-17 | 5:46 | 12.32 | 68 | 44 | 10.29 | 96.2 | 8.05
2009-08-17 | 6:01 | 12.21 | 68 | 44 | 10.35 | 96.5 | 8.04
2009-08-17 | 6:16 | 12.11 | 68 | 44 | 10.34 | 96.2 | 8.05
2009-08-17 | 6:31 | 12.02 | 68 | 44 | 10.37 | 96.2 | 8.05
2009-08-17 | 6:46 | 11.93 | 68 | 44 | 10.43 | 96.6 | 8.06
2009-08-17 | 7:01 | 11.84 | 68 | 44 | 10.45 | 96.6 | 8.06
2009-08-17 | 7:16 | 11.76 | 68 | 45 | 10.52 | 97.1 | 8.07
2009-08-17 | 7:31 | 11.7 | 69 | 45 | 10.56 | 97.3 | 8.08
2009-08-17 | 7:46 | 11.65 | 69 | 45 | 10.62 | 97.8 | 8.09
2009-08-17 | 8:01 | 11.63 | 69 | 45 | 10.6 | 97.6 | 8.11
2009-08-17 | 8:16 | 11.61 | 69 | 45 | 10.66 | 98 | 8.12
2009-08-17 | 8:31 | 11.62 | 69 | 45 | 10.7 | 98.4 | 8.13
2009-08-17 | 8:46 | 11.65 | 69 | 45 | 10.66 | 98.2 | 8.14
2009-08-17 | 9:01 | 11.7 | 69 | 45 | 10.7 | 98.6 | 8.14
2009-08-17 | 9:16 | 11.73 | 69 | 45 | 10.73 | 98.9 | 8.14
2009-08-17 | 9:31 | 11.8 | 69 | 45 | 10.72 | 99 | 8.15
2009-08-17 | 9:46 | 11.88 | 69 | 45 | 10.72 | 99.2 | 8.16
2009-08-17 | 10:01 | 11.97 | 69 | 45 | 10.72 | 99.4 | 8.17
2009-08-17 | 10:16 | 12.08 | 69 | 45 | 10.71 | 99.6 | 8.18
2009-08-17 | 10:31 | 12.22 | 69 | 45 | 10.7 | 99.8 | 8.18
2009-08-17 | 10:46 | 12.37 | 69 | 45 | 10.7 | 100.1 | 8.18
2009-08-17 | 11:01 | 12.44 | 69 | 45 | 10.68 | 100.1 | 8.19
2009-08-17 | 11:16 | 12.61 | 69 | 45 | 10.61 | 99.8 | 8.2
2009-08-17 | 11:31 | 12.89 | 69 | 45 | 10.58 | 100.1 | 8.21
2009-08-17 | 11:46 | 13.23 | 69 | 45 | 10.51 | 100.3 | 8.21
2009-08-17 | 12:01 | 13.6 | 68 | 44 | 10.44 | 100.4 | 8.21
2009-08-17 | 12:16 | 14.02 | 68 | 44 | 10.28 | 99.8 | 8.22
2009-08-17 | 12:31 | 14.5 | 68 | 44 | 10.2 | 100 | 8.22
2009-08-17 | 12:46 | 14.97 | 68 | 44 | 10.1 | 100.1 | 8.24
2009-08-17 | 13:01 | 15.48 | 68 | 44 | 9.94 | 99.6 | 8.25
2009-08-17 | 13:16 | 16.01 | 67 | 44 | 9.84 | 99.7 | 8.26
2009-08-17 | 13:31 | 16.55 | 67 | 44 | 9.69 | 99.3 | 8.26
2009-08-17 | 13:46 | 17.09 | 67 | 44 | 9.59 | 99.5 | 8.29
2009-08-17 | 14:01 | 17.6 | 67 | 43 | 9.46 | 99.1 | 8.29
2009-08-17 | 14:16 | 18.03 | 67 | 43 | 9.36 | 98.9 | 8.31
2009-08-17 | 14:31 | 18.46 | 66 | 43 | 9.26 | 98.8 | 8.33
2009-08-17 | 14:46 | 18.83 | 66 | 43 | 9.17 | 98.5 | 8.34
2009-08-17 | 15:01 | 19.1 | 66 | 43 | 9.12 | 98.5 | 8.34
2009-08-17 | 15:16 | 19.29 | 66 | 43 | 9.08 | 98.5 | 8.34
2009-08-17 | 15:31 | 19.59 | 66 | 43 | 8.99 | 98.1 | 8.35
2009-08-17 | 15:46 | 19.79 | 66 | 43 | 8.97 | 98.2 | 8.35
2009-08-17 | 16:01 | 19.89 | 66 | 43 | 8.92 | 97.8 | 8.36
2009-08-17 | 16:16 | 19.97 | 66 | 43 | 8.89 | 97.7 | 8.34
2009-08-17 | 16:31 | 19.97 | 65 | 43 | 8.86 | 97.4 | 8.33
2009-08-17 | 16:46 | 19.88 | 65 | 42 | 8.86 | 97.3 | 8.32
2009-08-17 | 17:01 | 19.78 | 65 | 42 | 8.87 | 97.1 | 8.3
2009-08-17 | 17:16 | 19.7 | 65 | 42 | 8.84 | 96.6 | 8.28
2009-08-17 | 17:31 | 19.58 | 65 | 42 | 8.85 | 96.5 | 8.26
2009-08-17 | 17:46 | 19.48 | 65 | 42 | 8.81 | 95.9 | 8.24
2009-08-17 | 18:01 | 19.33 | 65 | 42 | 8.84 | 95.9 | 8.22
2009-08-17 | 18:16 | 19.18 | 65 | 42 | 8.85 | 95.7 | 8.19
2009-08-17 | 18:31 | 19.07 | 65 | 42 | 8.85 | 95.5 | 8.17
2009-08-17 | 18:46 | 18.97 | 65 | 42 | 8.83 | 95.1 | 8.15
2009-08-17 | 19:01 | 18.89 | 65 | 42 | 8.86 | 95.3 | 8.14
2009-08-17 | 19:16 | 18.83 | 65 | 42 | 8.88 | 95.4 | 8.13
2009-08-17 | 19:31 | 18.76 | 65 | 42 | 8.84 | 94.8 | 8.12
2009-08-17 | 19:46 | 18.68 | 65 | 42 | 8.82 | 94.5 | 8.11
2009-08-17 | 20:01 | 18.58 | 65 | 42 | 8.84 | 94.4 | 8.09
2009-08-17 | 20:16 | 18.46 | 65 | 42 | 8.81 | 94 | 8.08
2009-08-17 | 20:31 | 18.34 | 65 | 42 | 8.82 | 93.8 | 8.06
2009-08-17 | 20:46 | 18.22 | 65 | 42 | 8.83 | 93.7 | 8.06
2009-08-17 | 21:01 | 18.1 | 65 | 42 | 8.86 | 93.8 | 8.05
2009-08-17 | 21:16 | 17.97 | 65 | 42 | 8.91 | 94 | 8.05
2009-08-17 | 21:31 | 17.82 | 65 | 42 | 8.97 | 94.4 | 8.05
2009-08-17 | 21:46 | 17.67 | 65 | 42 | 9.01 | 94.5 | 8.04
2009-08-17 | 22:01 | 17.49 | 65 | 42 | 9 | 94.1 | 8.04
2009-08-17 | 22:16 | 17.31 | 65 | 42 | 9.04 | 94.2 | 8.04
2009-08-17 | 22:31 | 17.13 | 65 | 42 | 9.11 | 94.5 | 8.04
2009-08-17 | 22:46 | 16.94 | 65 | 42 | 9.15 | 94.6 | 8.04
2009-08-17 | 23:01 | 16.74 | 64 | 42 | 9.19 | 94.6 | 8.03
2009-08-17 | 23:16 | 16.54 | 65 | 42 | 9.22 | 94.5 | 8.04
2009-08-17 | 23:31 | 16.35 | 64 | 42 | 9.26 | 94.5 | 8.03
2009-08-17 | 23:46 | 16.15 | 64 | 42 | 9.3 | 94.5 | 8.03
2009-08-18 | 0:01 | 15.96 | 64 | 42 | 9.36 | 94.8 | 8.03
2009-08-18 | 0:16 | 15.78 | 64 | 42 | 9.42 | 95 | 8.03
2009-08-18 | 0:31 | 15.6 | 64 | 42 | 9.47 | 95.1 | 8.03
2009-08-18 | 0:46 | 15.44 | 64 | 42 | 9.49 | 95 | 8.03
2009-08-18 | 1:01 | 15.26 | 64 | 42 | 9.55 | 95.2 | 8.03
2009-08-18 | 1:16 | 15.1 | 64 | 42 | 9.55 | 94.9 | 8.03
2009-08-18 | 1:31 | 14.95 | 64 | 42 | 9.59 | 95 | 8.03
2009-08-18 | 1:46 | 14.81 | 64 | 42 | 9.63 | 95.1 | 8.03
2009-08-18 | 2:01 | 14.66 | 64 | 42 | 9.7 | 95.5 | 8.03
2009-08-18 | 2:16 | 14.53 | 64 | 42 | 9.72 | 95.4 | 8.03
2009-08-18 | 2:31 | 14.39 | 64 | 42 | 9.73 | 95.2 | 8.02
2009-08-18 | 2:46 | 14.25 | 64 | 42 | 9.76 | 95.3 | 8.02
2009-08-18 | 3:01 | 14.12 | 64 | 42 | 9.82 | 95.5 | 8.03
2009-08-18 | 3:16 | 14 | 64 | 42 | 9.82 | 95.3 | 8.03
2009-08-18 | 3:31 | 13.88 | 65 | 42 | 9.88 | 95.6 | 8.02
2009-08-18 | 3:46 | 13.76 | 65 | 42 | 9.91 | 95.6 | 8.03
2009-08-18 | 4:01 | 13.65 | 65 | 42 | 9.95 | 95.8 | 8.02
2009-08-18 | 4:16 | 13.53 | 65 | 42 | 9.99 | 96 | 8.02
2009-08-18 | 4:31 | 13.42 | 65 | 42 | 10.01 | 95.9 | 8.02
2009-08-18 | 4:46 | 13.3 | 65 | 42 | 10.02 | 95.7 | 8.02
2009-08-18 | 5:01 | 13.2 | 65 | 42 | 10 | 95.4 | 8.02
2009-08-18 | 5:16 | 13.1 | 65 | 42 | 10.07 | 95.7 | 8.03
2009-08-18 | 5:31 | 13 | 65 | 42 | 10.11 | 95.9 | 8.02
2009-08-18 | 5:46 | 12.9 | 66 | 43 | 10.13 | 96 | 8.03
2009-08-18 | 6:01 | 12.81 | 66 | 43 | 10.15 | 96 | 8.03
2009-08-18 | 6:16 | 12.72 | 66 | 43 | 10.19 | 96.1 | 8.04
2009-08-18 | 6:31 | 12.63 | 66 | 43 | 10.22 | 96.2 | 8.04
2009-08-18 | 6:46 | 12.55 | 66 | 43 | 10.27 | 96.5 | 8.04
2009-08-18 | 7:01 | 12.47 | 67 | 43 | 10.34 | 97 | 8.05
2009-08-18 | 7:16 | 12.39 | 67 | 43 | 10.37 | 97.1 | 8.06
2009-08-18 | 7:31 | 12.32 | 67 | 44 | 10.39 | 97.1 | 8.07
2009-08-18 | 7:46 | 12.29 | 67 | 44 | 10.43 | 97.4 | 8.08
2009-08-18 | 8:01 | 12.26 | 68 | 44 | 10.47 | 97.8 | 8.09
2009-08-18 | 8:16 | 12.24 | 68 | 44 | 10.46 | 97.6 | 8.09
2009-08-18 | 8:31 | 12.25 | 68 | 44 | 10.45 | 97.5 | 8.11
2009-08-18 | 8:46 | 12.29 | 69 | 45 | 10.49 | 98 | 8.12
2009-08-18 | 9:01 | 12.37 | 69 | 45 | 10.45 | 97.8 | 8.13
2009-08-18 | 9:16 | 12.43 | 69 | 45 | 10.5 | 98.4 | 8.14
2009-08-18 | 9:31 | 12.5 | 70 | 45 | 10.5 | 98.6 | 8.15
2009-08-18 | 9:46 | 12.59 | 70 | 45 | 10.52 | 98.9 | 8.16
2009-08-18 | 10:01 | 12.67 | 70 | 46 | 10.52 | 99.1 | 8.16
2009-08-18 | 10:16 | 12.79 | 70 | 46 | 10.45 | 98.8 | 8.17
2009-08-18 | 10:31 | 12.92 | 71 | 46 | 10.43 | 98.8 | 8.19
2009-08-18 | 10:46 | 13.03 | 71 | 46 | 10.48 | 99.5 | 8.19
2009-08-18 | 11:01 | 13.23 | 71 | 46 | 10.48 | 100 | 8.21
2009-08-18 | 11:16 | 13.46 | 71 | 46 | 10.44 | 100.1 | 8.21
2009-08-18 | 11:31 | 13.71 | 71 | 46 | 10.38 | 100.1 | 8.22
2009-08-18 | 11:46 | 14.02 | 71 | 46 | 10.28 | 99.8 | 8.22
2009-08-18 | 12:01 | 14.39 | 71 | 46 | 10.17 | 99.5 | 8.23
2009-08-18 | 12:16 | 14.83 | 71 | 46 | 10.06 | 99.4 | 8.24
2009-08-18 | 12:31 | 15.3 | 71 | 46 | 9.94 | 99.3 | 8.25
2009-08-18 | 12:46 | 15.79 | 71 | 46 | 9.86 | 99.4 | 8.27
2009-08-18 | 13:01 | 16.31 | 71 | 46 | 9.76 | 99.5 | 8.27
2009-08-18 | 13:16 | 16.86 | 71 | 46 | 9.65 | 99.5 | 8.29
2009-08-18 | 13:31 | 17.42 | 71 | 46 | 9.5 | 99.1 | 8.3
2009-08-18 | 13:46 | 17.98 | 71 | 46 | 9.4 | 99.3 | 8.31
2009-08-18 | 14:01 | 18.55 | 71 | 46 | 9.29 | 99.3 | 8.32
2009-08-18 | 14:16 | 19.13 | 71 | 46 | 9.17 | 99.1 | 8.33
2009-08-18 | 14:31 | 19.71 | 70 | 46 | 9.05 | 98.9 | 8.34
2009-08-18 | 14:46 | 20.21 | 70 | 46 | 8.89 | 98.3 | 8.35
2009-08-18 | 15:01 | 20.67 | 70 | 46 | 8.79 | 97.9 | 8.36
2009-08-18 | 15:16 | 21.02 | 70 | 45 | 8.72 | 97.9 | 8.37
2009-08-18 | 15:31 | 21.38 | 70 | 45 | 8.6 | 97.2 | 8.36
2009-08-18 | 15:46 | 21.71 | 70 | 45 | 8.56 | 97.3 | 8.36
2009-08-18 | 16:01 | 21.95 | 70 | 45 | 8.5 | 97.2 | 8.36
2009-08-18 | 16:16 | 22.06 | 70 | 45 | 8.45 | 96.8 | 8.35
2009-08-18 | 16:31 | 22.08 | 69 | 45 | 8.43 | 96.6 | 8.35
2009-08-18 | 16:46 | 22.05 | 69 | 45 | 8.37 | 95.8 | 8.33
2009-08-18 | 17:01 | 21.95 | 69 | 45 | 8.37 | 95.6 | 8.32
2009-08-18 | 17:16 | 21.81 | 69 | 45 | 8.35 | 95.1 | 8.29
2009-08-18 | 17:31 | 21.61 | 69 | 45 | 8.37 | 95 | 8.26
2009-08-18 | 17:46 | 21.45 | 69 | 45 | 8.39 | 94.9 | 8.24
2009-08-18 | 18:01 | 21.27 | 69 | 45 | 8.36 | 94.3 | 8.22
2009-08-18 | 18:16 | 21.11 | 69 | 45 | 8.38 | 94.2 | 8.21
2009-08-18 | 18:31 | 20.95 | 69 | 45 | 8.39 | 94 | 8.19
2009-08-18 | 18:46 | 20.82 | 68 | 45 | 8.4 | 93.9 | 8.17
2009-08-18 | 19:01 | 20.73 | 68 | 44 | 8.4 | 93.8 | 8.16
2009-08-18 | 19:16 | 20.67 | 68 | 44 | 8.42 | 93.8 | 8.14
2009-08-18 | 19:31 | 20.6 | 68 | 44 | 8.43 | 93.9 | 8.13
2009-08-18 | 19:46 | 20.54 | 68 | 44 | 8.39 | 93.3 | 8.12
2009-08-18 | 20:01 | 20.46 | 68 | 44 | 8.42 | 93.4 | 8.11
2009-08-18 | 20:16 | 20.36 | 68 | 44 | 8.4 | 93 | 8.1
2009-08-18 | 20:31 | 20.26 | 68 | 44 | 8.41 | 93 | 8.09
2009-08-18 | 20:46 | 20.16 | 68 | 44 | 8.41 | 92.8 | 8.08
2009-08-18 | 21:01 | 20.04 | 68 | 44 | 8.42 | 92.7 | 8.07
2009-08-18 | 21:16 | 19.91 | 68 | 44 | 8.46 | 92.8 | 8.07
2009-08-18 | 21:31 | 19.78 | 68 | 44 | 8.48 | 92.9 | 8.07
2009-08-18 | 21:46 | 19.62 | 68 | 44 | 8.5 | 92.8 | 8.07
2009-08-18 | 22:01 | 19.46 | 68 | 44 | 8.54 | 92.9 | 8.06
2009-08-18 | 22:16 | 19.27 | 68 | 44 | 8.59 | 93.1 | 8.06
2009-08-18 | 22:31 | 19.07 | 68 | 44 | 8.64 | 93.3 | 8.07
2009-08-18 | 22:46 | 18.86 | 68 | 44 | 8.7 | 93.5 | 8.06
2009-08-18 | 23:01 | 18.65 | 68 | 44 | 8.74 | 93.5 | 8.06
2009-08-18 | 23:16 | 18.45 | 68 | 44 | 8.75 | 93.3 | 8.06
2009-08-18 | 23:31 | 18.24 | 68 | 44 | 8.79 | 93.3 | 8.05
2009-08-18 | 23:46 | 18.04 | 68 | 44 | 8.84 | 93.4 | 8.06
2009-08-19 | 0:01 | 17.83 | 68 | 44 | 8.89 | 93.6 | 8.06
2009-08-19 | 0:16 | 17.64 | 68 | 44 | 8.97 | 94 | 8.05
2009-08-19 | 0:31 | 17.44 | 68 | 44 | 8.98 | 93.8 | 8.05
2009-08-19 | 0:46 | 17.25 | 68 | 44 | 9.03 | 94 | 8.05
2009-08-19 | 1:01 | 17.06 | 68 | 44 | 9.1 | 94.3 | 8.05
2009-08-19 | 1:16 | 16.88 | 68 | 44 | 9.13 | 94.2 | 8.05
2009-08-19 | 1:31 | 16.71 | 68 | 44 | 9.18 | 94.4 | 8.05
2009-08-19 | 1:46 | 16.54 | 68 | 44 | 9.21 | 94.4 | 8.05
2009-08-19 | 2:01 | 16.38 | 68 | 44 | 9.25 | 94.5 | 8.05
2009-08-19 | 2:16 | 16.22 | 68 | 44 | 9.28 | 94.5 | 8.05
2009-08-19 | 2:31 | 16.07 | 68 | 44 | 9.33 | 94.6 | 8.04
2009-08-19 | 2:46 | 15.92 | 68 | 44 | 9.36 | 94.7 | 8.05
2009-08-19 | 3:01 | 15.78 | 68 | 44 | 9.4 | 94.8 | 8.04
2009-08-19 | 3:16 | 15.64 | 68 | 44 | 9.45 | 95 | 8.04
2009-08-19 | 3:31 | 15.5 | 68 | 44 | 9.43 | 94.6 | 8.04
2009-08-19 | 3:46 | 15.37 | 68 | 44 | 9.48 | 94.8 | 8.05
2009-08-19 | 4:01 | 15.24 | 68 | 44 | 9.53 | 95 | 8.05
2009-08-19 | 4:16 | 15.11 | 68 | 44 | 9.55 | 94.9 | 8.04
2009-08-19 | 4:31 | 14.99 | 68 | 44 | 9.59 | 95.1 | 8.04
2009-08-19 | 4:46 | 14.87 | 68 | 44 | 9.59 | 94.9 | 8.04
2009-08-19 | 5:01 | 14.75 | 68 | 44 | 9.63 | 95 | 8.04
2009-08-19 | 5:16 | 14.63 | 68 | 44 | 9.67 | 95.2 | 8.05
2009-08-19 | 5:31 | 14.52 | 68 | 44 | 9.72 | 95.4 | 8.05
2009-08-19 | 5:46 | 14.4 | 69 | 45 | 9.7 | 95 | 8.04
2009-08-19 | 6:01 | 14.29 | 69 | 45 | 9.74 | 95.1 | 8.05
2009-08-19 | 6:16 | 14.18 | 69 | 45 | 9.78 | 95.3 | 8.05
2009-08-19 | 6:31 | 14.09 | 69 | 45 | 9.84 | 95.7 | 8.05
2009-08-19 | 6:46 | 14 | 69 | 45 | 9.83 | 95.4 | 8.05
2009-08-19 | 7:01 | 13.92 | 69 | 45 | 9.84 | 95.3 | 8.07
2009-08-19 | 7:16 | 13.84 | 70 | 45 | 9.92 | 95.9 | 8.07
2009-08-19 | 7:31 | 13.77 | 70 | 45 | 9.97 | 96.3 | 8.07
2009-08-19 | 7:46 | 13.73 | 70 | 46 | 9.96 | 96.1 | 8.09
2009-08-19 | 8:01 | 13.69 | 70 | 46 | 9.98 | 96.2 | 8.09
2009-08-19 | 8:16 | 13.65 | 71 | 46 | 9.98 | 96.1 | 8.1
2009-08-19 | 8:31 | 13.64 | 71 | 46 | 10.01 | 96.4 | 8.11
2009-08-19 | 8:46 | 13.67 | 71 | 46 | 10.06 | 96.9 | 8.12
2009-08-19 | 9:01 | 13.74 | 71 | 46 | 10.07 | 97.1 | 8.13
2009-08-19 | 9:16 | 13.78 | 71 | 46 | 10.11 | 97.6 | 8.14
2009-08-19 | 9:31 | 13.84 | 71 | 46 | 10.11 | 97.8 | 8.15
2009-08-19 | 9:46 | 13.92 | 72 | 47 | 10.13 | 98.1 | 8.16
2009-08-19 | 10:01 | 14.01 | 72 | 47 | 10.15 | 98.6 | 8.18
2009-08-19 | 10:16 | 14.1 | 72 | 47 | 10.15 | 98.7 | 8.19
2009-08-19 | 10:31 | 14.23 | 72 | 47 | 10.1 | 98.5 | 8.2
2009-08-19 | 10:46 | 14.36 | 72 | 47 | 10.11 | 98.9 | 8.22
2009-08-19 | 11:01 | 14.49 | 72 | 47 | 10.11 | 99.1 | 8.22
2009-08-19 | 11:16 | 14.68 | 72 | 47 | 10.06 | 99 | 8.24
2009-08-19 | 11:31 | 14.92 | 72 | 47 | 10 | 99 | 8.24
2009-08-19 | 11:46 | 15.21 | 72 | 47 | 9.93 | 98.9 | 8.26
2009-08-19 | 12:01 | 15.58 | 71 | 46 | 9.89 | 99.3 | 8.26
2009-08-19 | 12:16 | 16 | 71 | 46 | 9.81 | 99.4 | 8.29
2009-08-19 | 12:31 | 16.44 | 71 | 46 | 9.75 | 99.7 | 8.29
2009-08-19 | 12:46 | 16.92 | 71 | 46 | 9.59 | 99.1 | 8.31
2009-08-19 | 13:01 | 17.44 | 70 | 46 | 9.46 | 98.8 | 8.32
2009-08-19 | 13:16 | 17.97 | 70 | 46 | 9.34 | 98.6 | 8.33
2009-08-19 | 13:31 | 18.51 | 70 | 45 | 9.26 | 98.9 | 8.35
2009-08-19 | 13:46 | 19.02 | 70 | 45 | 9.12 | 98.4 | 8.37
2009-08-19 | 14:01 | 19.52 | 69 | 45 | 9.03 | 98.3 | 8.37
2009-08-19 | 14:16 | 20.07 | 69 | 45 | 8.91 | 98.1 | 8.39
2009-08-19 | 14:31 | 20.58 | 69 | 45 | 8.84 | 98.3 | 8.41
2009-08-19 | 14:46 | 21.06 | 69 | 45 | 8.72 | 97.9 | 8.41
2009-08-19 | 15:01 | 21.49 | 69 | 45 | 8.63 | 97.8 | 8.41
2009-08-19 | 15:16 | 21.88 | 68 | 45 | 8.55 | 97.5 | 8.41
2009-08-19 | 15:31 | 22.26 | 68 | 44 | 8.46 | 97.3 | 8.4
2009-08-19 | 15:46 | 22.53 | 68 | 44 | 8.4 | 97 | 8.42
2009-08-19 | 16:01 | 22.73 | 68 | 44 | 8.34 | 96.8 | 8.42
2009-08-19 | 16:16 | 22.88 | 68 | 44 | 8.3 | 96.6 | 8.41
2009-08-19 | 16:31 | 22.93 | 68 | 44 | 8.28 | 96.4 | 8.4
2009-08-19 | 16:46 | 22.91 | 68 | 44 | 8.25 | 96 | 8.38
2009-08-19 | 17:01 | 22.83 | 68 | 44 | 8.23 | 95.6 | 8.35
2009-08-19 | 17:16 | 22.72 | 68 | 44 | 8.21 | 95.2 | 8.32
2009-08-19 | 17:31 | 22.52 | 67 | 44 | 8.21 | 94.8 | 8.29
2009-08-19 | 17:46 | 22.36 | 67 | 44 | 8.2 | 94.5 | 8.26
2009-08-19 | 18:01 | 22.18 | 67 | 44 | 8.22 | 94.3 | 8.22
2009-08-19 | 18:16 | 22 | 67 | 44 | 8.24 | 94.2 | 8.2
2009-08-19 | 18:31 | 21.86 | 67 | 44 | 8.27 | 94.3 | 8.18
2009-08-19 | 18:46 | 21.7 | 67 | 44 | 8.26 | 93.9 | 8.17
2009-08-19 | 19:01 | 21.64 | 67 | 44 | 8.27 | 93.9 | 8.15
2009-08-19 | 19:16 | 21.58 | 67 | 44 | 8.23 | 93.4 | 8.14
2009-08-19 | 19:31 | 21.53 | 67 | 44 | 8.25 | 93.5 | 8.12
2009-08-19 | 19:46 | 21.5 | 67 | 44 | 8.21 | 93 | 8.11
2009-08-19 | 20:01 | 21.46 | 67 | 44 | 8.22 | 93.1 | 8.09
2009-08-19 | 20:16 | 21.4 | 67 | 44 | 8.23 | 93 | 8.08
2009-08-19 | 20:31 | 21.33 | 67 | 43 | 8.23 | 92.9 | 8.07
2009-08-19 | 20:46 | 21.24 | 67 | 43 | 8.19 | 92.3 | 8.06
2009-08-19 | 21:01 | 21.14 | 67 | 44 | 8.24 | 92.7 | 8.06
2009-08-19 | 21:16 | 21.04 | 67 | 44 | 8.23 | 92.4 | 8.05
2009-08-19 | 21:31 | 20.92 | 67 | 44 | 8.25 | 92.4 | 8.05
2009-08-19 | 21:46 | 20.77 | 67 | 44 | 8.26 | 92.3 | 8.04
2009-08-19 | 22:01 | 20.61 | 67 | 44 | 8.32 | 92.7 | 8.04
2009-08-19 | 22:16 | 20.43 | 67 | 43 | 8.34 | 92.5 | 8.04
2009-08-19 | 22:31 | 20.24 | 67 | 44 | 8.38 | 92.7 | 8.04
2009-08-19 | 22:46 | 20.05 | 67 | 43 | 8.42 | 92.7 | 8.03
2009-08-19 | 23:01 | 19.85 | 67 | 43 | 8.48 | 93 | 8.04
2009-08-19 | 23:16 | 19.67 | 67 | 43 | 8.54 | 93.3 | 8.04
2009-08-19 | 23:31 | 19.48 | 67 | 43 | 8.56 | 93.2 | 8.03
2009-08-19 | 23:46 | 19.28 | 67 | 43 | 8.58 | 93 | 8.03
2009-08-20 | 0:01 | 19.09 | 67 | 43 | 8.64 | 93.3 | 8.03
2009-08-20 | 0:16 | 18.91 | 67 | 43 | 8.66 | 93.2 | 8.03
2009-08-20 | 0:31 | 18.73 | 67 | 43 | 8.69 | 93.2 | 8.03
2009-08-20 | 0:46 | 18.56 | 67 | 43 | 8.76 | 93.5 | 8.03
2009-08-20 | 1:01 | 18.39 | 67 | 43 | 8.76 | 93.3 | 8.03
2009-08-20 | 1:16 | 18.23 | 67 | 43 | 8.81 | 93.5 | 8.03
2009-08-20 | 1:31 | 18.08 | 67 | 43 | 8.83 | 93.4 | 8.03
2009-08-20 | 1:46 | 17.93 | 67 | 43 | 8.89 | 93.8 | 8.03
2009-08-20 | 2:01 | 17.79 | 67 | 43 | 8.92 | 93.8 | 8.03
2009-08-20 | 2:16 | 17.66 | 67 | 43 | 8.97 | 94.1 | 8.03
2009-08-20 | 2:31 | 17.52 | 67 | 43 | 8.95 | 93.7 | 8.03
2009-08-20 | 2:46 | 17.39 | 67 | 43 | 8.98 | 93.7 | 8.03
2009-08-20 | 3:01 | 17.26 | 67 | 43 | 9.03 | 93.9 | 8.04
2009-08-20 | 3:16 | 17.13 | 67 | 43 | 9.06 | 94 | 8.03
2009-08-20 | 3:31 | 16.99 | 67 | 43 | 9.09 | 94.1 | 8.04
2009-08-20 | 3:46 | 16.86 | 67 | 43 | 9.14 | 94.3 | 8.04
2009-08-20 | 4:01 | 16.74 | 67 | 43 | 9.17 | 94.4 | 8.04
2009-08-20 | 4:16 | 16.61 | 67 | 44 | 9.21 | 94.6 | 8.04
2009-08-20 | 4:31 | 16.49 | 67 | 44 | 9.2 | 94.1 | 8.04
2009-08-20 | 4:46 | 16.37 | 67 | 44 | 9.27 | 94.6 | 8.04
2009-08-20 | 5:01 | 16.24 | 67 | 44 | 9.28 | 94.5 | 8.04
2009-08-20 | 5:16 | 16.13 | 68 | 44 | 9.33 | 94.8 | 8.04
2009-08-20 | 5:31 | 16.01 | 68 | 44 | 9.32 | 94.5 | 8.04
2009-08-20 | 5:46 | 15.91 | 68 | 44 | 9.33 | 94.3 | 8.04
2009-08-20 | 6:01 | 15.81 | 68 | 44 | 9.39 | 94.8 | 8.04
2009-08-20 | 6:16 | 15.7 | 68 | 44 | 9.43 | 94.9 | 8.05
2009-08-20 | 6:31 | 15.61 | 69 | 45 | 9.43 | 94.8 | 8.05
2009-08-20 | 6:46 | 15.52 | 69 | 45 | 9.45 | 94.8 | 8.06
2009-08-20 | 7:01 | 15.44 | 69 | 45 | 9.52 | 95.3 | 8.06
2009-08-20 | 7:16 | 15.36 | 69 | 45 | 9.58 | 95.7 | 8.08
2009-08-20 | 7:31 | 15.3 | 70 | 45 | 9.61 | 95.9 | 8.08
2009-08-20 | 7:46 | 15.23 | 70 | 45 | 9.62 | 95.9 | 8.1
2009-08-20 | 8:01 | 15.17 | 70 | 46 | 9.66 | 96.1 | 8.11
2009-08-20 | 8:16 | 15.11 | 70 | 46 | 9.69 | 96.3 | 8.12
2009-08-20 | 8:31 | 15.08 | 70 | 46 | 9.7 | 96.3 | 8.12
2009-08-20 | 8:46 | 15.08 | 70 | 46 | 9.7 | 96.4 | 8.13
2009-08-20 | 9:01 | 15.11 | 71 | 46 | 9.75 | 97 | 8.14
2009-08-20 | 9:16 | 15.13 | 70 | 46 | 9.78 | 97.3 | 8.15
2009-08-20 | 9:31 | 15.16 | 70 | 46 | 9.81 | 97.6 | 8.15
2009-08-20 | 9:46 | 15.21 | 70 | 46 | 9.78 | 97.5 | 8.17
2009-08-20 | 10:01 | 15.29 | 70 | 46 | 9.77 | 97.5 | 8.18
2009-08-20 | 10:16 | 15.38 | 70 | 45 | 9.77 | 97.7 | 8.19
2009-08-20 | 10:31 | 15.48 | 70 | 45 | 9.76 | 97.8 | 8.18
2009-08-20 | 10:46 | 15.6 | 69 | 45 | 9.75 | 98 | 8.2
2009-08-20 | 11:01 | 15.73 | 69 | 45 | 9.8 | 98.8 | 8.21
2009-08-20 | 11:16 | 15.92 | 69 | 45 | 9.79 | 99 | 8.22
2009-08-20 | 11:31 | 16.15 | 69 | 45 | 9.75 | 99.1 | 8.22
2009-08-20 | 11:46 | 16.46 | 68 | 44 | 9.69 | 99.2 | 8.24
2009-08-20 | 12:01 | 16.83 | 68 | 44 | 9.58 | 98.8 | 8.25
2009-08-20 | 12:16 | 17.26 | 68 | 44 | 9.52 | 99.1 | 8.26
2009-08-20 | 12:31 | 17.73 | 67 | 44 | 9.39 | 98.6 | 8.27
2009-08-20 | 12:46 | 18.22 | 67 | 44 | 9.27 | 98.4 | 8.29
2009-08-20 | 13:01 | 18.74 | 67 | 43 | 9.2 | 98.7 | 8.3
2009-08-20 | 13:16 | 19.27 | 67 | 43 | 9.11 | 98.8 | 8.31
2009-08-20 | 13:31 | 19.8 | 66 | 43 | 9.01 | 98.7 | 8.33
2009-08-20 | 13:46 | 20.35 | 66 | 43 | 8.9 | 98.6 | 8.35
2009-08-20 | 14:01 | 20.89 | 66 | 43 | 8.77 | 98.2 | 8.35
2009-08-20 | 14:16 | 21.42 | 66 | 43 | 8.65 | 97.9 | 8.37
2009-08-20 | 14:31 | 21.94 | 66 | 43 | 8.57 | 97.9 | 8.38
2009-08-20 | 14:46 | 22.4 | 66 | 43 | 8.47 | 97.6 | 8.4
2009-08-20 | 15:01 | 22.79 | 65 | 43 | 8.37 | 97.2 | 8.4
2009-08-20 | 15:16 | 23.13 | 65 | 42 | 8.31 | 97.1 | 8.4
2009-08-20 | 15:31 | 23.43 | 65 | 42 | 8.25 | 97 | 8.41
2009-08-20 | 15:46 | 23.66 | 65 | 42 | 8.2 | 96.8 | 8.42
2009-08-20 | 16:01 | 23.81 | 65 | 42 | 8.16 | 96.6 | 8.42
2009-08-20 | 16:16 | 23.86 | 65 | 42 | 8.13 | 96.3 | 8.42
2009-08-20 | 16:31 | 23.82 | 65 | 42 | 8.12 | 96.1 | 8.4
2009-08-20 | 16:46 | 23.67 | 65 | 42 | 8.13 | 96 | 8.38
2009-08-20 | 17:01 | 23.52 | 65 | 42 | 8.12 | 95.6 | 8.36
2009-08-20 | 17:16 | 23.38 | 65 | 42 | 8.12 | 95.3 | 8.33
2009-08-20 | 17:31 | 23.14 | 65 | 42 | 8.13 | 95 | 8.29
2009-08-20 | 17:46 | 22.93 | 65 | 42 | 8.13 | 94.7 | 8.27
2009-08-20 | 18:01 | 22.77 | 65 | 42 | 8.15 | 94.6 | 8.23
2009-08-20 | 18:16 | 22.61 | 65 | 42 | 8.16 | 94.5 | 8.21
2009-08-20 | 18:31 | 22.45 | 65 | 42 | 8.17 | 94.3 | 8.2
2009-08-20 | 18:46 | 22.31 | 65 | 42 | 8.16 | 93.9 | 8.18
2009-08-20 | 19:01 | 22.16 | 65 | 42 | 8.19 | 94 | 8.17
2009-08-20 | 19:16 | 22.07 | 64 | 42 | 8.16 | 93.4 | 8.15
2009-08-20 | 19:31 | 21.96 | 64 | 42 | 8.14 | 93.1 | 8.13
2009-08-20 | 19:46 | 21.84 | 64 | 42 | 8.15 | 92.9 | 8.12
2009-08-20 | 20:01 | 21.7 | 64 | 42 | 8.16 | 92.8 | 8.11
2009-08-20 | 20:16 | 21.61 | 64 | 42 | 8.18 | 92.8 | 8.09
2009-08-20 | 20:31 | 21.49 | 64 | 42 | 8.2 | 92.8 | 8.08
2009-08-20 | 20:46 | 21.34 | 64 | 42 | 8.23 | 92.9 | 8.07
2009-08-20 | 21:01 | 21.19 | 64 | 42 | 8.25 | 92.9 | 8.06
2009-08-20 | 21:16 | 21.11 | 64 | 42 | 8.27 | 93 | 8.05
2009-08-20 | 21:31 | 20.99 | 64 | 42 | 8.28 | 92.8 | 8.05
2009-08-20 | 21:46 | 20.83 | 64 | 42 | 8.32 | 93 | 8.04
2009-08-20 | 22:01 | 20.67 | 64 | 42 | 8.36 | 93.2 | 8.05
2009-08-20 | 22:16 | 20.5 | 64 | 42 | 8.37 | 93 | 8.04
2009-08-20 | 22:31 | 20.34 | 64 | 42 | 8.43 | 93.4 | 8.04
2009-08-20 | 22:46 | 20.13 | 64 | 42 | 8.48 | 93.6 | 8.03
2009-08-20 | 23:01 | 19.93 | 64 | 42 | 8.49 | 93.3 | 8.04
2009-08-20 | 23:16 | 19.74 | 64 | 42 | 8.54 | 93.5 | 8.03
2009-08-20 | 23:31 | 19.57 | 64 | 42 | 8.6 | 93.8 | 8.03
2009-08-20 | 23:46 | 19.38 | 64 | 42 | 8.65 | 94 | 8.03
2009-08-21 | 0:01 | 19.18 | 64 | 42 | 8.69 | 94 | 8.03
2009-08-21 | 0:16 | 18.99 | 64 | 42 | 8.75 | 94.4 | 8.03
2009-08-21 | 0:31 | 18.83 | 64 | 42 | 8.79 | 94.4 | 8.03
2009-08-21 | 0:46 | 18.67 | 64 | 42 | 8.79 | 94.2 | 8.03
2009-08-21 | 1:01 | 18.52 | 64 | 42 | 8.83 | 94.3 | 8.03
2009-08-21 | 1:16 | 18.36 | 64 | 42 | 8.86 | 94.3 | 8.03
2009-08-21 | 1:31 | 18.22 | 64 | 42 | 8.88 | 94.2 | 8.03
2009-08-21 | 1:46 | 18.1 | 64 | 42 | 8.93 | 94.5 | 8.03
2009-08-21 | 2:01 | 18.01 | 64 | 42 | 8.95 | 94.5 | 8.03
2009-08-21 | 2:16 | 17.93 | 64 | 42 | 8.96 | 94.5 | 8.04
2009-08-21 | 2:31 | 17.84 | 65 | 42 | 8.96 | 94.3 | 8.04
2009-08-21 | 2:46 | 17.73 | 65 | 42 | 8.98 | 94.4 | 8.04
2009-08-21 | 3:01 | 17.62 | 65 | 42 | 8.99 | 94.2 | 8.04
2009-08-21 | 3:16 | 17.5 | 65 | 42 | 9.05 | 94.7 | 8.04
2009-08-21 | 3:31 | 17.39 | 65 | 42 | 9.07 | 94.6 | 8.04
2009-08-21 | 3:46 | 17.28 | 65 | 42 | 9.12 | 94.9 | 8.04
2009-08-21 | 4:01 | 17.16 | 65 | 42 | 9.15 | 95 | 8.04
2009-08-21 | 4:16 | 17.05 | 65 | 42 | 9.18 | 95.1 | 8.04
2009-08-21 | 4:31 | 16.95 | 66 | 43 | 9.17 | 94.7 | 8.04
2009-08-21 | 4:46 | 16.79 | 66 | 43 | 9.2 | 94.8 | 8.05
2009-08-21 | 5:01 | 16.62 | 66 | 43 | 9.26 | 95.1 | 8.05
2009-08-21 | 5:16 | 16.44 | 66 | 43 | 9.32 | 95.3 | 8.05
2009-08-21 | 5:31 | 16.28 | 66 | 43 | 9.36 | 95.4 | 8.05
2009-08-21 | 5:46 | 16.12 | 67 | 43 | 9.36 | 95.1 | 8.05
2009-08-21 | 6:01 | 15.97 | 67 | 44 | 9.45 | 95.6 | 8.05
2009-08-21 | 6:16 | 15.82 | 67 | 44 | 9.47 | 95.6 | 8.06
2009-08-21 | 6:31 | 15.67 | 67 | 44 | 9.51 | 95.7 | 8.06
2009-08-21 | 6:46 | 15.53 | 68 | 44 | 9.54 | 95.7 | 8.07
2009-08-21 | 7:01 | 15.38 | 68 | 44 | 9.61 | 96.1 | 8.08
2009-08-21 | 7:16 | 15.24 | 68 | 44 | 9.67 | 96.4 | 8.08
2009-08-21 | 7:31 | 15.11 | 68 | 45 | 9.71 | 96.5 | 8.09
2009-08-21 | 7:46 | 15 | 69 | 45 | 9.72 | 96.4 | 8.1
2009-08-21 | 8:01 | 14.89 | 69 | 45 | 9.78 | 96.8 | 8.1
2009-08-21 | 8:16 | 14.81 | 69 | 45 | 9.8 | 96.8 | 8.11
2009-08-21 | 8:31 | 14.74 | 69 | 45 | 9.85 | 97.2 | 8.11
2009-08-21 | 8:46 | 14.7 | 69 | 45 | 9.85 | 97.1 | 8.12
2009-08-21 | 9:01 | 14.71 | 69 | 45 | 9.91 | 97.6 | 8.13
2009-08-21 | 9:16 | 14.7 | 69 | 45 | 9.95 | 98 | 8.14
2009-08-21 | 9:31 | 14.72 | 68 | 44 | 9.96 | 98.2 | 8.15
2009-08-21 | 9:46 | 14.74 | 68 | 44 | 9.98 | 98.4 | 8.16
2009-08-21 | 10:01 | 14.79 | 68 | 44 | 9.99 | 98.7 | 8.17
2009-08-21 | 10:16 | 14.86 | 67 | 44 | 9.99 | 98.8 | 8.18
2009-08-21 | 10:31 | 14.94 | 67 | 44 | 10 | 99 | 8.18
2009-08-21 | 10:46 | 15.06 | 67 | 43 | 9.98 | 99.1 | 8.19
2009-08-21 | 11:01 | 15.19 | 66 | 43 | 9.94 | 99 | 8.2
2009-08-21 | 11:16 | 15.35 | 66 | 43 | 9.92 | 99.1 | 8.2
2009-08-21 | 11:31 | 15.6 | 66 | 43 | 9.9 | 99.4 | 8.21
2009-08-21 | 11:46 | 15.89 | 65 | 42 | 9.83 | 99.3 | 8.22
2009-08-21 | 12:01 | 16.25 | 65 | 42 | 9.76 | 99.5 | 8.23
2009-08-21 | 12:16 | 16.67 | 65 | 42 | 9.64 | 99 | 8.24
2009-08-21 | 12:31 | 17.1 | 64 | 42 | 9.55 | 99 | 8.25
2009-08-21 | 12:46 | 17.57 | 64 | 42 | 9.45 | 98.9 | 8.26
2009-08-21 | 13:01 | 18.05 | 64 | 41 | 9.34 | 98.8 | 8.27
2009-08-21 | 13:16 | 18.55 | 63 | 41 | 9.24 | 98.8 | 8.28
2009-08-21 | 13:31 | 19.02 | 63 | 41 | 9.17 | 98.9 | 8.31
2009-08-21 | 13:46 | 19.51 | 63 | 41 | 9.09 | 99 | 8.32
2009-08-21 | 14:01 | 20.02 | 63 | 41 | 9 | 99 | 8.33
2009-08-21 | 14:16 | 20.52 | 63 | 41 | 8.9 | 98.9 | 8.34
2009-08-21 | 14:31 | 20.98 | 63 | 41 | 8.82 | 98.9 | 8.35
2009-08-21 | 14:46 | 21.35 | 63 | 41 | 8.7 | 98.2 | 8.37
2009-08-21 | 15:01 | 21.69 | 63 | 41 | 8.6 | 97.7 | 8.39
2009-08-21 | 15:16 | 21.96 | 63 | 41 | 8.57 | 97.9 | 8.38
2009-08-21 | 15:31 | 22.2 | 63 | 41 | 8.48 | 97.3 | 8.38
2009-08-21 | 15:46 | 22.34 | 63 | 41 | 8.4 | 96.7 | 8.38
2009-08-21 | 16:01 | 22.45 | 63 | 41 | 8.39 | 96.8 | 8.38
2009-08-21 | 16:16 | 22.47 | 63 | 41 | 8.39 | 96.8 | 8.37
2009-08-21 | 16:31 | 22.38 | 62 | 41 | 8.39 | 96.7 | 8.35
2009-08-21 | 16:46 | 22.25 | 62 | 41 | 8.38 | 96.3 | 8.33
2009-08-21 | 17:01 | 22.08 | 62 | 41 | 8.37 | 95.8 | 8.31
2009-08-21 | 17:16 | 21.86 | 62 | 41 | 8.4 | 95.8 | 8.28
2009-08-21 | 17:31 | 21.62 | 62 | 41 | 8.41 | 95.5 | 8.24
2009-08-21 | 17:46 | 21.45 | 62 | 41 | 8.4 | 95.1 | 8.21
2009-08-21 | 18:01 | 21.23 | 62 | 41 | 8.41 | 94.8 | 8.19
2009-08-21 | 18:16 | 21.04 | 62 | 41 | 8.42 | 94.5 | 8.16
2009-08-21 | 18:31 | 20.84 | 62 | 40 | 8.46 | 94.6 | 8.14
2009-08-21 | 18:46 | 20.67 | 62 | 40 | 8.46 | 94.3 | 8.13
2009-08-21 | 19:01 | 20.52 | 62 | 40 | 8.47 | 94.2 | 8.12
2009-08-21 | 19:16 | 20.42 | 62 | 40 | 8.48 | 94.1 | 8.11
2009-08-21 | 19:31 | 20.35 | 62 | 40 | 8.5 | 94.2 | 8.1
2009-08-21 | 19:46 | 20.25 | 62 | 40 | 8.53 | 94.3 | 8.09
2009-08-21 | 20:01 | 20.18 | 62 | 40 | 8.53 | 94.2 | 8.09
2009-08-21 | 20:16 | 20.1 | 62 | 40 | 8.52 | 93.9 | 8.07
2009-08-21 | 20:31 | 19.99 | 62 | 40 | 8.51 | 93.5 | 8.06
2009-08-21 | 20:46 | 19.85 | 62 | 40 | 8.53 | 93.5 | 8.05
2009-08-21 | 21:01 | 19.73 | 62 | 40 | 8.58 | 93.9 | 8.05
2009-08-21 | 21:16 | 19.63 | 62 | 40 | 8.61 | 94 | 8.04
2009-08-21 | 21:31 | 19.51 | 62 | 41 | 8.63 | 94 | 8.04
2009-08-21 | 21:46 | 19.34 | 62 | 40 | 8.67 | 94.1 | 8.03
2009-08-21 | 22:01 | 19.15 | 62 | 40 | 8.72 | 94.3 | 8.03
2009-08-21 | 22:16 | 18.96 | 62 | 40 | 8.76 | 94.3 | 8.03
2009-08-21 | 22:31 | 18.77 | 62 | 40 | 8.81 | 94.6 | 8.03
2009-08-21 | 22:46 | 18.58 | 62 | 40 | 8.85 | 94.6 | 8.03
2009-08-21 | 23:01 | 18.39 | 62 | 40 | 8.91 | 94.8 | 8.02
2009-08-21 | 23:16 | 18.19 | 62 | 40 | 8.92 | 94.6 | 8.03
2009-08-21 | 23:31 | 17.99 | 62 | 40 | 8.95 | 94.6 | 8.02
2009-08-21 | 23:46 | 17.8 | 62 | 40 | 8.99 | 94.6 | 8.02
2009-08-22 | 0:01 | 17.6 | 62 | 40 | 9.06 | 94.9 | 8.02
2009-08-22 | 0:16 | 17.42 | 62 | 40 | 9.09 | 94.9 | 8.03
2009-08-22 | 0:31 | 17.24 | 62 | 40 | 9.11 | 94.8 | 8.03
2009-08-22 | 0:46 | 17.03 | 62 | 40 | 9.16 | 94.9 | 8.03
2009-08-22 | 1:01 | 16.84 | 62 | 40 | 9.23 | 95.1 | 8.02
2009-08-22 | 1:16 | 16.64 | 62 | 40 | 9.28 | 95.3 | 8.02
2009-08-22 | 1:31 | 16.45 | 62 | 40 | 9.3 | 95.2 | 8.02
2009-08-22 | 1:46 | 16.26 | 62 | 40 | 9.36 | 95.3 | 8.02
2009-08-22 | 2:01 | 16.09 | 62 | 41 | 9.37 | 95.1 | 8.02
2009-08-22 | 2:16 | 15.91 | 62 | 41 | 9.45 | 95.5 | 8.02
2009-08-22 | 2:31 | 15.74 | 63 | 41 | 9.46 | 95.3 | 8.02
2009-08-22 | 2:46 | 15.58 | 63 | 41 | 9.52 | 95.5 | 8.03
2009-08-22 | 3:01 | 15.42 | 63 | 41 | 9.57 | 95.7 | 8.03
2009-08-22 | 3:16 | 15.26 | 63 | 41 | 9.62 | 95.9 | 8.03
2009-08-22 | 3:31 | 15.12 | 63 | 41 | 9.65 | 96 | 8.03
2009-08-22 | 3:46 | 14.97 | 63 | 41 | 9.67 | 95.8 | 8.03
2009-08-22 | 4:01 | 14.83 | 63 | 41 | 9.71 | 95.9 | 8.02
2009-08-22 | 4:16 | 14.69 | 64 | 41 | 9.72 | 95.7 | 8.03
2009-08-22 | 4:31 | 14.57 | 64 | 41 | 9.77 | 96 | 8.03
2009-08-22 | 4:46 | 14.44 | 64 | 42 | 9.78 | 95.8 | 8.03
2009-08-22 | 5:01 | 14.31 | 64 | 42 | 9.84 | 96.1 | 8.03
2009-08-22 | 5:16 | 14.18 | 64 | 42 | 9.86 | 96.1 | 8.04
2009-08-22 | 5:31 | 14.05 | 65 | 42 | 9.88 | 96 | 8.04
2009-08-22 | 5:46 | 13.92 | 65 | 42 | 9.95 | 96.4 | 8.04
2009-08-22 | 6:01 | 13.79 | 65 | 42 | 9.95 | 96.1 | 8.04
2009-08-22 | 6:16 | 13.66 | 65 | 43 | 10.01 | 96.4 | 8.05
2009-08-22 | 6:31 | 13.53 | 66 | 43 | 10.06 | 96.6 | 8.05
2009-08-22 | 6:46 | 13.4 | 66 | 43 | 10.1 | 96.7 | 8.05
2009-08-22 | 7:01 | 13.27 | 66 | 43 | 10.16 | 97 | 8.06
2009-08-22 | 7:16 | 13.15 | 66 | 43 | 10.18 | 96.9 | 8.07
2009-08-22 | 7:31 | 13.04 | 67 | 43 | 10.26 | 97.4 | 8.07
2009-08-22 | 7:46 | 12.94 | 67 | 43 | 10.29 | 97.6 | 8.08
2009-08-22 | 8:01 | 12.85 | 67 | 43 | 10.34 | 97.8 | 8.09
2009-08-22 | 8:16 | 12.79 | 67 | 44 | 10.38 | 98.1 | 8.09
2009-08-22 | 8:31 | 12.74 | 67 | 44 | 10.37 | 97.9 | 8.1
2009-08-22 | 8:46 | 12.73 | 67 | 44 | 10.38 | 97.9 | 8.11
2009-08-22 | 9:01 | 12.76 | 67 | 44 | 10.45 | 98.6 | 8.12
2009-08-22 | 9:16 | 12.79 | 67 | 44 | 10.46 | 98.8 | 8.13
2009-08-22 | 9:31 | 12.83 | 67 | 44 | 10.43 | 98.7 | 8.14
2009-08-22 | 9:46 | 12.89 | 67 | 44 | 10.47 | 99.1 | 8.15
2009-08-22 | 10:01 | 12.98 | 67 | 44 | 10.49 | 99.5 | 8.16
2009-08-22 | 10:16 | 13.06 | 67 | 44 | 10.48 | 99.6 | 8.16
2009-08-22 | 10:31 | 13.15 | 67 | 43 | 10.45 | 99.6 | 8.18
2009-08-22 | 10:46 | 13.27 | 66 | 43 | 10.47 | 100 | 8.18
2009-08-22 | 11:01 | 13.4 | 66 | 43 | 10.47 | 100.2 | 8.19
2009-08-22 | 11:16 | 13.58 | 66 | 43 | 10.41 | 100.1 | 8.2
2009-08-22 | 11:31 | 13.81 | 66 | 43 | 10.4 | 100.5 | 8.2
2009-08-22 | 11:46 | 14.11 | 65 | 43 | 10.35 | 100.6 | 8.21
2009-08-22 | 12:01 | 14.48 | 65 | 42 | 10.24 | 100.5 | 8.23
2009-08-22 | 12:16 | 14.91 | 65 | 42 | 10.15 | 100.5 | 8.23
2009-08-22 | 12:31 | 15.37 | 65 | 42 | 10.04 | 100.3 | 8.24
2009-08-22 | 12:46 | 15.85 | 64 | 42 | 9.9 | 100 | 8.26
2009-08-22 | 13:01 | 16.35 | 64 | 42 | 9.78 | 99.8 | 8.27
2009-08-22 | 13:16 | 16.87 | 64 | 42 | 9.69 | 100 | 8.29
2009-08-22 | 13:31 | 17.39 | 64 | 42 | 9.56 | 99.7 | 8.3
2009-08-22 | 13:46 | 17.91 | 64 | 42 | 9.42 | 99.3 | 8.31
2009-08-22 | 14:01 | 18.44 | 64 | 42 | 9.34 | 99.6 | 8.33
2009-08-22 | 14:16 | 18.98 | 64 | 41 | 9.21 | 99.2 | 8.35
2009-08-22 | 14:31 | 19.47 | 64 | 41 | 9.11 | 99.2 | 8.37
2009-08-22 | 14:46 | 19.89 | 64 | 41 | 8.99 | 98.6 | 8.37
2009-08-22 | 15:01 | 20.27 | 63 | 41 | 8.9 | 98.4 | 8.38
2009-08-22 | 15:16 | 20.62 | 63 | 41 | 8.84 | 98.4 | 8.39
2009-08-22 | 15:31 | 20.89 | 63 | 41 | 8.77 | 98.2 | 8.38
2009-08-22 | 15:46 | 21.13 | 63 | 41 | 8.72 | 98.1 | 8.4
2009-08-22 | 16:01 | 21.28 | 63 | 41 | 8.67 | 97.8 | 8.4
2009-08-22 | 16:16 | 21.33 | 63 | 41 | 8.65 | 97.7 | 8.39
2009-08-22 | 16:31 | 21.29 | 63 | 41 | 8.63 | 97.4 | 8.38
2009-08-22 | 16:46 | 21.2 | 63 | 41 | 8.59 | 96.7 | 8.35
2009-08-22 | 17:01 | 21.03 | 63 | 41 | 8.61 | 96.6 | 8.33
2009-08-22 | 17:16 | 20.82 | 63 | 41 | 8.59 | 96 | 8.3
2009-08-22 | 17:31 | 20.62 | 63 | 41 | 8.59 | 95.6 | 8.27
2009-08-22 | 17:46 | 20.44 | 63 | 41 | 8.59 | 95.3 | 8.24
2009-08-22 | 18:01 | 20.27 | 63 | 41 | 8.62 | 95.4 | 8.21
2009-08-22 | 18:16 | 20.12 | 63 | 41 | 8.61 | 94.9 | 8.19
2009-08-22 | 18:31 | 19.99 | 63 | 41 | 8.64 | 95 | 8.16
2009-08-22 | 18:46 | 19.87 | 63 | 41 | 8.66 | 94.9 | 8.14
2009-08-22 | 19:01 | 19.76 | 63 | 41 | 8.67 | 94.9 | 8.13
2009-08-22 | 19:16 | 19.65 | 63 | 41 | 8.68 | 94.8 | 8.12
2009-08-22 | 19:31 | 19.57 | 63 | 41 | 8.65 | 94.4 | 8.11
2009-08-22 | 19:46 | 19.52 | 63 | 41 | 8.66 | 94.4 | 8.1
2009-08-22 | 20:01 | 19.48 | 63 | 41 | 8.66 | 94.3 | 8.1
2009-08-22 | 20:16 | 19.44 | 63 | 41 | 8.67 | 94.3 | 8.08
2009-08-22 | 20:31 | 19.38 | 63 | 41 | 8.67 | 94.2 | 8.07
2009-08-22 | 20:46 | 19.27 | 63 | 41 | 8.69 | 94.2 | 8.06
2009-08-22 | 21:01 | 19.14 | 64 | 41 | 8.72 | 94.3 | 8.05
2009-08-22 | 21:16 | 19 | 64 | 41 | 8.74 | 94.3 | 8.05
2009-08-22 | 21:31 | 18.85 | 64 | 41 | 8.78 | 94.4 | 8.05
2009-08-22 | 21:46 | 18.68 | 64 | 41 | 8.81 | 94.4 | 8.04
2009-08-22 | 22:01 | 18.5 | 64 | 41 | 8.85 | 94.5 | 8.04
2009-08-22 | 22:16 | 18.31 | 64 | 41 | 8.91 | 94.7 | 8.04
2009-08-22 | 22:31 | 18.13 | 64 | 41 | 8.95 | 94.8 | 8.04
2009-08-22 | 22:46 | 17.96 | 64 | 42 | 8.98 | 94.8 | 8.04
2009-08-22 | 23:01 | 17.78 | 64 | 42 | 8.99 | 94.6 | 8.04
2009-08-22 | 23:16 | 17.59 | 64 | 41 | 9.02 | 94.5 | 8.04
2009-08-22 | 23:31 | 17.4 | 64 | 42 | 9.09 | 94.8 | 8.04
2009-08-22 | 23:46 | 17.2 | 64 | 42 | 9.14 | 95 | 8.04
2009-08-23 | 0:01 | 17.01 | 64 | 42 | 9.18 | 95 | 8.04
2009-08-23 | 0:16 | 16.82 | 64 | 42 | 9.2 | 94.8 | 8.04
2009-08-23 | 0:31 | 16.64 | 64 | 42 | 9.26 | 95.1 | 8.04
2009-08-23 | 0:46 | 16.46 | 64 | 42 | 9.31 | 95.2 | 8.04
2009-08-23 | 1:01 | 16.29 | 64 | 42 | 9.35 | 95.3 | 8.04
2009-08-23 | 1:16 | 16.13 | 64 | 42 | 9.39 | 95.4 | 8.04
2009-08-23 | 1:31 | 15.98 | 64 | 42 | 9.4 | 95.2 | 8.04
2009-08-23 | 1:46 | 15.83 | 65 | 42 | 9.45 | 95.4 | 8.05
2009-08-23 | 2:01 | 15.68 | 65 | 42 | 9.49 | 95.6 | 8.04
2009-08-23 | 2:16 | 15.53 | 65 | 42 | 9.51 | 95.4 | 8.04
2009-08-23 | 2:31 | 15.4 | 65 | 42 | 9.55 | 95.5 | 8.04
2009-08-23 | 2:46 | 15.27 | 65 | 42 | 9.55 | 95.3 | 8.05
2009-08-23 | 3:01 | 15.15 | 65 | 42 | 9.61 | 95.7 | 8.04
2009-08-23 | 3:16 | 15.03 | 65 | 42 | 9.62 | 95.5 | 8.05
2009-08-23 | 3:31 | 14.92 | 65 | 42 | 9.68 | 95.8 | 8.05
2009-08-23 | 3:46 | 14.81 | 65 | 42 | 9.72 | 96 | 8.05
2009-08-23 | 4:01 | 14.68 | 65 | 43 | 9.74 | 95.9 | 8.05
2009-08-23 | 4:16 | 14.57 | 66 | 43 | 9.75 | 95.7 | 8.05
2009-08-23 | 4:31 | 14.45 | 66 | 43 | 9.76 | 95.6 | 8.05
2009-08-23 | 4:46 | 14.32 | 66 | 43 | 9.83 | 96 | 8.05
2009-08-23 | 5:01 | 14.18 | 66 | 43 | 9.84 | 95.9 | 8.06
2009-08-23 | 5:16 | 14.03 | 66 | 43 | 9.9 | 96.1 | 8.05
2009-08-23 | 5:31 | 13.88 | 67 | 43 | 9.97 | 96.5 | 8.06
2009-08-23 | 5:46 | 13.73 | 67 | 43 | 10.01 | 96.6 | 8.06
2009-08-23 | 6:01 | 13.57 | 67 | 44 | 10.04 | 96.5 | 8.06
2009-08-23 | 6:16 | 13.42 | 67 | 44 | 10.06 | 96.4 | 8.06
2009-08-23 | 6:31 | 13.27 | 67 | 44 | 10.13 | 96.7 | 8.06
2009-08-23 | 6:46 | 13.11 | 68 | 44 | 10.2 | 97.1 | 8.07
2009-08-23 | 7:01 | 12.98 | 68 | 44 | 10.24 | 97.1 | 8.08
2009-08-23 | 7:16 | 12.85 | 68 | 44 | 10.31 | 97.5 | 8.09
2009-08-23 | 7:31 | 12.72 | 68 | 44 | 10.35 | 97.6 | 8.09
2009-08-23 | 7:46 | 12.62 | 68 | 44 | 10.36 | 97.5 | 8.1
2009-08-23 | 8:01 | 12.53 | 69 | 45 | 10.43 | 97.9 | 8.11
2009-08-23 | 8:16 | 12.45 | 69 | 45 | 10.47 | 98.2 | 8.12
2009-08-23 | 8:31 | 12.4 | 69 | 45 | 10.49 | 98.3 | 8.13
2009-08-23 | 8:46 | 12.4 | 69 | 45 | 10.52 | 98.5 | 8.14
2009-08-23 | 9:01 | 12.43 | 69 | 45 | 10.53 | 98.7 | 8.15
2009-08-23 | 9:16 | 12.45 | 69 | 45 | 10.54 | 98.8 | 8.16
2009-08-23 | 9:31 | 12.47 | 69 | 45 | 10.56 | 99 | 8.17
2009-08-23 | 9:46 | 12.51 | 70 | 45 | 10.56 | 99.1 | 8.19
2009-08-23 | 10:01 | 12.58 | 70 | 45 | 10.56 | 99.3 | 8.2
2009-08-23 | 10:16 | 12.64 | 70 | 45 | 10.57 | 99.5 | 8.2
2009-08-23 | 10:31 | 12.71 | 70 | 45 | 10.57 | 99.7 | 8.22
2009-08-23 | 10:46 | 12.8 | 70 | 45 | 10.61 | 100.3 | 8.22
2009-08-23 | 11:01 | 12.92 | 69 | 45 | 10.61 | 100.5 | 8.23
2009-08-23 | 11:16 | 13.08 | 69 | 45 | 10.59 | 100.7 | 8.25
2009-08-23 | 11:31 | 13.27 | 69 | 45 | 10.58 | 101 | 8.25
2009-08-23 | 11:46 | 13.55 | 69 | 45 | 10.49 | 100.8 | 8.25
2009-08-23 | 12:01 | 13.92 | 69 | 45 | 10.39 | 100.7 | 8.27
2009-08-23 | 12:16 | 14.33 | 69 | 45 | 10.31 | 100.7 | 8.27
2009-08-23 | 12:31 | 14.77 | 68 | 44 | 10.21 | 100.8 | 8.28
2009-08-23 | 12:46 | 15.23 | 68 | 44 | 10.11 | 100.8 | 8.3
2009-08-23 | 13:01 | 15.71 | 68 | 44 | 10 | 100.7 | 8.31
2009-08-23 | 13:16 | 16.2 | 68 | 44 | 9.85 | 100.2 | 8.32
2009-08-23 | 13:31 | 16.73 | 67 | 44 | 9.73 | 100.1 | 8.34
2009-08-23 | 13:46 | 17.22 | 67 | 44 | 9.6 | 99.8 | 8.35
2009-08-23 | 14:01 | 17.75 | 67 | 44 | 9.53 | 100.2 | 8.37
2009-08-23 | 14:16 | 18.25 | 67 | 44 | 9.39 | 99.7 | 8.39
2009-08-23 | 14:31 | 18.72 | 67 | 43 | 9.31 | 99.8 | 8.42
2009-08-23 | 14:46 | 19.13 | 67 | 43 | 9.19 | 99.3 | 8.42
2009-08-23 | 15:01 | 19.48 | 67 | 43 | 9.09 | 99 | 8.43
2009-08-23 | 15:16 | 19.8 | 67 | 43 | 9.04 | 99 | 8.44
2009-08-23 | 15:31 | 20.03 | 66 | 43 | 8.95 | 98.5 | 8.45
2009-08-23 | 15:46 | 20.29 | 66 | 43 | 8.9 | 98.5 | 8.45
2009-08-23 | 16:01 | 20.44 | 66 | 43 | 8.84 | 98.1 | 8.45
2009-08-23 | 16:16 | 20.48 | 66 | 43 | 8.8 | 97.7 | 8.44
2009-08-23 | 16:31 | 20.43 | 66 | 43 | 8.81 | 97.7 | 8.43
2009-08-23 | 16:46 | 20.32 | 66 | 43 | 8.78 | 97.2 | 8.41
2009-08-23 | 17:01 | 20.14 | 66 | 43 | 8.77 | 96.7 | 8.38
2009-08-23 | 17:16 | 19.91 | 66 | 43 | 8.83 | 96.9 | 8.35
2009-08-23 | 17:31 | 19.66 | 66 | 43 | 8.83 | 96.5 | 8.31
2009-08-23 | 17:46 | 19.44 | 66 | 43 | 8.83 | 96.1 | 8.29
2009-08-23 | 18:01 | 19.24 | 66 | 43 | 8.83 | 95.7 | 8.26
2009-08-23 | 18:16 | 19.05 | 66 | 43 | 8.85 | 95.5 | 8.23
2009-08-23 | 18:31 | 18.91 | 66 | 43 | 8.88 | 95.6 | 8.21
2009-08-23 | 18:46 | 18.78 | 66 | 43 | 8.88 | 95.3 | 8.2
2009-08-23 | 19:01 | 18.67 | 66 | 43 | 8.91 | 95.5 | 8.17
2009-08-23 | 19:16 | 18.57 | 67 | 43 | 8.9 | 95.1 | 8.16
2009-08-23 | 19:31 | 18.48 | 67 | 43 | 8.92 | 95.2 | 8.14
2009-08-23 | 19:46 | 18.43 | 67 | 43 | 8.89 | 94.7 | 8.13
2009-08-23 | 20:01 | 18.36 | 67 | 43 | 8.88 | 94.5 | 8.12
2009-08-23 | 20:16 | 18.27 | 67 | 43 | 8.88 | 94.4 | 8.11
2009-08-23 | 20:31 | 18.16 | 67 | 43 | 8.93 | 94.6 | 8.1
2009-08-23 | 20:46 | 18.05 | 67 | 43 | 8.92 | 94.3 | 8.09
2009-08-23 | 21:01 | 17.91 | 67 | 43 | 8.95 | 94.4 | 8.08
2009-08-23 | 21:16 | 17.75 | 67 | 43 | 9 | 94.6 | 8.09
2009-08-23 | 21:31 | 17.59 | 67 | 43 | 9.05 | 94.8 | 8.08
2009-08-23 | 21:46 | 17.4 | 67 | 43 | 9.09 | 94.9 | 8.07
2009-08-23 | 22:01 | 17.21 | 67 | 43 | 9.14 | 95 | 8.07
2009-08-23 | 22:16 | 16.99 | 67 | 43 | 9.18 | 95 | 8.07
2009-08-23 | 22:31 | 16.78 | 67 | 44 | 9.23 | 95.1 | 8.07
2009-08-23 | 22:46 | 16.55 | 67 | 44 | 9.3 | 95.3 | 8.06
2009-08-23 | 23:01 | 16.32 | 67 | 44 | 9.32 | 95 | 8.07
2009-08-23 | 23:16 | 16.09 | 67 | 44 | 9.39 | 95.3 | 8.06
2009-08-23 | 23:31 | 15.85 | 67 | 44 | 9.46 | 95.5 | 8.07
2009-08-23 | 23:46 | 15.61 | 67 | 44 | 9.54 | 95.8 | 8.06
2009-08-24 | 0:01 | 15.38 | 67 | 44 | 9.55 | 95.5 | 8.06
2009-08-24 | 0:16 | 15.16 | 67 | 44 | 9.64 | 95.9 | 8.06
2009-08-24 | 0:31 | 14.93 | 67 | 43 | 9.68 | 95.9 | 8.06
2009-08-24 | 0:46 | 14.72 | 67 | 44 | 9.72 | 95.8 | 8.06
2009-08-24 | 1:01 | 14.52 | 67 | 44 | 9.76 | 95.8 | 8.06
2009-08-24 | 1:16 | 14.32 | 67 | 44 | 9.84 | 96.1 | 8.06
2009-08-24 | 1:31 | 14.13 | 67 | 44 | 9.89 | 96.2 | 8.06
2009-08-24 | 1:46 | 13.97 | 67 | 44 | 9.91 | 96.1 | 8.06
2009-08-24 | 2:01 | 13.79 | 67 | 44 | 9.98 | 96.4 | 8.05
2009-08-24 | 2:16 | 13.63 | 67 | 44 | 10 | 96.3 | 8.06
2009-08-24 | 2:31 | 13.47 | 67 | 44 | 10.07 | 96.5 | 8.06
2009-08-24 | 2:46 | 13.32 | 67 | 44 | 10.08 | 96.4 | 8.06
2009-08-24 | 3:01 | 13.17 | 67 | 44 | 10.11 | 96.3 | 8.05
2009-08-24 | 3:16 | 13.02 | 67 | 44 | 10.14 | 96.3 | 8.05
2009-08-24 | 3:31 | 12.87 | 67 | 44 | 10.22 | 96.8 | 8.06
2009-08-24 | 3:46 | 12.73 | 67 | 44 | 10.26 | 96.8 | 8.05
2009-08-24 | 4:01 | 12.59 | 67 | 44 | 10.3 | 96.9 | 8.05
2009-08-24 | 4:16 | 12.44 | 68 | 44 | 10.35 | 97 | 8.05
2009-08-24 | 4:31 | 12.31 | 68 | 44 | 10.36 | 96.9 | 8.06
2009-08-24 | 4:46 | 12.18 | 68 | 44 | 10.42 | 97.1 | 8.06
2009-08-24 | 5:01 | 12.05 | 68 | 44 | 10.47 | 97.2 | 8.05
2009-08-24 | 5:16 | 11.91 | 68 | 44 | 10.49 | 97.1 | 8.05
2009-08-24 | 5:31 | 11.77 | 68 | 44 | 10.54 | 97.3 | 8.05
2009-08-24 | 5:46 | 11.64 | 68 | 45 | 10.59 | 97.5 | 8.06
2009-08-24 | 6:01 | 11.54 | 69 | 45 | 10.62 | 97.5 | 8.06
2009-08-24 | 6:16 | 11.43 | 69 | 45 | 10.65 | 97.5 | 8.05
2009-08-24 | 6:31 | 11.33 | 69 | 45 | 10.66 | 97.4 | 8.06
2009-08-24 | 6:46 | 11.23 | 69 | 45 | 10.69 | 97.5 | 8.07
2009-08-24 | 7:01 | 11.14 | 69 | 45 | 10.73 | 97.6 | 8.07
2009-08-24 | 7:16 | 11.06 | 70 | 45 | 10.8 | 98.1 | 8.08
2009-08-24 | 7:31 | 10.98 | 70 | 45 | 10.85 | 98.4 | 8.08
2009-08-24 | 7:46 | 10.91 | 70 | 46 | 10.87 | 98.4 | 8.1
2009-08-24 | 8:01 | 10.86 | 70 | 46 | 10.92 | 98.7 | 8.1
2009-08-24 | 8:16 | 10.82 | 70 | 46 | 10.95 | 98.9 | 8.11
2009-08-24 | 8:31 | 10.81 | 71 | 46 | 10.95 | 98.9 | 8.12
2009-08-24 | 8:46 | 10.81 | 71 | 46 | 10.96 | 98.9 | 8.13
2009-08-24 | 9:01 | 10.85 | 71 | 46 | 11.02 | 99.6 | 8.14
2009-08-24 | 9:16 | 10.89 | 71 | 46 | 11 | 99.5 | 8.15
2009-08-24 | 9:31 | 10.94 | 71 | 46 | 11 | 99.6 | 8.16
2009-08-24 | 9:46 | 11.01 | 71 | 46 | 11.02 | 100 | 8.17
2009-08-24 | 10:01 | 11.11 | 72 | 47 | 11.04 | 100.4 | 8.19
2009-08-24 | 10:16 | 11.21 | 72 | 47 | 11.04 | 100.6 | 8.19
2009-08-24 | 10:31 | 11.31 | 72 | 47 | 11.01 | 100.6 | 8.2
2009-08-24 | 10:46 | 11.44 | 72 | 47 | 11.02 | 101 | 8.21
2009-08-24 | 11:01 | 11.57 | 72 | 47 | 11.01 | 101.2 | 8.21
2009-08-24 | 11:16 | 11.77 | 72 | 47 | 10.98 | 101.4 | 8.22
2009-08-24 | 11:31 | 12 | 72 | 47 | 10.91 | 101.2 | 8.23
2009-08-24 | 11:46 | 12.28 | 72 | 47 | 10.87 | 101.5 | 8.24
2009-08-24 | 12:01 | 12.66 | 71 | 46 | 10.78 | 101.6 | 8.24
2009-08-24 | 12:16 | 13.1 | 71 | 46 | 10.63 | 101.1 | 8.27
2009-08-24 | 12:31 | 13.58 | 71 | 46 | 10.5 | 100.9 | 8.27