## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB audited against Loner Pro Series, Sonde AA, NIST
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 11:28 |  | 11:28 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.04 |  | A
DO | Winkler titration of surface water |  | 9.85 |  | 10.21 |  | B
SpC | 148 µS/mL |  | 54.6 |  | 55 |  | A
Ph | pH 7 standard solution |  | 8.23 |  | 8.38 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-05-19 13:38 | 12.1 | 7.61 | 10.48 | 97.5 | 33 | 22.1
Sonde | 2009-05-19 13:40 | 12.16 | 7.7 | 10.91 | 101.6 | 33 | 22
Grade |  | A | A | B |  | A
Audit 2 | 2009-05-22 13:06 | 13.1 | 7.78 | 10.67 | 101.5 | 46.5 | 29
Sonde | 2009-05-22 13:15 | 13.35 | 7.88 | 11.16 | 106.8 | 44 | 29
Grade |  | A | A | B |  | A
Audit 3 | 2009-05-31 12:04 | 9.7 | 7.51 | 11.32 | 99.6 | 29.8 | 18.8
Sonde | 2009-05-31 12:00 | 9.68 | 7.56 | 11.96 | 105.2 | 27 | 18
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB audited against Loner Pro Series, Sonde AA, NIST
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-31 |  | 2009-05-31 |  | A
Time | Computer set to atomic time |  | 12:33 |  | 12:40 |  | A
Temp | VWR NIST traceable surface water |  | 10.1 |  | 9.95 |  | A
DO | In Situ |  | 11.19 |  | 11.82 |  | B
SpC | In SItu |  | 29.6 |  | 28 |  | A
Ph | In SItu |  | 7.47 |  | 7.63 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner Pro Series
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 11:24 |  | 11:24 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17 |  | A
DO | Winkler titration of surface water |  | 10.09 |  | 9.85 |  | A
SpC | 148 µS/mL |  | 53 |  | 54.6 |  | A
Ph | pH 7 standard solution |  | 8.4 |  | 8.23 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-19 13:38 | 12.1 | 7.61 | 10.48 | 97.5 | 33 | 22.1
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 13:45 |  | 13:45 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.2 |  | A
DO | Winkler titration of surface water |  | 9.9 |  | 10.42 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 228 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
I trust Sonde SpC data despite failing Pro-Series Audit |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner Pro Series
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-19 |  | 2009-05-19 |  | A
Time | Computer set to atomic time |  | 13:45 |  | 13:45 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.2 |  | A
DO | Winkler titration of surface water |  | 9.9 |  | 10.42 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 228 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-22 13:06 | 13.1 | 7.78 | 10.67 | 101.5 | 46.5 | 29
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-22 |  | 2009-05-22 |  | A
Time | Computer set to atomic time |  | 17:59 |  | 17:59 |  | A
Temp | VWR NIST traceable surface water |  | 16.5 |  | 16.4 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.48 |  | A
SpC | 148 µS/mL |  | 153.6 |  | 168 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
I trust Sonde SpC data despite failing Pro-Series Audit |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-26 |  | 2009-05-26 |  | A
Time | Computer set to atomic time |  | 12:41 |  | 12:41 |  | A
Temp | VWR NIST traceable surface water |  | 21.5 |  | 21.3 |  | A
DO | Winkler titration of surface water |  | 9.3 |  | 8.79 |  | B
SpC | 148 µS/mL |  | 153.6 |  | 153.9 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-05-31 12:04 | 9.7 | 7.51 | 11.32 | 99.6 | 29.8 | 18.8
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-31 |  | 2009-05-31 |  | A
Time | Computer set to atomic time |  | 15:15 |  | 15:15 |  | A
Temp | VWR NIST traceable surface water |  | 19.4 |  | 19.3 |  | A
DO | Winkler titration of surface water |  | 7.9 |  | 8.03 |  | A
SpC | 148 µS/mL |  | 153.6 |  | 177.1 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
I trust Sonde SpC data despite failing Pro-Series Audit |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | SpCond | TDS | DO Conc | DO% | pH
M/D/Y |  | °C | uS/cm | mg/L | mg/L | %
2009-05-19 | 13:30 | 12.01 | 33 | 22 | 12.2 | 113.2 | 7.7
2009-05-19 | 13:45 | 12.26 | 33 | 22 | 11.22 | 104.7 | 7.7
2009-05-19 | 14:00 | 12.51 | 33 | 22 | 11.39 | 106.9 | 7.71
2009-05-19 | 14:15 | 12.75 | 33 | 22 | 11.29 | 106.6 | 7.72
2009-05-19 | 14:30 | 12.93 | 33 | 22 | 11.17 | 105.9 | 7.73
2009-05-19 | 14:45 | 13.08 | 33 | 22 | 11.01 | 104.7 | 7.72
2009-05-19 | 15:00 | 13.23 | 33 | 22 | 10.94 | 104.4 | 7.73
2009-05-19 | 15:15 | 13.33 | 33 | 22 | 10.87 | 103.9 | 7.73
2009-05-19 | 15:30 | 13.4 | 33 | 22 | 10.8 | 103.4 | 7.72
2009-05-19 | 15:45 | 13.44 | 33 | 22 | 10.73 | 102.8 | 7.72
2009-05-19 | 16:00 | 13.49 | 33 | 22 | 10.67 | 102.4 | 7.72
2009-05-19 | 16:15 | 13.52 | 33 | 22 | 10.63 | 102.1 | 7.72
2009-05-19 | 16:30 | 13.52 | 33 | 22 | 10.6 | 101.8 | 7.72
2009-05-19 | 16:45 | 13.45 | 33 | 22 | 10.58 | 101.5 | 7.72
2009-05-19 | 17:00 | 13.43 | 33 | 22 | 10.52 | 100.8 | 7.72
2009-05-19 | 17:15 | 13.34 | 33 | 21 | 10.53 | 100.8 | 7.71
2009-05-19 | 17:30 | 13.23 | 33 | 21 | 10.55 | 100.6 | 7.69
2009-05-19 | 17:45 | 13.11 | 33 | 21 | 10.57 | 100.6 | 7.69
2009-05-19 | 18:00 | 13 | 32 | 21 | 10.53 | 100 | 7.69
2009-05-19 | 18:15 | 12.89 | 32 | 21 | 10.57 | 100.1 | 7.68
2009-05-19 | 18:30 | 12.74 | 32 | 21 | 10.61 | 100.1 | 7.67
2009-05-19 | 18:45 | 12.58 | 33 | 21 | 10.6 | 99.7 | 7.67
2009-05-19 | 19:00 | 12.41 | 33 | 21 | 10.63 | 99.5 | 7.67
2009-05-19 | 19:15 | 12.25 | 33 | 21 | 10.65 | 99.4 | 7.66
2009-05-19 | 19:30 | 12.09 | 33 | 21 | 10.73 | 99.8 | 7.66
2009-05-19 | 19:45 | 11.9 | 33 | 21 | 10.8 | 100 | 7.66
2009-05-19 | 20:00 | 11.74 | 33 | 21 | 10.79 | 99.5 | 7.66
2009-05-19 | 20:15 | 11.6 | 33 | 21 | 10.8 | 99.3 | 7.66
2009-05-19 | 20:30 | 11.45 | 33 | 21 | 10.87 | 99.6 | 7.65
2009-05-19 | 20:45 | 11.33 | 33 | 21 | 10.86 | 99.3 | 7.65
2009-05-19 | 21:00 | 11.22 | 33 | 22 | 10.92 | 99.6 | 7.64
2009-05-19 | 21:15 | 11.11 | 33 | 22 | 10.96 | 99.7 | 7.64
2009-05-19 | 21:30 | 10.99 | 33 | 22 | 10.95 | 99.3 | 7.64
2009-05-19 | 21:45 | 10.88 | 33 | 22 | 11.02 | 99.6 | 7.64
2009-05-19 | 22:00 | 10.78 | 33 | 22 | 11.01 | 99.3 | 7.64
2009-05-19 | 22:15 | 10.67 | 33 | 22 | 11.02 | 99.2 | 7.64
2009-05-19 | 22:30 | 10.54 | 33 | 22 | 11.1 | 99.6 | 7.63
2009-05-19 | 22:45 | 10.39 | 33 | 22 | 11.17 | 99.9 | 7.63
2009-05-19 | 23:00 | 10.22 | 33 | 22 | 11.23 | 100 | 7.63
2009-05-19 | 23:15 | 10.05 | 33 | 22 | 11.29 | 100.1 | 7.63
2009-05-19 | 23:30 | 9.87 | 33 | 22 | 11.35 | 100.2 | 7.63
2009-05-19 | 23:45 | 9.68 | 33 | 22 | 11.41 | 100.3 | 7.62
2009-05-20 | 0:00 | 9.51 | 33 | 22 | 11.41 | 100 | 7.62
2009-05-20 | 0:15 | 9.35 | 33 | 22 | 11.44 | 99.8 | 7.62
2009-05-20 | 0:30 | 9.2 | 33 | 22 | 11.48 | 99.8 | 7.62
2009-05-20 | 0:45 | 9.04 | 33 | 22 | 11.57 | 100.2 | 7.62
2009-05-20 | 1:01 | 8.87 | 33 | 22 | 11.65 | 100.4 | 7.61
2009-05-20 | 1:16 | 8.7 | 33 | 22 | 11.71 | 100.6 | 7.6
2009-05-20 | 1:31 | 8.56 | 33 | 22 | 11.71 | 100.2 | 7.6
2009-05-20 | 1:46 | 8.42 | 33 | 22 | 11.74 | 100.2 | 7.6
2009-05-20 | 2:01 | 8.26 | 33 | 22 | 11.79 | 100.1 | 7.59
2009-05-20 | 2:16 | 8.09 | 33 | 22 | 11.84 | 100.2 | 7.59
2009-05-20 | 2:31 | 7.91 | 33 | 22 | 11.95 | 100.7 | 7.59
2009-05-20 | 2:46 | 7.72 | 33 | 22 | 11.99 | 100.5 | 7.59
2009-05-20 | 3:01 | 7.51 | 33 | 22 | 12.05 | 100.5 | 7.58
2009-05-20 | 3:16 | 7.3 | 33 | 22 | 12.16 | 101 | 7.58
2009-05-20 | 3:31 | 7.1 | 33 | 22 | 12.26 | 101.3 | 7.58
2009-05-20 | 3:46 | 6.9 | 33 | 22 | 12.29 | 101 | 7.58
2009-05-20 | 4:01 | 6.71 | 33 | 22 | 12.34 | 100.9 | 7.57
2009-05-20 | 4:16 | 6.54 | 33 | 22 | 12.39 | 100.9 | 7.57
2009-05-20 | 4:31 | 6.37 | 33 | 22 | 12.45 | 101 | 7.57
2009-05-20 | 4:46 | 6.21 | 33 | 22 | 12.56 | 101.4 | 7.57
2009-05-20 | 5:01 | 6.06 | 33 | 22 | 12.64 | 101.7 | 7.57
2009-05-20 | 5:16 | 5.92 | 33 | 22 | 12.71 | 101.9 | 7.56
2009-05-20 | 5:31 | 5.79 | 33 | 22 | 12.77 | 102 | 7.56
2009-05-20 | 5:46 | 5.66 | 33 | 22 | 12.83 | 102.2 | 7.56
2009-05-20 | 6:01 | 5.53 | 33 | 22 | 12.83 | 101.9 | 7.56
2009-05-20 | 6:16 | 5.41 | 33 | 22 | 12.87 | 101.8 | 7.57
2009-05-20 | 6:31 | 5.29 | 34 | 22 | 12.98 | 102.4 | 7.57
2009-05-20 | 6:46 | 5.2 | 34 | 22 | 13.05 | 102.7 | 7.57
2009-05-20 | 7:01 | 5.12 | 34 | 22 | 13.06 | 102.6 | 7.57
2009-05-20 | 7:16 | 5.05 | 34 | 22 | 13.15 | 103.1 | 7.58
2009-05-20 | 7:31 | 5 | 34 | 22 | 13.15 | 103 | 7.59
2009-05-20 | 7:46 | 4.97 | 34 | 22 | 13.23 | 103.5 | 7.59
2009-05-20 | 8:01 | 4.94 | 34 | 22 | 13.29 | 103.9 | 7.59
2009-05-20 | 8:16 | 4.92 | 34 | 22 | 13.28 | 103.8 | 7.6
2009-05-20 | 8:31 | 4.9 | 34 | 22 | 13.39 | 104.6 | 7.6
2009-05-20 | 8:46 | 4.9 | 34 | 22 | 13.49 | 105.4 | 7.61
2009-05-20 | 9:01 | 4.91 | 34 | 22 | 13.55 | 105.9 | 7.61
2009-05-20 | 9:16 | 4.95 | 34 | 22 | 13.58 | 106.2 | 7.62
2009-05-20 | 9:31 | 5.04 | 34 | 22 | 13.56 | 106.3 | 7.62
2009-05-20 | 9:46 | 5.16 | 34 | 22 | 13.44 | 105.7 | 7.63
2009-05-20 | 10:01 | 5.33 | 34 | 22 | 13.4 | 105.8 | 7.63
2009-05-20 | 10:16 | 5.54 | 35 | 22 | 13.33 | 105.9 | 7.64
2009-05-20 | 10:31 | 5.79 | 35 | 22 | 13.25 | 105.8 | 7.65
2009-05-20 | 10:46 | 6.07 | 35 | 23 | 13.19 | 106.1 | 7.65
2009-05-20 | 11:01 | 6.35 | 35 | 23 | 13.17 | 106.8 | 7.66
2009-05-20 | 11:16 | 6.66 | 35 | 23 | 13.08 | 106.8 | 7.66
2009-05-20 | 11:31 | 6.98 | 35 | 23 | 13.05 | 107.4 | 7.67
2009-05-20 | 11:46 | 7.32 | 35 | 23 | 12.92 | 107.3 | 7.67
2009-05-20 | 12:01 | 7.68 | 35 | 23 | 12.86 | 107.7 | 7.68
2009-05-20 | 12:16 | 8.04 | 35 | 23 | 12.77 | 107.9 | 7.68
2009-05-20 | 12:31 | 8.41 | 35 | 23 | 12.66 | 108 | 7.69
2009-05-20 | 12:46 | 8.79 | 35 | 23 | 12.54 | 107.9 | 7.7
2009-05-20 | 13:00 | 9.19 | 36 | 23 | 12.35 | 107.3 | 7.7
2009-05-20 | 13:15 | 9.59 | 36 | 23 | 12.19 | 107 | 7.71
2009-05-20 | 13:30 | 9.98 | 36 | 23 | 12.1 | 107.2 | 7.72
2009-05-20 | 13:45 | 10.36 | 36 | 23 | 11.99 | 107.1 | 7.73
2009-05-20 | 14:00 | 10.7 | 36 | 23 | 11.88 | 107 | 7.73
2009-05-20 | 14:15 | 11.03 | 36 | 23 | 11.76 | 106.7 | 7.74
2009-05-20 | 14:30 | 11.31 | 36 | 23 | 11.6 | 105.9 | 7.75
2009-05-20 | 14:45 | 11.58 | 36 | 24 | 11.52 | 105.8 | 7.75
2009-05-20 | 15:00 | 11.81 | 36 | 24 | 11.43 | 105.6 | 7.75
2009-05-20 | 15:15 | 12 | 36 | 24 | 11.34 | 105.3 | 7.76
2009-05-20 | 15:30 | 12.16 | 37 | 24 | 11.26 | 104.9 | 7.76
2009-05-20 | 15:45 | 12.31 | 37 | 24 | 11.19 | 104.5 | 7.76
2009-05-20 | 16:00 | 12.42 | 37 | 24 | 11.07 | 103.7 | 7.77
2009-05-20 | 16:15 | 12.51 | 37 | 24 | 11.05 | 103.7 | 7.77
2009-05-20 | 16:30 | 12.58 | 37 | 24 | 10.96 | 103 | 7.77
2009-05-20 | 16:45 | 12.61 | 37 | 24 | 10.95 | 103.1 | 7.77
2009-05-20 | 17:00 | 12.63 | 37 | 24 | 10.93 | 102.9 | 7.78
2009-05-20 | 17:15 | 12.62 | 37 | 24 | 10.91 | 102.7 | 7.77
2009-05-20 | 17:30 | 12.62 | 37 | 24 | 10.89 | 102.5 | 7.77
2009-05-20 | 17:45 | 12.59 | 37 | 24 | 10.83 | 101.8 | 7.77
2009-05-20 | 18:00 | 12.56 | 38 | 24 | 10.84 | 101.9 | 7.76
2009-05-20 | 18:15 | 12.52 | 38 | 24 | 10.84 | 101.8 | 7.76
2009-05-20 | 18:30 | 12.49 | 38 | 25 | 10.84 | 101.7 | 7.76
2009-05-20 | 18:45 | 12.46 | 38 | 25 | 10.83 | 101.6 | 7.75
2009-05-20 | 19:00 | 12.43 | 38 | 25 | 10.78 | 101 | 7.75
2009-05-20 | 19:15 | 12.39 | 38 | 25 | 10.81 | 101.2 | 7.75
2009-05-20 | 19:30 | 12.32 | 38 | 25 | 10.83 | 101.2 | 7.74
2009-05-20 | 19:45 | 12.24 | 38 | 25 | 10.79 | 100.7 | 7.74
2009-05-20 | 20:00 | 12.14 | 38 | 25 | 10.79 | 100.4 | 7.74
2009-05-20 | 20:15 | 12.02 | 38 | 25 | 10.84 | 100.7 | 7.73
2009-05-20 | 20:30 | 11.87 | 39 | 25 | 10.89 | 100.7 | 7.72
2009-05-20 | 20:45 | 11.72 | 39 | 25 | 10.92 | 100.7 | 7.71
2009-05-20 | 21:00 | 11.58 | 39 | 25 | 10.9 | 100.1 | 7.71
2009-05-20 | 21:15 | 11.44 | 39 | 25 | 10.95 | 100.4 | 7.71
2009-05-20 | 21:30 | 11.31 | 39 | 25 | 11 | 100.4 | 7.71
2009-05-20 | 21:45 | 11.19 | 39 | 25 | 10.98 | 100 | 7.7
2009-05-20 | 22:00 | 11.07 | 39 | 25 | 10.99 | 99.9 | 7.7
2009-05-20 | 22:15 | 10.96 | 39 | 25 | 11.06 | 100.2 | 7.7
2009-05-20 | 22:30 | 10.86 | 39 | 26 | 11.1 | 100.4 | 7.7
2009-05-20 | 22:45 | 10.74 | 39 | 26 | 11.1 | 100 | 7.7
2009-05-20 | 23:00 | 10.61 | 39 | 26 | 11.17 | 100.4 | 7.7
2009-05-20 | 23:15 | 10.49 | 40 | 26 | 11.22 | 100.5 | 7.69
2009-05-20 | 23:30 | 10.36 | 40 | 26 | 11.26 | 100.6 | 7.69
2009-05-20 | 23:45 | 10.24 | 40 | 26 | 11.3 | 100.7 | 7.69
2009-05-21 | 0:00 | 10.11 | 44 | 28 | 11.34 | 100.7 | 7.69
2009-05-21 | 0:15 | 9.95 | 43 | 28 | 11.39 | 100.8 | 7.7
2009-05-21 | 0:30 | 9.8 | 43 | 28 | 11.44 | 100.9 | 7.7
2009-05-21 | 0:45 | 9.65 | 43 | 28 | 11.49 | 100.9 | 7.69
2009-05-21 | 1:01 | 9.51 | 43 | 28 | 11.48 | 100.5 | 7.69
2009-05-21 | 1:16 | 9.36 | 43 | 28 | 11.56 | 100.9 | 7.69
2009-05-21 | 1:31 | 9.22 | 43 | 28 | 11.62 | 101 | 7.69
2009-05-21 | 1:46 | 9.07 | 43 | 28 | 11.67 | 101.1 | 7.69
2009-05-21 | 2:01 | 8.92 | 43 | 28 | 11.72 | 101.2 | 7.69
2009-05-21 | 2:16 | 8.78 | 44 | 28 | 11.77 | 101.3 | 7.69
2009-05-21 | 2:31 | 8.64 | 43 | 28 | 11.77 | 100.9 | 7.69
2009-05-21 | 2:46 | 8.5 | 43 | 28 | 11.85 | 101.3 | 7.69
2009-05-21 | 3:01 | 8.36 | 44 | 28 | 11.85 | 100.9 | 7.68
2009-05-21 | 3:16 | 8.21 | 43 | 28 | 11.94 | 101.3 | 7.68
2009-05-21 | 3:31 | 8.05 | 43 | 28 | 11.95 | 101.1 | 7.68
2009-05-21 | 3:46 | 7.9 | 43 | 28 | 11.98 | 101 | 7.68
2009-05-21 | 4:01 | 7.76 | 43 | 28 | 12.08 | 101.4 | 7.68
2009-05-21 | 4:16 | 7.61 | 43 | 28 | 12.15 | 101.6 | 7.67
2009-05-21 | 4:31 | 7.47 | 41 | 27 | 12.16 | 101.3 | 7.67
2009-05-21 | 4:46 | 7.34 | 41 | 27 | 12.24 | 101.7 | 7.67
2009-05-21 | 5:01 | 7.21 | 41 | 27 | 12.24 | 101.4 | 7.66
2009-05-21 | 5:16 | 7.08 | 41 | 27 | 12.33 | 101.8 | 7.66
2009-05-21 | 5:31 | 6.96 | 43 | 28 | 12.33 | 101.5 | 7.66
2009-05-21 | 5:46 | 6.84 | 41 | 27 | 12.36 | 101.4 | 7.66
2009-05-21 | 6:01 | 6.73 | 42 | 27 | 12.4 | 101.5 | 7.66
2009-05-21 | 6:16 | 6.62 | 42 | 27 | 12.5 | 102 | 7.66
2009-05-21 | 6:31 | 6.52 | 42 | 27 | 12.58 | 102.4 | 7.66
2009-05-21 | 6:46 | 6.44 | 42 | 27 | 12.65 | 102.7 | 7.66
2009-05-21 | 7:01 | 6.36 | 42 | 27 | 12.7 | 103 | 7.66
2009-05-21 | 7:16 | 6.3 | 42 | 27 | 12.76 | 103.3 | 7.67
2009-05-21 | 7:31 | 6.25 | 42 | 27 | 12.75 | 103.1 | 7.67
2009-05-21 | 7:46 | 6.21 | 42 | 27 | 12.76 | 103.1 | 7.68
2009-05-21 | 8:01 | 6.19 | 42 | 27 | 12.85 | 103.7 | 7.68
2009-05-21 | 8:16 | 6.16 | 42 | 27 | 12.92 | 104.2 | 7.68
2009-05-21 | 8:31 | 6.14 | 42 | 27 | 13.02 | 105 | 7.69
2009-05-21 | 8:46 | 6.14 | 41 | 27 | 13.07 | 105.4 | 7.69
2009-05-21 | 9:01 | 6.14 | 41 | 27 | 13.12 | 105.8 | 7.7
2009-05-21 | 9:16 | 6.18 | 42 | 27 | 13.21 | 106.6 | 7.7
2009-05-21 | 9:31 | 6.25 | 42 | 27 | 13.14 | 106.2 | 7.71
2009-05-21 | 9:46 | 6.36 | 42 | 27 | 13.12 | 106.3 | 7.72
2009-05-21 | 10:01 | 6.52 | 42 | 27 | 13.06 | 106.3 | 7.73
2009-05-21 | 10:16 | 6.75 | 42 | 27 | 12.97 | 106.2 | 7.73
2009-05-21 | 10:31 | 7.04 | 42 | 27 | 12.8 | 105.6 | 7.73
2009-05-21 | 10:46 | 7.35 | 42 | 27 | 12.72 | 105.7 | 7.74
2009-05-21 | 11:01 | 7.69 | 42 | 27 | 12.75 | 106.8 | 7.74
2009-05-21 | 11:16 | 8.03 | 42 | 27 | 12.72 | 107.5 | 7.75
2009-05-21 | 11:31 | 8.39 | 42 | 27 | 12.65 | 107.8 | 7.75
2009-05-21 | 11:46 | 8.77 | 42 | 27 | 12.56 | 108.1 | 7.76
2009-05-21 | 12:01 | 9.17 | 42 | 27 | 12.45 | 108.1 | 7.76
2009-05-21 | 12:16 | 9.56 | 42 | 27 | 12.33 | 108.1 | 7.77
2009-05-21 | 12:31 | 9.96 | 42 | 27 | 12.21 | 108.1 | 7.78
2009-05-21 | 12:46 | 10.38 | 42 | 27 | 12.08 | 108 | 7.78
2009-05-21 | 13:00 | 10.81 | 42 | 28 | 11.88 | 107.3 | 7.79
2009-05-21 | 13:15 | 11.24 | 42 | 27 | 11.77 | 107.4 | 7.8
2009-05-21 | 13:30 | 11.68 | 42 | 28 | 11.59 | 106.7 | 7.8
2009-05-21 | 13:45 | 12.11 | 42 | 28 | 11.43 | 106.3 | 7.82
2009-05-21 | 14:00 | 12.53 | 43 | 28 | 11.28 | 105.9 | 7.82
2009-05-21 | 14:15 | 12.92 | 43 | 28 | 11.14 | 105.5 | 7.83
2009-05-21 | 14:30 | 13.29 | 43 | 28 | 11 | 105.1 | 7.84
2009-05-21 | 14:45 | 13.62 | 43 | 28 | 10.88 | 104.7 | 7.84
2009-05-21 | 15:00 | 13.92 | 43 | 28 | 10.81 | 104.7 | 7.85
2009-05-21 | 15:15 | 14.16 | 43 | 28 | 10.73 | 104.5 | 7.85
2009-05-21 | 15:30 | 14.34 | 44 | 28 | 10.66 | 104.2 | 7.85
2009-05-21 | 15:45 | 14.49 | 44 | 28 | 10.59 | 103.9 | 7.85
2009-05-21 | 16:00 | 14.6 | 44 | 28 | 10.48 | 103.1 | 7.85
2009-05-21 | 16:15 | 14.68 | 44 | 29 | 10.42 | 102.6 | 7.85
2009-05-21 | 16:30 | 14.72 | 44 | 29 | 10.41 | 102.6 | 7.85
2009-05-21 | 16:45 | 14.74 | 44 | 29 | 10.34 | 102 | 7.85
2009-05-21 | 17:00 | 14.73 | 44 | 28 | 10.35 | 102.1 | 7.85
2009-05-21 | 17:15 | 14.68 | 44 | 29 | 10.3 | 101.4 | 7.85
2009-05-21 | 17:30 | 14.66 | 44 | 29 | 10.31 | 101.5 | 7.85
2009-05-21 | 17:45 | 14.62 | 43 | 28 | 10.31 | 101.4 | 7.85
2009-05-21 | 18:00 | 14.59 | 43 | 28 | 10.25 | 100.8 | 7.85
2009-05-21 | 18:15 | 14.55 | 44 | 29 | 10.27 | 100.9 | 7.84
2009-05-21 | 18:30 | 14.51 | 44 | 29 | 10.27 | 100.8 | 7.84
2009-05-21 | 18:45 | 14.48 | 45 | 29 | 10.27 | 100.7 | 7.84
2009-05-21 | 19:00 | 14.45 | 45 | 29 | 10.27 | 100.6 | 7.84
2009-05-21 | 19:15 | 14.42 | 45 | 30 | 10.26 | 100.5 | 7.84
2009-05-21 | 19:30 | 14.36 | 46 | 30 | 10.26 | 100.4 | 7.84
2009-05-21 | 19:45 | 14.29 | 46 | 30 | 10.27 | 100.3 | 7.83
2009-05-21 | 20:00 | 14.2 | 46 | 30 | 10.23 | 99.7 | 7.83
2009-05-21 | 20:15 | 14.12 | 46 | 30 | 10.27 | 99.9 | 7.82
2009-05-21 | 20:30 | 14 | 46 | 30 | 10.29 | 99.9 | 7.81
2009-05-21 | 20:45 | 13.84 | 47 | 30 | 10.28 | 99.4 | 7.81
2009-05-21 | 21:00 | 13.67 | 47 | 31 | 10.29 | 99.2 | 7.8
2009-05-21 | 21:15 | 13.49 | 47 | 30 | 10.37 | 99.5 | 7.8
2009-05-21 | 21:30 | 13.32 | 49 | 32 | 10.42 | 99.6 | 7.79
2009-05-21 | 21:45 | 13.14 | 46 | 30 | 10.43 | 99.3 | 7.79
2009-05-21 | 22:00 | 12.98 | 46 | 30 | 10.45 | 99.1 | 7.79
2009-05-21 | 22:15 | 12.83 | 46 | 30 | 10.52 | 99.5 | 7.79
2009-05-21 | 22:30 | 12.7 | 46 | 30 | 10.57 | 99.6 | 7.79
2009-05-21 | 22:45 | 12.57 | 46 | 30 | 10.56 | 99.3 | 7.79
2009-05-21 | 23:00 | 12.44 | 47 | 30 | 10.63 | 99.6 | 7.79
2009-05-21 | 23:15 | 12.32 | 47 | 30 | 10.67 | 99.8 | 7.79
2009-05-21 | 23:30 | 12.2 | 47 | 30 | 10.67 | 99.4 | 7.79
2009-05-21 | 23:45 | 12.07 | 47 | 30 | 10.68 | 99.3 | 7.79
2009-05-22 | 0:00 | 11.92 | 47 | 31 | 10.72 | 99.3 | 7.79
2009-05-22 | 0:15 | 11.78 | 47 | 31 | 10.8 | 99.7 | 7.8
2009-05-22 | 0:30 | 11.63 | 47 | 31 | 10.86 | 99.9 | 7.79
2009-05-22 | 0:45 | 11.49 | 47 | 31 | 10.91 | 100.1 | 7.79
2009-05-22 | 1:01 | 11.36 | 47 | 31 | 10.96 | 100.2 | 7.79
2009-05-22 | 1:16 | 11.23 | 48 | 31 | 11 | 100.3 | 7.78
2009-05-22 | 1:31 | 11.09 | 48 | 31 | 10.99 | 99.9 | 7.79
2009-05-22 | 1:46 | 10.95 | 48 | 31 | 11.06 | 100.2 | 7.79
2009-05-22 | 2:01 | 10.82 | 48 | 31 | 11.11 | 100.3 | 7.79
2009-05-22 | 2:16 | 10.68 | 48 | 31 | 11.16 | 100.5 | 7.78
2009-05-22 | 2:31 | 10.55 | 48 | 31 | 11.16 | 100.1 | 7.78
2009-05-22 | 2:46 | 10.41 | 48 | 31 | 11.18 | 100 | 7.78
2009-05-22 | 3:01 | 10.26 | 48 | 31 | 11.21 | 99.9 | 7.79
2009-05-22 | 3:16 | 10.12 | 49 | 32 | 11.3 | 100.4 | 7.78
2009-05-22 | 3:31 | 9.97 | 49 | 32 | 11.32 | 100.2 | 7.78
2009-05-22 | 3:46 | 9.83 | 49 | 32 | 11.35 | 100.1 | 7.78
2009-05-22 | 4:01 | 9.7 | 49 | 32 | 11.44 | 100.6 | 7.78
2009-05-22 | 4:16 | 9.56 | 48 | 32 | 11.5 | 100.8 | 7.78
2009-05-22 | 4:31 | 9.43 | 48 | 31 | 11.5 | 100.5 | 7.77
2009-05-22 | 4:46 | 9.3 | 48 | 31 | 11.58 | 100.9 | 7.77
2009-05-22 | 5:01 | 9.16 | 48 | 31 | 11.64 | 101.1 | 7.77
2009-05-22 | 5:16 | 9.02 | 47 | 31 | 11.64 | 100.8 | 7.76
2009-05-22 | 5:31 | 8.87 | 47 | 31 | 11.73 | 101.1 | 7.76
2009-05-22 | 5:46 | 8.72 | 47 | 30 | 11.74 | 100.9 | 7.75
2009-05-22 | 6:01 | 8.58 | 46 | 30 | 11.84 | 101.4 | 7.75
2009-05-22 | 6:16 | 8.44 | 46 | 30 | 11.92 | 101.7 | 7.75
2009-05-22 | 6:31 | 8.31 | 46 | 30 | 11.93 | 101.5 | 7.75
2009-05-22 | 6:46 | 8.2 | 45 | 29 | 12.03 | 102 | 7.75
2009-05-22 | 7:01 | 8.1 | 45 | 29 | 12.1 | 102.4 | 7.75
2009-05-22 | 7:16 | 8.02 | 45 | 29 | 12.15 | 102.7 | 7.75
2009-05-22 | 7:31 | 7.96 | 44 | 29 | 12.2 | 103 | 7.75
2009-05-22 | 7:46 | 7.94 | 44 | 29 | 12.19 | 102.8 | 7.76
2009-05-22 | 8:01 | 7.91 | 44 | 29 | 12.26 | 103.3 | 7.76
2009-05-22 | 8:16 | 7.91 | 44 | 29 | 12.27 | 103.4 | 7.77
2009-05-22 | 8:31 | 7.91 | 44 | 29 | 12.39 | 104.4 | 7.77
2009-05-22 | 8:46 | 7.91 | 44 | 28 | 12.51 | 105.4 | 7.77
2009-05-22 | 9:01 | 7.94 | 44 | 28 | 12.54 | 105.7 | 7.78
2009-05-22 | 9:16 | 7.98 | 44 | 28 | 12.55 | 105.9 | 7.78
2009-05-22 | 9:31 | 8.06 | 44 | 28 | 12.5 | 105.7 | 7.79
2009-05-22 | 9:46 | 8.17 | 44 | 28 | 12.49 | 105.9 | 7.79
2009-05-22 | 10:01 | 8.34 | 44 | 28 | 12.44 | 105.9 | 7.8
2009-05-22 | 10:16 | 8.55 | 44 | 28 | 12.37 | 105.8 | 7.81
2009-05-22 | 10:31 | 8.84 | 44 | 28 | 12.27 | 105.7 | 7.81
2009-05-22 | 10:46 | 9.18 | 44 | 28 | 12.21 | 106.1 | 7.82
2009-05-22 | 11:01 | 9.55 | 44 | 28 | 12.14 | 106.4 | 7.83
2009-05-22 | 11:16 | 9.94 | 44 | 28 | 12.11 | 107.2 | 7.83
2009-05-22 | 11:31 | 10.3 | 44 | 28 | 12.05 | 107.5 | 7.84
2009-05-22 | 11:46 | 10.71 | 44 | 28 | 11.96 | 107.7 | 7.85
2009-05-22 | 12:01 | 11.13 | 44 | 28 | 11.8 | 107.3 | 7.85
2009-05-22 | 12:16 | 11.56 | 44 | 29 | 11.71 | 107.6 | 7.86
2009-05-22 | 12:31 | 11.99 | 44 | 29 | 11.6 | 107.6 | 7.87
2009-05-22 | 12:46 | 12.45 | 44 | 29 | 11.47 | 107.5 | 7.87
2009-05-22 | 13:00 | 12.9 | 44 | 29 | 11.28 | 106.8 | 7.88
2009-05-22 | 13:15 | 13.35 | 44 | 29 | 11.16 | 106.8 | 7.88
2009-05-22 | 13:30 | 13.76 | 44 | 29 | 11.04 | 106.6 | 7.9
2009-05-22 | 13:45 | 14.2 | 44 | 29 | 10.87 | 106 | 7.9
2009-05-22 | 14:00 | 14.53 | 44 | 29 | 10.75 | 105.5 | 7.91
2009-05-22 | 14:15 | 14.81 | 45 | 29 | 10.69 | 105.6 | 7.92
2009-05-22 | 14:30 | 15.02 | 45 | 29 | 10.57 | 104.8 | 7.92
2009-05-22 | 14:45 | 15.2 | 45 | 29 | 10.48 | 104.3 | 7.93
2009-05-22 | 15:00 | 15.33 | 45 | 29 | 10.47 | 104.5 | 7.93
2009-05-22 | 15:15 | 15.58 | 45 | 29 | 10.41 | 104.5 | 7.94
2009-05-22 | 15:30 | 15.76 | 45 | 29 | 10.35 | 104.3 | 7.93
2009-05-22 | 15:45 | 15.95 | 45 | 29 | 10.22 | 103.5 | 7.94
2009-05-22 | 16:00 | 15.96 | 45 | 29 | 10.2 | 103.3 | 7.94
2009-05-22 | 16:15 | 15.99 | 45 | 29 | 10.12 | 102.5 | 7.93
2009-05-22 | 16:30 | 15.94 | 45 | 29 | 10.14 | 102.6 | 7.92
2009-05-22 | 16:45 | 15.93 | 45 | 30 | 10.12 | 102.4 | 7.92
2009-05-22 | 17:00 | 15.88 | 45 | 30 | 10.06 | 101.6 | 7.91
2009-05-22 | 17:15 | 15.83 | 46 | 30 | 10.02 | 101.1 | 7.91
2009-05-22 | 17:30 | 15.77 | 46 | 30 | 10 | 100.8 | 7.9
2009-05-22 | 17:45 | 15.69 | 46 | 30 | 9.99 | 100.6 | 7.9
2009-05-22 | 18:00 | 15.62 | 46 | 30 | 10.03 | 100.8 | 7.9
2009-05-22 | 18:15 | 15.54 | 46 | 30 | 10 | 100.3 | 7.89
2009-05-22 | 18:30 | 15.48 | 46 | 30 | 9.98 | 100 | 7.88
2009-05-22 | 18:45 | 15.43 | 46 | 30 | 10.01 | 100.2 | 7.88
2009-05-22 | 19:00 | 15.38 | 46 | 30 | 9.98 | 99.7 | 7.87
2009-05-22 | 19:15 | 15.35 | 46 | 30 | 10 | 99.9 | 7.87
2009-05-22 | 19:30 | 15.31 | 46 | 30 | 10.01 | 100 | 7.86
2009-05-22 | 19:45 | 15.25 | 46 | 30 | 9.98 | 99.5 | 7.85
2009-05-22 | 20:00 | 15.15 | 47 | 30 | 10.02 | 99.7 | 7.85
2009-05-22 | 20:15 | 15.05 | 47 | 30 | 10.01 | 99.4 | 7.84
2009-05-22 | 20:30 | 14.92 | 47 | 30 | 10.05 | 99.5 | 7.83
2009-05-22 | 20:45 | 14.74 | 47 | 30 | 10.09 | 99.5 | 7.83
2009-05-22 | 21:00 | 14.54 | 47 | 31 | 10.08 | 99 | 7.82
2009-05-22 | 21:15 | 14.33 | 47 | 31 | 10.11 | 98.8 | 7.81
2009-05-22 | 21:30 | 14.13 | 47 | 31 | 10.15 | 98.8 | 7.81
2009-05-22 | 21:45 | 13.95 | 47 | 31 | 10.19 | 98.7 | 7.81
2009-05-22 | 22:00 | 13.78 | 47 | 31 | 10.27 | 99.1 | 7.81
2009-05-22 | 22:15 | 13.62 | 48 | 31 | 10.32 | 99.3 | 7.81
2009-05-22 | 22:30 | 13.47 | 48 | 31 | 10.37 | 99.5 | 7.8
2009-05-22 | 22:45 | 13.33 | 48 | 31 | 10.37 | 99.1 | 7.81
2009-05-22 | 23:00 | 13.18 | 48 | 31 | 10.43 | 99.4 | 7.8
2009-05-22 | 23:15 | 13.02 | 48 | 31 | 10.5 | 99.7 | 7.81
2009-05-22 | 23:30 | 12.86 | 48 | 31 | 10.54 | 99.7 | 7.8
2009-05-22 | 23:45 | 12.7 | 48 | 31 | 10.54 | 99.3 | 7.81
2009-05-23 | 0:00 | 12.54 | 48 | 31 | 10.56 | 99.2 | 7.8
2009-05-23 | 0:15 | 12.39 | 49 | 32 | 10.59 | 99.2 | 7.8
2009-05-23 | 0:30 | 12.24 | 49 | 32 | 10.63 | 99.2 | 7.8
2009-05-23 | 0:45 | 12.1 | 49 | 32 | 10.67 | 99.2 | 7.8
2009-05-23 | 1:01 | 11.97 | 49 | 32 | 10.7 | 99.2 | 7.8
2009-05-23 | 1:16 | 11.83 | 49 | 32 | 10.79 | 99.7 | 7.8
2009-05-23 | 1:31 | 11.69 | 49 | 32 | 10.85 | 100 | 7.8
2009-05-23 | 1:46 | 11.55 | 49 | 32 | 10.85 | 99.7 | 7.8
2009-05-23 | 2:01 | 11.41 | 49 | 32 | 10.88 | 99.6 | 7.8
2009-05-23 | 2:16 | 11.26 | 49 | 32 | 10.96 | 100 | 7.8
2009-05-23 | 2:31 | 11.11 | 50 | 32 | 10.98 | 99.8 | 7.8
2009-05-23 | 2:46 | 10.95 | 50 | 32 | 11.02 | 99.8 | 7.8
2009-05-23 | 3:01 | 10.8 | 50 | 32 | 11.1 | 100.2 | 7.8
2009-05-23 | 3:16 | 10.65 | 50 | 32 | 11.17 | 100.5 | 7.79
2009-05-23 | 3:31 | 10.51 | 50 | 32 | 11.18 | 100.2 | 7.79
2009-05-23 | 3:46 | 10.36 | 50 | 32 | 11.25 | 100.5 | 7.79
2009-05-23 | 4:01 | 10.22 | 50 | 32 | 11.31 | 100.7 | 7.79
2009-05-23 | 4:16 | 10.07 | 50 | 32 | 11.36 | 100.8 | 7.79
2009-05-23 | 4:31 | 9.93 | 49 | 32 | 11.41 | 100.9 | 7.78
2009-05-23 | 4:46 | 9.79 | 49 | 32 | 11.46 | 101 | 7.78
2009-05-23 | 5:01 | 9.65 | 49 | 32 | 11.51 | 101.1 | 7.77
2009-05-23 | 5:16 | 9.5 | 48 | 31 | 11.56 | 101.2 | 7.77
2009-05-23 | 5:31 | 9.35 | 48 | 31 | 11.61 | 101.3 | 7.76
2009-05-23 | 5:46 | 9.19 | 47 | 31 | 11.67 | 101.4 | 7.75
2009-05-23 | 6:01 | 9.04 | 47 | 30 | 11.68 | 101.1 | 7.75
2009-05-23 | 6:16 | 8.89 | 46 | 30 | 11.73 | 101.2 | 7.75
2009-05-23 | 6:31 | 8.76 | 46 | 30 | 11.83 | 101.8 | 7.75
2009-05-23 | 6:46 | 8.64 | 45 | 29 | 11.91 | 102.1 | 7.75
2009-05-23 | 7:01 | 8.54 | 45 | 29 | 11.98 | 102.5 | 7.75
2009-05-23 | 7:16 | 8.47 | 45 | 29 | 12.03 | 102.8 | 7.75
2009-05-23 | 7:31 | 8.42 | 44 | 29 | 12.08 | 103 | 7.75
2009-05-23 | 7:46 | 8.39 | 44 | 29 | 12.06 | 102.8 | 7.75
2009-05-23 | 8:01 | 8.38 | 44 | 29 | 12.08 | 102.9 | 7.76
2009-05-23 | 8:16 | 8.37 | 44 | 28 | 12.12 | 103.2 | 7.76
2009-05-23 | 8:31 | 8.37 | 44 | 28 | 12.19 | 103.8 | 7.77
2009-05-23 | 8:46 | 8.39 | 44 | 28 | 12.33 | 105.1 | 7.77
2009-05-23 | 9:01 | 8.41 | 43 | 28 | 12.43 | 106 | 7.78
2009-05-23 | 9:16 | 8.46 | 43 | 28 | 12.48 | 106.6 | 7.78
2009-05-23 | 9:31 | 8.55 | 43 | 28 | 12.44 | 106.5 | 7.79
2009-05-23 | 9:46 | 8.68 | 43 | 28 | 12.32 | 105.8 | 7.79
2009-05-23 | 10:01 | 8.86 | 43 | 28 | 12.22 | 105.4 | 7.8
2009-05-23 | 10:16 | 9.08 | 43 | 28 | 12.13 | 105.1 | 7.81
2009-05-23 | 10:31 | 9.39 | 43 | 28 | 12.01 | 104.9 | 7.82
2009-05-23 | 10:46 | 9.75 | 43 | 28 | 11.99 | 105.6 | 7.83
2009-05-23 | 11:01 | 10.12 | 43 | 28 | 11.99 | 106.5 | 7.83
2009-05-23 | 11:16 | 10.51 | 43 | 28 | 11.89 | 106.6 | 7.84
2009-05-23 | 11:31 | 10.91 | 43 | 28 | 11.79 | 106.7 | 7.85
2009-05-23 | 11:46 | 11.33 | 43 | 28 | 11.69 | 106.8 | 7.85
2009-05-23 | 12:01 | 11.75 | 43 | 28 | 11.63 | 107.3 | 7.86
2009-05-23 | 12:16 | 12.18 | 43 | 28 | 11.53 | 107.4 | 7.87
2009-05-23 | 12:31 | 12.63 | 43 | 28 | 11.36 | 106.9 | 7.87
2009-05-23 | 12:46 | 13.08 | 43 | 28 | 11.21 | 106.6 | 7.89
2009-05-23 | 13:00 | 13.54 | 44 | 28 | 11.12 | 106.8 | 7.89
2009-05-23 | 13:15 | 14 | 44 | 28 | 10.95 | 106.2 | 7.9
2009-05-23 | 13:30 | 14.46 | 44 | 28 | 10.84 | 106.3 | 7.9
2009-05-23 | 13:45 | 14.9 | 44 | 28 | 10.67 | 105.6 | 7.91
2009-05-23 | 14:00 | 15.3 | 44 | 29 | 10.58 | 105.6 | 7.92
2009-05-23 | 14:15 | 15.68 | 44 | 29 | 10.43 | 104.9 | 7.92
2009-05-23 | 14:30 | 16.02 | 44 | 29 | 10.34 | 104.8 | 7.93
2009-05-23 | 14:45 | 16.32 | 44 | 29 | 10.24 | 104.5 | 7.93
2009-05-23 | 15:00 | 16.56 | 44 | 29 | 10.11 | 103.7 | 7.94
2009-05-23 | 15:15 | 16.78 | 44 | 29 | 10.05 | 103.5 | 7.94
2009-05-23 | 15:30 | 16.92 | 45 | 29 | 9.96 | 102.9 | 7.94
2009-05-23 | 15:45 | 17 | 45 | 29 | 9.93 | 102.7 | 7.93
2009-05-23 | 16:00 | 17.07 | 45 | 29 | 9.89 | 102.5 | 7.93
2009-05-23 | 16:15 | 17.09 | 45 | 29 | 9.87 | 102.3 | 7.93
2009-05-23 | 16:30 | 17.06 | 45 | 29 | 9.84 | 102 | 7.93
2009-05-23 | 16:45 | 17.04 | 45 | 29 | 9.82 | 101.7 | 7.92
2009-05-23 | 17:00 | 16.99 | 45 | 29 | 9.8 | 101.3 | 7.92
2009-05-23 | 17:15 | 16.92 | 45 | 29 | 9.75 | 100.7 | 7.91
2009-05-23 | 17:30 | 16.81 | 45 | 29 | 9.77 | 100.7 | 7.91
2009-05-23 | 17:45 | 16.78 | 45 | 29 | 9.78 | 100.7 | 7.9
2009-05-23 | 18:00 | 16.7 | 45 | 30 | 9.76 | 100.3 | 7.9
2009-05-23 | 18:15 | 16.71 | 45 | 30 | 9.74 | 100.2 | 7.9
2009-05-23 | 18:30 | 16.71 | 46 | 30 | 9.79 | 100.7 | 7.91
2009-05-23 | 18:45 | 16.68 | 46 | 30 | 9.79 | 100.6 | 7.9
2009-05-23 | 19:00 | 16.65 | 46 | 30 | 9.77 | 100.4 | 7.9
2009-05-23 | 19:15 | 16.56 | 46 | 30 | 9.73 | 99.8 | 7.89
2009-05-23 | 19:30 | 16.44 | 46 | 30 | 9.77 | 99.9 | 7.88
2009-05-23 | 19:45 | 16.3 | 46 | 30 | 9.75 | 99.4 | 7.86
2009-05-23 | 20:00 | 16.16 | 46 | 30 | 9.75 | 99.1 | 7.86
2009-05-23 | 20:15 | 16.04 | 45 | 29 | 9.75 | 98.9 | 7.85
2009-05-23 | 20:30 | 15.91 | 45 | 29 | 9.76 | 98.7 | 7.84
2009-05-23 | 20:45 | 15.77 | 45 | 29 | 9.81 | 99 | 7.82
2009-05-23 | 21:00 | 15.61 | 44 | 29 | 9.82 | 98.6 | 7.81
2009-05-23 | 21:15 | 15.45 | 44 | 29 | 9.88 | 98.9 | 7.8
2009-05-23 | 21:30 | 15.28 | 44 | 29 | 9.88 | 98.6 | 7.79
2009-05-23 | 21:45 | 15.13 | 44 | 29 | 9.91 | 98.6 | 7.8
2009-05-23 | 22:00 | 15 | 44 | 29 | 9.98 | 98.9 | 7.79
2009-05-23 | 22:15 | 14.85 | 44 | 29 | 9.98 | 98.7 | 7.79
2009-05-23 | 22:30 | 14.7 | 44 | 29 | 10.05 | 99 | 7.79
2009-05-23 | 22:45 | 14.55 | 44 | 29 | 10.06 | 98.8 | 7.79
2009-05-23 | 23:00 | 14.41 | 44 | 29 | 10.12 | 99.1 | 7.79
2009-05-23 | 23:15 | 14.28 | 45 | 29 | 10.17 | 99.3 | 7.79
2009-05-23 | 23:30 | 14.15 | 45 | 29 | 10.17 | 99 | 7.79
2009-05-23 | 23:45 | 14.02 | 45 | 29 | 10.19 | 98.9 | 7.79
2009-05-24 | 0:00 | 13.88 | 45 | 29 | 10.26 | 99.3 | 7.78
2009-05-24 | 0:15 | 13.74 | 45 | 29 | 10.31 | 99.5 | 7.78
2009-05-24 | 0:30 | 13.58 | 45 | 29 | 10.37 | 99.6 | 7.79
2009-05-24 | 0:45 | 13.41 | 45 | 29 | 10.41 | 99.7 | 7.79
2009-05-24 | 1:01 | 13.23 | 45 | 29 | 10.47 | 99.8 | 7.78
2009-05-24 | 1:16 | 13.05 | 45 | 29 | 10.52 | 100 | 7.78
2009-05-24 | 1:31 | 12.89 | 45 | 29 | 10.57 | 100 | 7.78
2009-05-24 | 1:46 | 12.75 | 45 | 29 | 10.61 | 100.1 | 7.77
2009-05-24 | 2:01 | 12.6 | 45 | 29 | 10.65 | 100.2 | 7.77
2009-05-24 | 2:16 | 12.46 | 45 | 29 | 10.7 | 100.3 | 7.77
2009-05-24 | 2:31 | 12.31 | 45 | 29 | 10.69 | 99.9 | 7.77
2009-05-24 | 2:46 | 12.15 | 44 | 29 | 10.72 | 99.8 | 7.76
2009-05-24 | 3:01 | 11.98 | 44 | 29 | 10.75 | 99.7 | 7.75
2009-05-24 | 3:16 | 11.79 | 44 | 28 | 10.8 | 99.8 | 7.75
2009-05-24 | 3:31 | 11.6 | 43 | 28 | 10.91 | 100.3 | 7.74
2009-05-24 | 3:46 | 11.42 | 43 | 28 | 10.93 | 100.1 | 7.74
2009-05-24 | 4:01 | 11.24 | 43 | 28 | 11.02 | 100.5 | 7.73
2009-05-24 | 4:16 | 11.07 | 42 | 27 | 11.09 | 100.7 | 7.73
2009-05-24 | 4:31 | 10.89 | 42 | 27 | 11.15 | 100.9 | 7.72
2009-05-24 | 4:46 | 10.72 | 42 | 27 | 11.16 | 100.6 | 7.72
2009-05-24 | 5:01 | 10.56 | 41 | 27 | 11.19 | 100.5 | 7.72
2009-05-24 | 5:16 | 10.4 | 41 | 27 | 11.28 | 100.9 | 7.71
2009-05-24 | 5:31 | 10.24 | 41 | 26 | 11.35 | 101.1 | 7.7
2009-05-24 | 5:46 | 10.1 | 41 | 26 | 11.36 | 100.8 | 7.7
2009-05-24 | 6:01 | 9.95 | 40 | 26 | 11.4 | 100.9 | 7.7
2009-05-24 | 6:16 | 9.82 | 40 | 26 | 11.45 | 101 | 7.7
2009-05-24 | 6:31 | 9.69 | 40 | 26 | 11.5 | 101.1 | 7.7
2009-05-24 | 6:46 | 9.58 | 40 | 26 | 11.6 | 101.8 | 7.7
2009-05-24 | 7:01 | 9.47 | 40 | 26 | 11.63 | 101.7 | 7.71
2009-05-24 | 7:16 | 9.36 | 40 | 26 | 11.73 | 102.3 | 7.7
2009-05-24 | 7:31 | 9.27 | 39 | 26 | 11.8 | 102.7 | 7.71
2009-05-24 | 7:46 | 9.2 | 39 | 26 | 11.81 | 102.6 | 7.71
2009-05-24 | 8:01 | 9.14 | 39 | 25 | 11.89 | 103.2 | 7.71
2009-05-24 | 8:16 | 9.08 | 39 | 25 | 11.92 | 103.3 | 7.72
2009-05-24 | 8:31 | 9.03 | 39 | 25 | 12 | 103.9 | 7.73
2009-05-24 | 8:46 | 8.99 | 39 | 25 | 12.1 | 104.6 | 7.73
2009-05-24 | 9:01 | 8.97 | 39 | 25 | 12.19 | 105.3 | 7.74
2009-05-24 | 9:16 | 8.98 | 39 | 25 | 12.23 | 105.8 | 7.74
2009-05-24 | 9:31 | 9.03 | 39 | 25 | 12.25 | 106.1 | 7.74
2009-05-24 | 9:46 | 9.12 | 39 | 25 | 12.16 | 105.5 | 7.74
2009-05-24 | 10:01 | 9.26 | 38 | 25 | 12.14 | 105.7 | 7.75
2009-05-24 | 10:16 | 9.46 | 38 | 25 | 12.08 | 105.6 | 7.75
2009-05-24 | 10:31 | 9.7 | 38 | 25 | 11.94 | 105.1 | 7.77
2009-05-24 | 10:46 | 9.98 | 38 | 25 | 11.93 | 105.6 | 7.77
2009-05-24 | 11:01 | 10.28 | 38 | 25 | 11.95 | 106.5 | 7.78
2009-05-24 | 11:16 | 10.58 | 38 | 25 | 11.93 | 107.1 | 7.78
2009-05-24 | 11:31 | 10.91 | 39 | 25 | 11.82 | 107 | 7.79
2009-05-24 | 11:46 | 11.26 | 39 | 25 | 11.78 | 107.4 | 7.79
2009-05-24 | 12:01 | 11.62 | 39 | 25 | 11.64 | 107.1 | 7.8
2009-05-24 | 12:16 | 12 | 39 | 25 | 11.52 | 106.9 | 7.8
2009-05-24 | 12:31 | 12.39 | 39 | 25 | 11.4 | 106.7 | 7.81
2009-05-24 | 12:46 | 12.78 | 39 | 25 | 11.33 | 107 | 7.81
2009-05-24 | 13:00 | 13.18 | 39 | 25 | 11.18 | 106.5 | 7.83
2009-05-24 | 13:15 | 13.58 | 39 | 25 | 11.09 | 106.7 | 7.84
2009-05-24 | 13:30 | 13.97 | 39 | 25 | 10.99 | 106.6 | 7.83
2009-05-24 | 13:45 | 14.34 | 39 | 25 | 10.83 | 105.9 | 7.85
2009-05-24 | 14:00 | 14.69 | 39 | 25 | 10.75 | 105.9 | 7.85
2009-05-24 | 14:15 | 15 | 39 | 25 | 10.65 | 105.6 | 7.85
2009-05-24 | 14:30 | 15.27 | 39 | 26 | 10.55 | 105.2 | 7.85
2009-05-24 | 14:45 | 15.51 | 39 | 26 | 10.45 | 104.8 | 7.86
2009-05-24 | 15:00 | 15.73 | 39 | 26 | 10.32 | 104 | 7.87
2009-05-24 | 15:15 | 15.91 | 39 | 26 | 10.28 | 103.9 | 7.87
2009-05-24 | 15:30 | 16.04 | 40 | 26 | 10.22 | 103.7 | 7.87
2009-05-24 | 15:45 | 16.15 | 40 | 26 | 10.17 | 103.4 | 7.87
2009-05-24 | 16:00 | 16.21 | 40 | 26 | 10.13 | 103.1 | 7.86
2009-05-24 | 16:15 | 16.27 | 40 | 26 | 10.09 | 102.8 | 7.87
2009-05-24 | 16:30 | 16.29 | 40 | 26 | 10.06 | 102.5 | 7.87
2009-05-24 | 16:45 | 16.3 | 40 | 26 | 10.03 | 102.3 | 7.86
2009-05-24 | 17:00 | 16.29 | 40 | 26 | 10.01 | 102 | 7.85
2009-05-24 | 17:15 | 16.25 | 40 | 26 | 10 | 101.8 | 7.85
2009-05-24 | 17:30 | 16.23 | 40 | 26 | 9.98 | 101.6 | 7.84
2009-05-24 | 17:45 | 16.15 | 40 | 26 | 9.93 | 101 | 7.84
2009-05-24 | 18:00 | 16.08 | 40 | 26 | 9.91 | 100.6 | 7.84
2009-05-24 | 18:15 | 16.03 | 40 | 26 | 9.9 | 100.3 | 7.83
2009-05-24 | 18:30 | 15.98 | 40 | 26 | 9.93 | 100.6 | 7.82
2009-05-24 | 18:45 | 15.94 | 41 | 26 | 9.9 | 100.1 | 7.82
2009-05-24 | 19:00 | 15.89 | 41 | 26 | 9.89 | 99.9 | 7.81
2009-05-24 | 19:15 | 15.82 | 41 | 26 | 9.93 | 100.2 | 7.81
2009-05-24 | 19:30 | 15.73 | 41 | 26 | 9.96 | 100.3 | 7.81
2009-05-24 | 19:45 | 15.64 | 41 | 27 | 9.98 | 100.3 | 7.8
2009-05-24 | 20:00 | 15.49 | 41 | 27 | 9.96 | 99.8 | 7.8
2009-05-24 | 20:15 | 15.34 | 41 | 27 | 9.97 | 99.6 | 7.79
2009-05-24 | 20:30 | 15.19 | 41 | 27 | 9.98 | 99.4 | 7.78
2009-05-24 | 20:45 | 15 | 41 | 27 | 10.06 | 99.7 | 7.78
2009-05-24 | 21:00 | 14.82 | 41 | 27 | 10.11 | 99.9 | 7.77
2009-05-24 | 21:15 | 14.64 | 41 | 27 | 10.11 | 99.5 | 7.77
2009-05-24 | 21:30 | 14.48 | 42 | 27 | 10.13 | 99.3 | 7.76
2009-05-24 | 21:45 | 14.32 | 42 | 27 | 10.2 | 99.7 | 7.76
2009-05-24 | 22:00 | 14.15 | 42 | 27 | 10.26 | 99.9 | 7.76
2009-05-24 | 22:15 | 14 | 42 | 27 | 10.31 | 100 | 7.76
2009-05-24 | 22:30 | 13.85 | 42 | 27 | 10.3 | 99.7 | 7.76
2009-05-24 | 22:45 | 13.72 | 42 | 27 | 10.37 | 100 | 7.76
2009-05-24 | 23:00 | 13.58 | 42 | 27 | 10.42 | 100.1 | 7.76
2009-05-24 | 23:15 | 13.43 | 42 | 27 | 10.42 | 99.8 | 7.76
2009-05-24 | 23:30 | 13.26 | 42 | 27 | 10.49 | 100.2 | 7.76
2009-05-24 | 23:45 | 13.1 | 42 | 27 | 10.5 | 99.9 | 7.76
2009-05-25 | 0:00 | 12.93 | 42 | 27 | 10.59 | 100.3 | 7.75
2009-05-25 | 0:15 | 12.76 | 42 | 27 | 10.65 | 100.5 | 7.75
2009-05-25 | 0:30 | 12.57 | 42 | 27 | 10.66 | 100.2 | 7.75
2009-05-25 | 0:45 | 12.38 | 42 | 28 | 10.75 | 100.6 | 7.75
2009-05-25 | 1:01 | 12.2 | 42 | 28 | 10.76 | 100.3 | 7.75
2009-05-25 | 1:16 | 12.04 | 42 | 28 | 10.85 | 100.7 | 7.74
2009-05-25 | 1:31 | 11.87 | 42 | 27 | 10.91 | 100.9 | 7.74
2009-05-25 | 1:46 | 11.7 | 42 | 27 | 10.97 | 101.1 | 7.74
2009-05-25 | 2:01 | 11.52 | 42 | 27 | 11.02 | 101.2 | 7.73
2009-05-25 | 2:16 | 11.35 | 42 | 27 | 11.08 | 101.3 | 7.73
2009-05-25 | 2:31 | 11.18 | 41 | 27 | 11.13 | 101.4 | 7.72
2009-05-25 | 2:46 | 11.01 | 41 | 27 | 11.14 | 101 | 7.72
2009-05-25 | 3:01 | 10.82 | 41 | 26 | 11.17 | 100.9 | 7.72
2009-05-25 | 3:16 | 10.63 | 40 | 26 | 11.27 | 101.4 | 7.7
2009-05-25 | 3:31 | 10.44 | 40 | 26 | 11.35 | 101.6 | 7.7
2009-05-25 | 3:46 | 10.26 | 39 | 25 | 11.37 | 101.3 | 7.69
2009-05-25 | 4:01 | 10.08 | 39 | 25 | 11.4 | 101.2 | 7.68
2009-05-25 | 4:16 | 9.91 | 39 | 25 | 11.5 | 101.7 | 7.67
2009-05-25 | 4:31 | 9.75 | 38 | 25 | 11.51 | 101.4 | 7.67
2009-05-25 | 4:46 | 9.6 | 38 | 25 | 11.6 | 101.8 | 7.67
2009-05-25 | 5:01 | 9.46 | 38 | 24 | 11.61 | 101.6 | 7.67
2009-05-25 | 5:16 | 9.34 | 37 | 24 | 11.69 | 102 | 7.66
2009-05-25 | 5:31 | 9.22 | 37 | 24 | 11.7 | 101.7 | 7.66
2009-05-25 | 5:46 | 9.09 | 37 | 24 | 11.73 | 101.7 | 7.66
2009-05-25 | 6:01 | 8.97 | 37 | 24 | 11.82 | 102.2 | 7.65
2009-05-25 | 6:16 | 8.84 | 37 | 24 | 11.9 | 102.6 | 7.66
2009-05-25 | 6:31 | 8.72 | 37 | 24 | 11.97 | 102.8 | 7.66
2009-05-25 | 6:46 | 8.61 | 37 | 24 | 11.98 | 102.6 | 7.66
2009-05-25 | 7:01 | 8.49 | 36 | 24 | 12.02 | 102.7 | 7.66
2009-05-25 | 7:16 | 8.38 | 36 | 24 | 12.12 | 103.3 | 7.66
2009-05-25 | 7:31 | 8.29 | 36 | 24 | 12.2 | 103.8 | 7.67
2009-05-25 | 7:46 | 8.21 | 36 | 24 | 12.27 | 104.1 | 7.67
2009-05-25 | 8:01 | 8.13 | 36 | 23 | 12.33 | 104.5 | 7.67
2009-05-25 | 8:16 | 8.06 | 36 | 23 | 12.4 | 104.9 | 7.67
2009-05-25 | 8:31 | 8 | 36 | 23 | 12.52 | 105.7 | 7.68
2009-05-25 | 8:46 | 7.94 | 36 | 23 | 12.58 | 106 | 7.68
2009-05-25 | 9:01 | 7.91 | 36 | 23 | 12.69 | 107 | 7.69
2009-05-25 | 9:16 | 7.91 | 36 | 23 | 12.71 | 107.1 | 7.69
2009-05-25 | 9:31 | 7.95 | 36 | 23 | 12.67 | 106.8 | 7.7
2009-05-25 | 9:46 | 8.04 | 36 | 23 | 12.65 | 107 | 7.7
2009-05-25 | 10:01 | 8.18 | 36 | 23 | 12.55 | 106.4 | 7.7
2009-05-25 | 10:16 | 8.38 | 36 | 23 | 12.51 | 106.6 | 7.72
2009-05-25 | 10:31 | 8.61 | 36 | 23 | 12.44 | 106.6 | 7.72
2009-05-25 | 10:46 | 8.87 | 36 | 23 | 12.39 | 106.9 | 7.73
2009-05-25 | 11:01 | 9.15 | 36 | 24 | 12.4 | 107.7 | 7.73
2009-05-25 | 11:16 | 9.44 | 36 | 24 | 12.32 | 107.7 | 7.74
2009-05-25 | 11:31 | 9.76 | 36 | 24 | 12.3 | 108.3 | 7.74
2009-05-25 | 11:46 | 10.08 | 36 | 24 | 12.23 | 108.6 | 7.75
2009-05-25 | 12:01 | 10.43 | 36 | 24 | 12.09 | 108.2 | 7.76
2009-05-25 | 12:16 | 10.79 | 36 | 24 | 12.02 | 108.4 | 7.76
2009-05-25 | 12:31 | 11.16 | 37 | 24 | 11.86 | 108 | 7.77
2009-05-25 | 12:46 | 11.55 | 37 | 24 | 11.73 | 107.7 | 7.77
2009-05-25 | 13:00 | 11.94 | 37 | 24 | 11.59 | 107.4 | 7.78
2009-05-25 | 13:15 | 12.32 | 37 | 24 | 11.51 | 107.6 | 7.79
2009-05-25 | 13:30 | 12.71 | 37 | 24 | 11.35 | 107 | 7.79
2009-05-25 | 13:45 | 13.07 | 37 | 24 | 11.27 | 107.2 | 7.8
2009-05-25 | 14:00 | 13.43 | 37 | 24 | 11.17 | 107 | 7.8
2009-05-25 | 14:15 | 13.74 | 37 | 24 | 11.06 | 106.7 | 7.81
2009-05-25 | 14:30 | 14.04 | 37 | 24 | 10.89 | 105.8 | 7.82
2009-05-25 | 14:45 | 14.3 | 37 | 24 | 10.82 | 105.7 | 7.82
2009-05-25 | 15:00 | 14.54 | 37 | 24 | 10.69 | 105 | 7.82
2009-05-25 | 15:15 | 14.73 | 38 | 24 | 10.64 | 104.9 | 7.84
2009-05-25 | 15:30 | 14.9 | 38 | 24 | 10.58 | 104.6 | 7.83
2009-05-25 | 15:45 | 15.02 | 38 | 25 | 10.52 | 104.4 | 7.84
2009-05-25 | 16:00 | 15.14 | 38 | 25 | 10.46 | 104.1 | 7.84
2009-05-25 | 16:15 | 15.23 | 38 | 25 | 10.42 | 103.8 | 7.84
2009-05-25 | 16:30 | 15.27 | 38 | 25 | 10.38 | 103.5 | 7.83
2009-05-25 | 16:45 | 15.28 | 38 | 25 | 10.3 | 102.8 | 7.83
2009-05-25 | 17:00 | 15.27 | 38 | 25 | 10.3 | 102.7 | 7.83
2009-05-25 | 17:15 | 15.25 | 38 | 25 | 10.29 | 102.6 | 7.83
2009-05-25 | 17:30 | 15.21 | 39 | 25 | 10.28 | 102.4 | 7.82
2009-05-25 | 17:45 | 15.19 | 39 | 25 | 10.27 | 102.3 | 7.82
2009-05-25 | 18:00 | 15.16 | 39 | 25 | 10.21 | 101.6 | 7.81
2009-05-25 | 18:15 | 15.12 | 39 | 25 | 10.23 | 101.7 | 7.81
2009-05-25 | 18:30 | 15.09 | 39 | 25 | 10.23 | 101.7 | 7.8
2009-05-25 | 18:45 | 15.06 | 39 | 25 | 10.23 | 101.6 | 7.8
2009-05-25 | 19:00 | 15.04 | 39 | 25 | 10.18 | 101 | 7.79
2009-05-25 | 19:15 | 14.98 | 39 | 25 | 10.16 | 100.7 | 7.79
2009-05-25 | 19:30 | 14.91 | 39 | 26 | 10.2 | 101 | 7.79
2009-05-25 | 19:45 | 14.81 | 39 | 26 | 10.23 | 101.1 | 7.79
2009-05-25 | 20:00 | 14.7 | 39 | 26 | 10.21 | 100.6 | 7.78
2009-05-25 | 20:15 | 14.57 | 40 | 26 | 10.21 | 100.3 | 7.78
2009-05-25 | 20:30 | 14.42 | 40 | 26 | 10.28 | 100.6 | 7.77
2009-05-25 | 20:45 | 14.25 | 40 | 26 | 10.27 | 100.2 | 7.77
2009-05-25 | 21:00 | 14.09 | 40 | 26 | 10.33 | 100.5 | 7.76
2009-05-25 | 21:15 | 13.94 | 40 | 26 | 10.33 | 100.1 | 7.75
2009-05-25 | 21:30 | 13.8 | 40 | 26 | 10.35 | 100 | 7.75
2009-05-25 | 21:45 | 13.66 | 40 | 26 | 10.37 | 99.9 | 7.75
2009-05-25 | 22:00 | 13.54 | 40 | 26 | 10.45 | 100.3 | 7.75
2009-05-25 | 22:15 | 13.4 | 40 | 26 | 10.5 | 100.5 | 7.74
2009-05-25 | 22:30 | 13.26 | 40 | 26 | 10.5 | 100.2 | 7.74
2009-05-25 | 22:45 | 13.12 | 40 | 26 | 10.57 | 100.6 | 7.74
2009-05-25 | 23:00 | 12.99 | 41 | 26 | 10.57 | 100.2 | 7.74
2009-05-25 | 23:15 | 12.87 | 41 | 26 | 10.63 | 100.6 | 7.74
2009-05-25 | 23:30 | 12.73 | 41 | 26 | 10.68 | 100.8 | 7.74
2009-05-25 | 23:45 | 12.59 | 41 | 27 | 10.68 | 100.5 | 7.74
2009-05-26 | 0:00 | 12.45 | 41 | 27 | 10.71 | 100.4 | 7.74
2009-05-26 | 0:15 | 12.31 | 41 | 27 | 10.74 | 100.3 | 7.74
2009-05-26 | 0:30 | 12.16 | 41 | 27 | 10.83 | 100.9 | 7.74
2009-05-26 | 0:45 | 12 | 41 | 27 | 10.89 | 101.1 | 7.74
2009-05-26 | 1:01 | 11.82 | 41 | 27 | 10.95 | 101.2 | 7.74
2009-05-26 | 1:16 | 11.64 | 41 | 27 | 11.02 | 101.4 | 7.74
2009-05-26 | 1:31 | 11.48 | 41 | 27 | 11.07 | 101.5 | 7.74
2009-05-26 | 1:46 | 11.3 | 41 | 27 | 11.13 | 101.6 | 7.73
2009-05-26 | 2:01 | 11.14 | 41 | 27 | 11.18 | 101.7 | 7.73
2009-05-26 | 2:16 | 10.97 | 41 | 27 | 11.18 | 101.3 | 7.73
2009-05-26 | 2:31 | 10.82 | 41 | 27 | 11.21 | 101.2 | 7.72
2009-05-26 | 2:46 | 10.66 | 41 | 26 | 11.25 | 101.2 | 7.72
2009-05-26 | 3:01 | 10.5 | 40 | 26 | 11.29 | 101.2 | 7.72
2009-05-26 | 3:16 | 10.32 | 40 | 26 | 11.34 | 101.2 | 7.71
2009-05-26 | 3:31 | 10.15 | 40 | 26 | 11.39 | 101.3 | 7.71
2009-05-26 | 3:46 | 9.97 | 39 | 26 | 11.5 | 101.8 | 7.7
2009-05-26 | 4:01 | 9.81 | 39 | 25 | 11.58 | 102.1 | 7.69
2009-05-26 | 4:16 | 9.66 | 39 | 25 | 11.58 | 101.8 | 7.69
2009-05-26 | 4:31 | 9.52 | 39 | 25 | 11.67 | 102.2 | 7.68
2009-05-26 | 4:46 | 9.37 | 38 | 25 | 11.68 | 101.9 | 7.68
2009-05-26 | 5:01 | 9.24 | 38 | 25 | 11.76 | 102.3 | 7.68
2009-05-26 | 5:16 | 9.12 | 38 | 25 | 11.82 | 102.5 | 7.67
2009-05-26 | 5:31 | 9 | 38 | 25 | 11.81 | 102.2 | 7.67
2009-05-26 | 5:46 | 8.9 | 38 | 24 | 11.89 | 102.6 | 7.67
2009-05-26 | 6:01 | 8.79 | 37 | 24 | 11.95 | 102.9 | 7.67
2009-05-26 | 6:16 | 8.7 | 37 | 24 | 12.01 | 103.1 | 7.67
2009-05-26 | 6:31 | 8.6 | 37 | 24 | 12 | 102.8 | 7.67
2009-05-26 | 6:46 | 8.51 | 37 | 24 | 12.09 | 103.3 | 7.67
2009-05-26 | 7:01 | 8.42 | 37 | 24 | 12.1 | 103.2 | 7.67
2009-05-26 | 7:16 | 8.34 | 37 | 24 | 12.18 | 103.7 | 7.67
2009-05-26 | 7:31 | 8.27 | 37 | 24 | 12.19 | 103.6 | 7.68
2009-05-26 | 7:46 | 8.21 | 37 | 24 | 12.28 | 104.2 | 7.68
2009-05-26 | 8:01 | 8.16 | 37 | 24 | 12.29 | 104.2 | 7.68
2009-05-26 | 8:16 | 8.12 | 37 | 24 | 12.34 | 104.5 | 7.68
2009-05-26 | 8:31 | 8.09 | 37 | 24 | 12.49 | 105.7 | 7.7
2009-05-26 | 8:46 | 8.06 | 37 | 24 | 12.62 | 106.7 | 7.69
2009-05-26 | 9:01 | 8.06 | 37 | 24 | 12.72 | 107.6 | 7.7
2009-05-26 | 9:16 | 8.09 | 37 | 24 | 12.71 | 107.6 | 7.71
2009-05-26 | 9:31 | 8.16 | 37 | 24 | 12.65 | 107.2 | 7.71
2009-05-26 | 9:46 | 8.28 | 37 | 24 | 12.62 | 107.3 | 7.72
2009-05-26 | 10:01 | 8.45 | 37 | 24 | 12.56 | 107.2 | 7.72
2009-05-26 | 10:16 | 8.67 | 37 | 24 | 12.47 | 107.1 | 7.73
2009-05-26 | 10:31 | 8.93 | 37 | 24 | 12.38 | 107 | 7.73
2009-05-26 | 10:46 | 9.22 | 37 | 24 | 12.33 | 107.2 | 7.74
2009-05-26 | 11:01 | 9.52 | 37 | 24 | 12.27 | 107.4 | 7.74
2009-05-26 | 11:16 | 9.83 | 37 | 24 | 12.21 | 107.7 | 7.75
2009-05-26 | 11:31 | 10.16 | 37 | 24 | 12.19 | 108.4 | 7.76
2009-05-26 | 11:46 | 10.49 | 37 | 24 | 12.12 | 108.6 | 7.77
2009-05-26 | 12:01 | 10.84 | 37 | 24 | 12.03 | 108.7 | 7.77
2009-05-26 | 12:16 | 11.18 | 37 | 24 | 11.92 | 108.6 | 7.77
2009-05-26 | 12:31 | 11.54 | 37 | 24 | 11.77 | 108.1 | 7.78
2009-05-26 | 12:46 | 11.92 | 37 | 24 | 11.69 | 108.3 | 7.79
2009-05-26 | 13:00 | 12.31 | 37 | 24 | 11.58 | 108.2 | 7.79
2009-05-26 | 13:15 | 12.7 | 37 | 24 | 11.4 | 107.5 | 7.79
2009-05-26 | 13:30 | 13.08 | 37 | 24 | 11.31 | 107.5 | 7.8
2009-05-26 | 13:45 | 13.46 | 37 | 24 | 11.2 | 107.4 | 7.8
2009-05-26 | 14:00 | 13.82 | 38 | 24 | 11.04 | 106.7 | 7.81
2009-05-26 | 14:15 | 14.13 | 38 | 24 | 10.91 | 106.1 | 7.82
2009-05-26 | 14:30 | 14.42 | 38 | 25 | 10.84 | 106.1 | 7.82
2009-05-26 | 14:45 | 14.69 | 38 | 25 | 10.76 | 105.9 | 7.83
2009-05-26 | 15:00 | 14.92 | 38 | 25 | 10.67 | 105.7 | 7.83
2009-05-26 | 15:15 | 15.12 | 38 | 25 | 10.59 | 105.3 | 7.83
2009-05-26 | 15:30 | 15.28 | 38 | 25 | 10.52 | 104.9 | 7.83
2009-05-26 | 15:45 | 15.42 | 38 | 25 | 10.41 | 104.1 | 7.84
2009-05-26 | 16:00 | 15.51 | 38 | 25 | 10.33 | 103.6 | 7.84
2009-05-26 | 16:15 | 15.6 | 38 | 25 | 10.32 | 103.7 | 7.84
2009-05-26 | 16:30 | 15.64 | 38 | 25 | 10.29 | 103.4 | 7.83
2009-05-26 | 16:45 | 15.68 | 38 | 25 | 10.27 | 103.3 | 7.83
2009-05-26 | 17:00 | 15.7 | 39 | 25 | 10.24 | 103.1 | 7.83
2009-05-26 | 17:15 | 15.7 | 39 | 25 | 10.21 | 102.8 | 7.83
2009-05-26 | 17:30 | 15.69 | 39 | 25 | 10.19 | 102.6 | 7.82
2009-05-26 | 17:45 | 15.67 | 39 | 25 | 10.13 | 101.9 | 7.82
2009-05-26 | 18:00 | 15.63 | 39 | 25 | 10.1 | 101.5 | 7.82
2009-05-26 | 18:15 | 15.59 | 39 | 25 | 10.13 | 101.8 | 7.81
2009-05-26 | 18:30 | 15.53 | 39 | 25 | 10.1 | 101.3 | 7.8
2009-05-26 | 18:45 | 15.5 | 39 | 25 | 10.08 | 101 | 7.8
2009-05-26 | 19:00 | 15.45 | 39 | 25 | 10.07 | 100.8 | 7.79
2009-05-26 | 19:15 | 15.42 | 39 | 26 | 10.06 | 100.7 | 7.79
2009-05-26 | 19:30 | 15.35 | 39 | 26 | 10.11 | 101 | 7.78
2009-05-26 | 19:45 | 15.27 | 39 | 26 | 10.13 | 101.1 | 7.78
2009-05-26 | 20:00 | 15.17 | 39 | 26 | 10.16 | 101.1 | 7.78
2009-05-26 | 20:15 | 15.06 | 40 | 26 | 10.17 | 101 | 7.77
2009-05-26 | 20:30 | 14.93 | 40 | 26 | 10.19 | 100.9 | 7.76
2009-05-26 | 20:45 | 14.77 | 40 | 26 | 10.22 | 100.8 | 7.76
2009-05-26 | 21:00 | 14.62 | 40 | 26 | 10.25 | 100.8 | 7.75
2009-05-26 | 21:15 | 14.49 | 40 | 26 | 10.22 | 100.2 | 7.74
2009-05-26 | 21:30 | 14.36 | 40 | 26 | 10.23 | 100 | 7.74
2009-05-26 | 21:45 | 14.24 | 40 | 26 | 10.29 | 100.3 | 7.74
2009-05-26 | 22:00 | 14.12 | 40 | 26 | 10.29 | 100.1 | 7.74
2009-05-26 | 22:15 | 13.99 | 40 | 26 | 10.35 | 100.4 | 7.74
2009-05-26 | 22:30 | 13.87 | 41 | 26 | 10.39 | 100.6 | 7.73
2009-05-26 | 22:45 | 13.75 | 41 | 26 | 10.39 | 100.3 | 7.74
2009-05-26 | 23:00 | 13.62 | 41 | 26 | 10.45 | 100.6 | 7.74
2009-05-26 | 23:15 | 13.5 | 41 | 26 | 10.45 | 100.3 | 7.74
2009-05-26 | 23:30 | 13.38 | 41 | 27 | 10.52 | 100.7 | 7.74
2009-05-26 | 23:45 | 13.26 | 41 | 27 | 10.57 | 100.9 | 7.74
2009-05-27 | 0:00 | 13.14 | 41 | 27 | 10.61 | 101 | 7.74
2009-05-27 | 0:15 | 13.02 | 41 | 27 | 10.65 | 101.1 | 7.74
2009-05-27 | 0:30 | 12.88 | 41 | 27 | 10.69 | 101.2 | 7.73
2009-05-27 | 0:45 | 12.76 | 41 | 27 | 10.68 | 100.8 | 7.73
2009-05-27 | 1:01 | 12.64 | 41 | 27 | 10.74 | 101.2 | 7.74
2009-05-27 | 1:16 | 12.51 | 41 | 27 | 10.75 | 100.9 | 7.74
2009-05-27 | 1:31 | 12.37 | 41 | 27 | 10.82 | 101.3 | 7.74
2009-05-27 | 1:46 | 12.22 | 41 | 27 | 10.88 | 101.5 | 7.73
2009-05-27 | 2:01 | 12.06 | 41 | 26 | 10.89 | 101.2 | 7.73
2009-05-27 | 2:16 | 11.9 | 40 | 26 | 10.91 | 101 | 7.72
2009-05-27 | 2:31 | 11.74 | 40 | 26 | 10.95 | 101 | 7.72
2009-05-27 | 2:46 | 11.57 | 40 | 26 | 11.05 | 101.5 | 7.71
2009-05-27 | 3:01 | 11.4 | 39 | 26 | 11.12 | 101.8 | 7.71
2009-05-27 | 3:16 | 11.23 | 39 | 25 | 11.18 | 102 | 7.7
2009-05-27 | 3:31 | 11.06 | 38 | 25 | 11.24 | 102.1 | 7.69
2009-05-27 | 3:46 | 10.89 | 38 | 25 | 11.25 | 101.7 | 7.69
2009-05-27 | 4:01 | 10.72 | 38 | 25 | 11.33 | 102.1 | 7.68
2009-05-27 | 4:16 | 10.56 | 37 | 24 | 11.34 | 101.8 | 7.68
2009-05-27 | 4:31 | 10.41 | 37 | 24 | 11.43 | 102.2 | 7.67
2009-05-27 | 4:46 | 10.26 | 37 | 24 | 11.49 | 102.4 | 7.67
2009-05-27 | 5:01 | 10.12 | 37 | 24 | 11.54 | 102.5 | 7.67
2009-05-27 | 5:16 | 9.98 | 37 | 24 | 11.54 | 102.2 | 7.66
2009-05-27 | 5:31 | 9.85 | 36 | 24 | 11.56 | 102.1 | 7.66
2009-05-27 | 5:46 | 9.71 | 36 | 24 | 11.6 | 102.1 | 7.66
2009-05-27 | 6:01 | 9.56 | 36 | 23 | 11.71 | 102.7 | 7.65
2009-05-27 | 6:16 | 9.43 | 36 | 23 | 11.79 | 103 | 7.65
2009-05-27 | 6:31 | 9.29 | 36 | 23 | 11.8 | 102.8 | 7.65
2009-05-27 | 6:46 | 9.17 | 36 | 23 | 11.85 | 102.9 | 7.65
2009-05-27 | 7:01 | 9.06 | 36 | 23 | 11.94 | 103.5 | 7.65
2009-05-27 | 7:16 | 8.96 | 36 | 23 | 11.96 | 103.4 | 7.65
2009-05-27 | 7:31 | 8.87 | 36 | 23 | 12.05 | 103.9 | 7.65
2009-05-27 | 7:46 | 8.8 | 35 | 23 | 12.12 | 104.4 | 7.66
2009-05-27 | 8:01 | 8.73 | 35 | 23 | 12.18 | 104.6 | 7.66
2009-05-27 | 8:16 | 8.67 | 35 | 23 | 12.25 | 105.2 | 7.67
2009-05-27 | 8:31 | 8.62 | 35 | 23 | 12.29 | 105.4 | 7.67
2009-05-27 | 8:46 | 8.58 | 35 | 23 | 12.44 | 106.5 | 7.67
2009-05-27 | 9:01 | 8.56 | 35 | 23 | 12.5 | 107 | 7.68
2009-05-27 | 9:16 | 8.58 | 35 | 23 | 12.53 | 107.3 | 7.69
2009-05-27 | 9:31 | 8.65 | 35 | 23 | 12.53 | 107.5 | 7.69
2009-05-27 | 9:46 | 8.75 | 35 | 23 | 12.43 | 106.9 | 7.69
2009-05-27 | 10:01 | 8.91 | 35 | 23 | 12.33 | 106.4 | 7.7
2009-05-27 | 10:16 | 9.12 | 35 | 23 | 12.3 | 106.7 | 7.7
2009-05-27 | 10:31 | 9.35 | 35 | 23 | 12.17 | 106.2 | 7.71
2009-05-27 | 10:46 | 9.61 | 35 | 23 | 12.11 | 106.3 | 7.72
2009-05-27 | 11:01 | 9.9 | 35 | 23 | 12.11 | 107 | 7.72
2009-05-27 | 11:16 | 10.18 | 35 | 23 | 12.08 | 107.5 | 7.73
2009-05-27 | 11:31 | 10.5 | 35 | 23 | 12.02 | 107.8 | 7.73
2009-05-27 | 11:46 | 10.83 | 35 | 23 | 11.99 | 108.3 | 7.74
2009-05-27 | 12:01 | 11.16 | 35 | 23 | 11.93 | 108.6 | 7.74
2009-05-27 | 12:16 | 11.51 | 35 | 23 | 11.79 | 108.2 | 7.75
2009-05-27 | 12:31 | 11.88 | 36 | 23 | 11.66 | 107.9 | 7.76
2009-05-27 | 12:46 | 12.25 | 36 | 23 | 11.6 | 108.2 | 7.76
2009-05-27 | 13:00 | 12.63 | 36 | 23 | 11.45 | 107.7 | 7.76
2009-05-27 | 13:15 | 13.01 | 36 | 23 | 11.36 | 107.8 | 7.77
2009-05-27 | 13:30 | 13.4 | 36 | 23 | 11.19 | 107.2 | 7.78
2009-05-27 | 13:45 | 13.76 | 36 | 23 | 11.11 | 107.2 | 7.78
2009-05-27 | 14:00 | 14.1 | 36 | 23 | 10.96 | 106.6 | 7.78
2009-05-27 | 14:15 | 14.41 | 36 | 23 | 10.83 | 106 | 7.79
2009-05-27 | 14:30 | 14.69 | 36 | 23 | 10.75 | 105.9 | 7.8
2009-05-27 | 14:45 | 14.95 | 36 | 23 | 10.67 | 105.7 | 7.79
2009-05-27 | 15:00 | 15.18 | 36 | 24 | 10.59 | 105.5 | 7.8
2009-05-27 | 15:15 | 15.38 | 36 | 24 | 10.51 | 105.1 | 7.8
2009-05-27 | 15:30 | 15.54 | 36 | 24 | 10.39 | 104.3 | 7.8
2009-05-27 | 15:45 | 15.68 | 36 | 24 | 10.3 | 103.7 | 7.81
2009-05-27 | 16:00 | 15.78 | 37 | 24 | 10.29 | 103.8 | 7.8
2009-05-27 | 16:15 | 15.86 | 37 | 24 | 10.25 | 103.6 | 7.8
2009-05-27 | 16:30 | 15.91 | 37 | 24 | 10.22 | 103.4 | 7.8
2009-05-27 | 16:45 | 15.95 | 37 | 24 | 10.19 | 103.1 | 7.8
2009-05-27 | 17:00 | 15.95 | 37 | 24 | 10.12 | 102.4 | 7.8
2009-05-27 | 17:15 | 15.94 | 37 | 24 | 10.08 | 101.9 | 7.79
2009-05-27 | 17:30 | 15.94 | 37 | 24 | 10.09 | 102.1 | 7.79
2009-05-27 | 17:45 | 15.95 | 37 | 24 | 10.1 | 102.2 | 7.79
2009-05-27 | 18:00 | 15.93 | 37 | 24 | 10.1 | 102.2 | 7.79
2009-05-27 | 18:15 | 15.93 | 37 | 24 | 10.09 | 102.1 | 7.79
2009-05-27 | 18:30 | 15.91 | 37 | 24 | 10.08 | 102 | 7.78
2009-05-27 | 18:45 | 15.9 | 37 | 24 | 10.08 | 101.9 | 7.79
2009-05-27 | 19:00 | 15.89 | 38 | 24 | 10.07 | 101.8 | 7.78
2009-05-27 | 19:15 | 15.86 | 38 | 24 | 10.06 | 101.6 | 7.78
2009-05-27 | 19:30 | 15.81 | 38 | 25 | 10.06 | 101.5 | 7.77
2009-05-27 | 19:45 | 15.74 | 38 | 25 | 10.01 | 100.8 | 7.76
2009-05-27 | 20:00 | 15.66 | 38 | 25 | 9.99 | 100.5 | 7.76
2009-05-27 | 20:15 | 15.54 | 38 | 25 | 9.99 | 100.2 | 7.75
2009-05-27 | 20:30 | 15.4 | 38 | 25 | 10 | 100 | 7.74
2009-05-27 | 20:45 | 15.25 | 38 | 25 | 10.03 | 100 | 7.74
2009-05-27 | 21:00 | 15.12 | 38 | 25 | 10.09 | 100.3 | 7.73
2009-05-27 | 21:15 | 14.98 | 38 | 25 | 10.08 | 100 | 7.73
2009-05-27 | 21:30 | 14.87 | 38 | 25 | 10.15 | 100.3 | 7.72
2009-05-27 | 21:45 | 14.77 | 39 | 25 | 10.18 | 100.5 | 7.72
2009-05-27 | 22:00 | 14.67 | 39 | 25 | 10.16 | 100.1 | 7.72
2009-05-27 | 22:15 | 14.56 | 39 | 25 | 10.22 | 100.4 | 7.72
2009-05-27 | 22:30 | 14.47 | 39 | 25 | 10.22 | 100.1 | 7.71
2009-05-27 | 22:45 | 14.36 | 39 | 25 | 10.22 | 100 | 7.71
2009-05-27 | 23:00 | 14.24 | 39 | 25 | 10.29 | 100.4 | 7.71
2009-05-27 | 23:15 | 14.13 | 39 | 25 | 10.34 | 100.6 | 7.71
2009-05-27 | 23:30 | 14 | 39 | 25 | 10.33 | 100.3 | 7.71
2009-05-27 | 23:45 | 13.87 | 39 | 25 | 10.41 | 100.7 | 7.71
2009-05-28 | 0:00 | 13.73 | 39 | 26 | 10.41 | 100.4 | 7.72
2009-05-28 | 0:15 | 13.58 | 39 | 26 | 10.44 | 100.3 | 7.71
2009-05-28 | 0:30 | 13.44 | 39 | 26 | 10.52 | 100.8 | 7.71
2009-05-28 | 0:45 | 13.28 | 39 | 25 | 10.57 | 101 | 7.71
2009-05-28 | 1:01 | 13.13 | 39 | 25 | 10.58 | 100.7 | 7.71
2009-05-28 | 1:16 | 12.96 | 39 | 25 | 10.66 | 101 | 7.71
2009-05-28 | 1:31 | 12.78 | 39 | 25 | 10.72 | 101.3 | 7.69
2009-05-28 | 1:46 | 12.59 | 38 | 25 | 10.73 | 101 | 7.69
2009-05-28 | 2:01 | 12.41 | 38 | 25 | 10.82 | 101.3 | 7.68
2009-05-28 | 2:16 | 12.22 | 38 | 24 | 10.89 | 101.5 | 7.68
2009-05-28 | 2:31 | 12.03 | 37 | 24 | 10.91 | 101.3 | 7.66
2009-05-28 | 2:46 | 11.85 | 37 | 24 | 10.99 | 101.6 | 7.66
2009-05-28 | 3:01 | 11.66 | 36 | 24 | 11 | 101.3 | 7.65
2009-05-28 | 3:16 | 11.49 | 36 | 24 | 11.09 | 101.7 | 7.65
2009-05-28 | 3:31 | 11.33 | 36 | 23 | 11.16 | 101.9 | 7.64
2009-05-28 | 3:46 | 11.18 | 36 | 23 | 11.21 | 102.1 | 7.63
2009-05-28 | 4:01 | 11.03 | 35 | 23 | 11.21 | 101.7 | 7.63
2009-05-28 | 4:16 | 10.89 | 35 | 23 | 11.29 | 102.1 | 7.62
2009-05-28 | 4:31 | 10.76 | 35 | 23 | 11.34 | 102.3 | 7.62
2009-05-28 | 4:46 | 10.63 | 35 | 23 | 11.33 | 101.9 | 7.62
2009-05-28 | 5:01 | 10.49 | 35 | 23 | 11.41 | 102.2 | 7.62
2009-05-28 | 5:16 | 10.36 | 35 | 22 | 11.47 | 102.4 | 7.6
2009-05-28 | 5:31 | 10.22 | 34 | 22 | 11.47 | 102.1 | 7.6
2009-05-28 | 5:46 | 10.07 | 34 | 22 | 11.5 | 102 | 7.6
2009-05-28 | 6:01 | 9.92 | 34 | 22 | 11.6 | 102.6 | 7.6
2009-05-28 | 6:16 | 9.78 | 34 | 22 | 11.62 | 102.4 | 7.61
2009-05-28 | 6:31 | 9.64 | 34 | 22 | 11.67 | 102.5 | 7.61
2009-05-28 | 6:46 | 9.51 | 34 | 22 | 11.72 | 102.7 | 7.62
2009-05-28 | 7:01 | 9.41 | 34 | 22 | 11.78 | 102.9 | 7.62
2009-05-28 | 7:16 | 9.32 | 34 | 22 | 11.84 | 103.2 | 7.63
2009-05-28 | 7:31 | 9.25 | 34 | 22 | 11.94 | 103.9 | 7.63
2009-05-28 | 7:46 | 9.17 | 34 | 22 | 11.96 | 103.9 | 7.64
2009-05-28 | 8:01 | 9.12 | 34 | 22 | 12.06 | 104.6 | 7.65
2009-05-28 | 8:16 | 9.06 | 34 | 22 | 12.12 | 105 | 7.65
2009-05-28 | 8:31 | 9.03 | 34 | 22 | 12.14 | 105.1 | 7.66
2009-05-28 | 8:46 | 9 | 34 | 22 | 12.25 | 105.9 | 7.67
2009-05-28 | 9:01 | 9 | 34 | 22 | 12.34 | 106.7 | 7.67
2009-05-28 | 9:16 | 9.04 | 34 | 22 | 12.41 | 107.4 | 7.67
2009-05-28 | 9:31 | 9.09 | 34 | 22 | 12.38 | 107.3 | 7.68
2009-05-28 | 9:46 | 9.17 | 34 | 22 | 12.32 | 107 | 7.67
2009-05-28 | 10:01 | 9.29 | 34 | 22 | 12.21 | 106.4 | 7.68
2009-05-28 | 10:16 | 9.45 | 34 | 22 | 12.17 | 106.4 | 7.68
2009-05-28 | 10:31 | 9.66 | 34 | 22 | 12.11 | 106.4 | 7.68
2009-05-28 | 10:46 | 9.92 | 34 | 22 | 12.07 | 106.8 | 7.69
2009-05-28 | 11:01 | 10.19 | 34 | 22 | 12.08 | 107.5 | 7.7
2009-05-28 | 11:16 | 10.48 | 34 | 22 | 12.05 | 108 | 7.7
2009-05-28 | 11:31 | 10.8 | 34 | 22 | 11.94 | 107.8 | 7.72
2009-05-28 | 11:46 | 11.12 | 34 | 22 | 11.9 | 108.3 | 7.72
2009-05-28 | 12:01 | 11.46 | 34 | 22 | 11.83 | 108.4 | 7.72
2009-05-28 | 12:16 | 11.81 | 34 | 22 | 11.68 | 107.9 | 7.73
2009-05-28 | 12:31 | 12.16 | 34 | 22 | 11.6 | 108 | 7.74
2009-05-28 | 12:46 | 12.47 | 34 | 22 | 11.43 | 107.2 | 7.74
2009-05-28 | 13:00 | 12.85 | 34 | 22 | 11.37 | 107.5 | 7.75
2009-05-28 | 13:15 | 13.24 | 34 | 22 | 11.26 | 107.4 | 7.76
2009-05-28 | 13:30 | 13.6 | 34 | 22 | 11.1 | 106.7 | 7.76
2009-05-28 | 13:45 | 13.99 | 35 | 22 | 11.01 | 106.8 | 7.76
2009-05-28 | 14:00 | 14.33 | 35 | 22 | 10.85 | 106.1 | 7.77
2009-05-28 | 14:15 | 14.64 | 35 | 23 | 10.72 | 105.5 | 7.78
2009-05-28 | 14:30 | 14.93 | 35 | 23 | 10.6 | 105 | 7.79
2009-05-28 | 14:45 | 15.2 | 35 | 23 | 10.55 | 105 | 7.79
2009-05-28 | 15:00 | 15.45 | 35 | 23 | 10.42 | 104.3 | 7.79
2009-05-28 | 15:15 | 15.67 | 35 | 23 | 10.37 | 104.3 | 7.79
2009-05-28 | 15:30 | 15.86 | 35 | 23 | 10.3 | 104 | 7.8
2009-05-28 | 15:45 | 16.02 | 35 | 23 | 10.24 | 103.8 | 7.8
2009-05-28 | 16:00 | 16.15 | 35 | 23 | 10.14 | 103 | 7.79
2009-05-28 | 16:15 | 16.26 | 35 | 23 | 10.11 | 103 | 7.8
2009-05-28 | 16:30 | 16.34 | 35 | 23 | 10.08 | 102.9 | 7.8
2009-05-28 | 16:45 | 16.4 | 35 | 23 | 10.04 | 102.5 | 7.79
2009-05-28 | 17:00 | 16.43 | 36 | 23 | 10.02 | 102.4 | 7.79
2009-05-28 | 17:15 | 16.45 | 36 | 23 | 9.98 | 102.1 | 7.78
2009-05-28 | 17:30 | 16.46 | 36 | 23 | 9.92 | 101.4 | 7.79
2009-05-28 | 17:45 | 16.47 | 36 | 23 | 9.88 | 101.1 | 7.78
2009-05-28 | 18:00 | 16.48 | 36 | 23 | 9.85 | 100.9 | 7.77
2009-05-28 | 18:15 | 16.49 | 36 | 23 | 9.88 | 101.2 | 7.77
2009-05-28 | 18:30 | 16.5 | 36 | 23 | 9.89 | 101.2 | 7.77
2009-05-28 | 18:45 | 16.51 | 36 | 23 | 9.87 | 101.1 | 7.76
2009-05-28 | 19:00 | 16.51 | 36 | 24 | 9.87 | 101.1 | 7.76
2009-05-28 | 19:15 | 16.49 | 36 | 24 | 9.86 | 100.9 | 7.76
2009-05-28 | 19:30 | 16.44 | 36 | 24 | 9.86 | 100.8 | 7.76
2009-05-28 | 19:45 | 16.36 | 36 | 24 | 9.86 | 100.7 | 7.75
2009-05-28 | 20:00 | 16.22 | 37 | 24 | 9.84 | 100.1 | 7.74
2009-05-28 | 20:15 | 16.09 | 37 | 24 | 9.88 | 100.3 | 7.73
2009-05-28 | 20:30 | 15.96 | 37 | 24 | 9.88 | 100 | 7.73
2009-05-28 | 20:45 | 15.85 | 37 | 24 | 9.91 | 100.1 | 7.72
2009-05-28 | 21:00 | 15.74 | 37 | 24 | 9.94 | 100.2 | 7.71
2009-05-28 | 21:15 | 15.65 | 37 | 24 | 9.96 | 100.1 | 7.7
2009-05-28 | 21:30 | 15.55 | 37 | 24 | 9.94 | 99.8 | 7.7
2009-05-28 | 21:45 | 15.45 | 37 | 24 | 9.94 | 99.5 | 7.69
2009-05-28 | 22:00 | 15.36 | 37 | 24 | 9.99 | 99.9 | 7.69
2009-05-28 | 22:15 | 15.27 | 37 | 24 | 10.03 | 100.1 | 7.69
2009-05-28 | 22:30 | 15.18 | 37 | 24 | 10.03 | 99.8 | 7.69
2009-05-28 | 22:45 | 15.08 | 38 | 24 | 10.03 | 99.7 | 7.69
2009-05-28 | 23:00 | 14.97 | 38 | 24 | 10.09 | 100 | 7.69
2009-05-28 | 23:15 | 14.86 | 38 | 24 | 10.09 | 99.8 | 7.69
2009-05-28 | 23:30 | 14.76 | 38 | 25 | 10.11 | 99.7 | 7.69
2009-05-28 | 23:45 | 14.65 | 38 | 25 | 10.18 | 100.2 | 7.68
2009-05-29 | 0:00 | 14.53 | 38 | 24 | 10.23 | 100.4 | 7.68
2009-05-29 | 0:15 | 14.42 | 38 | 24 | 10.27 | 100.6 | 7.68
2009-05-29 | 0:30 | 14.32 | 37 | 24 | 10.3 | 100.7 | 7.68
2009-05-29 | 0:45 | 14.21 | 37 | 24 | 10.34 | 100.8 | 7.68
2009-05-29 | 1:01 | 14.09 | 37 | 24 | 10.32 | 100.3 | 7.67
2009-05-29 | 1:16 | 13.96 | 37 | 24 | 10.38 | 100.7 | 7.67
2009-05-29 | 1:31 | 13.82 | 36 | 24 | 10.39 | 100.4 | 7.66
2009-05-29 | 1:46 | 13.66 | 36 | 23 | 10.41 | 100.3 | 7.66
2009-05-29 | 2:01 | 13.5 | 36 | 23 | 10.45 | 100.3 | 7.65
2009-05-29 | 2:16 | 13.35 | 35 | 23 | 10.48 | 100.3 | 7.64
2009-05-29 | 2:31 | 13.21 | 35 | 23 | 10.52 | 100.4 | 7.63
2009-05-29 | 2:46 | 13.09 | 35 | 22 | 10.55 | 100.4 | 7.63
2009-05-29 | 3:01 | 12.96 | 34 | 22 | 10.63 | 100.8 | 7.62
2009-05-29 | 3:16 | 12.85 | 34 | 22 | 10.64 | 100.6 | 7.61
2009-05-29 | 3:31 | 12.74 | 34 | 22 | 10.66 | 100.6 | 7.61
2009-05-29 | 3:46 | 12.62 | 34 | 22 | 10.73 | 101 | 7.61
2009-05-29 | 4:01 | 12.5 | 33 | 22 | 10.74 | 100.8 | 7.61
2009-05-29 | 4:16 | 12.37 | 33 | 22 | 10.82 | 101.2 | 7.6
2009-05-29 | 4:31 | 12.24 | 33 | 21 | 10.88 | 101.5 | 7.59
2009-05-29 | 4:46 | 12.1 | 33 | 21 | 10.88 | 101.2 | 7.59
2009-05-29 | 5:01 | 11.95 | 33 | 21 | 10.91 | 101.2 | 7.59
2009-05-29 | 5:16 | 11.8 | 32 | 21 | 11 | 101.6 | 7.58
2009-05-29 | 5:31 | 11.63 | 32 | 21 | 11.07 | 101.8 | 7.58
2009-05-29 | 5:46 | 11.47 | 32 | 21 | 11.08 | 101.6 | 7.57
2009-05-29 | 6:01 | 11.3 | 32 | 21 | 11.17 | 102 | 7.58
2009-05-29 | 6:16 | 11.13 | 32 | 21 | 11.19 | 101.8 | 7.57
2009-05-29 | 6:31 | 10.97 | 32 | 21 | 11.28 | 102.2 | 7.57
2009-05-29 | 6:46 | 10.82 | 32 | 21 | 11.35 | 102.5 | 7.57
2009-05-29 | 7:01 | 10.68 | 32 | 21 | 11.37 | 102.3 | 7.57
2009-05-29 | 7:16 | 10.54 | 32 | 21 | 11.4 | 102.3 | 7.57
2009-05-29 | 7:31 | 10.43 | 32 | 21 | 11.5 | 102.9 | 7.58
2009-05-29 | 7:46 | 10.32 | 32 | 21 | 11.57 | 103.2 | 7.58
2009-05-29 | 8:01 | 10.22 | 31 | 20 | 11.63 | 103.5 | 7.58
2009-05-29 | 8:16 | 10.15 | 31 | 20 | 11.68 | 103.8 | 7.6
2009-05-29 | 8:31 | 10.08 | 31 | 20 | 11.76 | 104.4 | 7.58
2009-05-29 | 8:46 | 10.03 | 31 | 20 | 11.84 | 105 | 7.6
2009-05-29 | 9:01 | 10.01 | 31 | 20 | 11.87 | 105.2 | 7.61
2009-05-29 | 9:16 | 10.04 | 31 | 20 | 11.95 | 105.9 | 7.62
2009-05-29 | 9:31 | 10.09 | 31 | 20 | 11.93 | 105.9 | 7.63
2009-05-29 | 9:46 | 10.18 | 31 | 20 | 11.84 | 105.3 | 7.64
2009-05-29 | 10:01 | 10.32 | 31 | 20 | 11.76 | 105 | 7.64
2009-05-29 | 10:16 | 10.49 | 31 | 20 | 11.75 | 105.3 | 7.65
2009-05-29 | 10:31 | 10.7 | 32 | 20 | 11.7 | 105.3 | 7.66
2009-05-29 | 10:46 | 10.94 | 32 | 20 | 11.62 | 105.2 | 7.67
2009-05-29 | 11:01 | 11.18 | 32 | 21 | 11.61 | 105.7 | 7.67
2009-05-29 | 11:16 | 11.43 | 32 | 21 | 11.59 | 106.1 | 7.68
2009-05-29 | 11:31 | 11.7 | 32 | 21 | 11.54 | 106.4 | 7.68
2009-05-29 | 11:46 | 11.99 | 32 | 21 | 11.48 | 106.5 | 7.69
2009-05-29 | 12:01 | 12.28 | 32 | 21 | 11.45 | 107 | 7.69
2009-05-29 | 12:16 | 12.59 | 32 | 21 | 11.39 | 107.2 | 7.7
2009-05-29 | 12:31 | 12.8 | 32 | 21 | 11.3 | 106.8 | 7.71
2009-05-29 | 12:46 | 13.03 | 32 | 21 | 11.2 | 106.4 | 7.71
2009-05-29 | 13:00 | 13.05 | 32 | 21 | 11.07 | 105.2 | 7.71
2009-05-29 | 13:15 | 13.07 | 32 | 21 | 11.01 | 104.6 | 7.71
2009-05-29 | 13:30 | 13.21 | 32 | 21 | 10.98 | 104.7 | 7.72
2009-05-29 | 13:45 | 13.5 | 32 | 21 | 10.94 | 105 | 7.72
2009-05-29 | 14:00 | 13.8 | 32 | 21 | 10.94 | 105.7 | 7.72
2009-05-29 | 14:15 | 14.1 | 32 | 21 | 10.88 | 105.8 | 7.73
2009-05-29 | 14:30 | 14.27 | 32 | 21 | 10.8 | 105.4 | 7.73
2009-05-29 | 14:45 | 14.2 | 32 | 21 | 10.75 | 104.7 | 7.72
2009-05-29 | 15:00 | 14.15 | 32 | 21 | 10.64 | 103.6 | 7.71
2009-05-29 | 15:15 | 14.37 | 32 | 21 | 10.63 | 104 | 7.72
2009-05-29 | 15:30 | 14.59 | 32 | 21 | 10.61 | 104.3 | 7.72
2009-05-29 | 15:45 | 14.64 | 32 | 21 | 10.57 | 104 | 7.73
2009-05-29 | 16:00 | 14.61 | 32 | 21 | 10.57 | 103.9 | 7.73
2009-05-29 | 16:15 | 14.62 | 32 | 21 | 10.51 | 103.4 | 7.72
2009-05-29 | 16:30 | 14.69 | 32 | 21 | 10.51 | 103.5 | 7.72
2009-05-29 | 16:45 | 14.76 | 32 | 21 | 10.42 | 102.8 | 7.73
2009-05-29 | 17:00 | 14.69 | 32 | 21 | 10.43 | 102.8 | 7.72
2009-05-29 | 17:15 | 14.6 | 33 | 21 | 10.42 | 102.4 | 7.71
2009-05-29 | 17:30 | 14.68 | 33 | 21 | 10.38 | 102.2 | 7.71
2009-05-29 | 17:45 | 14.83 | 33 | 21 | 10.39 | 102.7 | 7.72
2009-05-29 | 18:00 | 14.94 | 33 | 21 | 10.35 | 102.5 | 7.71
2009-05-29 | 18:15 | 15.02 | 33 | 21 | 10.27 | 101.9 | 7.71
2009-05-29 | 18:30 | 15.06 | 33 | 21 | 10.26 | 101.8 | 7.7
2009-05-29 | 18:45 | 15.05 | 33 | 21 | 10.25 | 101.8 | 7.7
2009-05-29 | 19:00 | 15 | 33 | 21 | 10.21 | 101.2 | 7.69
2009-05-29 | 19:15 | 14.98 | 33 | 21 | 10.24 | 101.6 | 7.69
2009-05-29 | 19:30 | 14.93 | 33 | 21 | 10.21 | 101.1 | 7.68
2009-05-29 | 19:45 | 14.91 | 33 | 21 | 10.2 | 101 | 7.68
2009-05-29 | 20:00 | 14.92 | 33 | 21 | 10.22 | 101.2 | 7.69
2009-05-29 | 20:15 | 14.93 | 33 | 21 | 10.23 | 101.3 | 7.68
2009-05-29 | 20:30 | 14.97 | 33 | 21 | 10.22 | 101.3 | 7.68
2009-05-29 | 20:45 | 15.04 | 33 | 21 | 10.2 | 101.2 | 7.67
2009-05-29 | 21:00 | 15.09 | 33 | 21 | 10.18 | 101.1 | 7.67
2009-05-29 | 21:15 | 15.03 | 33 | 21 | 10.19 | 101.1 | 7.66
2009-05-29 | 21:30 | 14.91 | 33 | 21 | 10.22 | 101.2 | 7.63
2009-05-29 | 21:45 | 14.76 | 33 | 21 | 10.26 | 101.2 | 7.64
2009-05-29 | 22:00 | 14.59 | 33 | 21 | 10.29 | 101.2 | 7.63
2009-05-29 | 22:15 | 14.44 | 32 | 21 | 10.29 | 100.8 | 7.62
2009-05-29 | 22:30 | 14.31 | 32 | 21 | 10.35 | 101.1 | 7.62
2009-05-29 | 22:45 | 14.24 | 32 | 21 | 10.33 | 100.7 | 7.6
2009-05-29 | 23:00 | 14.17 | 32 | 21 | 10.33 | 100.6 | 7.6
2009-05-29 | 23:15 | 14.05 | 31 | 20 | 10.35 | 100.6 | 7.59
2009-05-29 | 23:30 | 13.89 | 31 | 20 | 10.44 | 101 | 7.59
2009-05-29 | 23:45 | 13.73 | 31 | 20 | 10.44 | 100.7 | 7.58
2009-05-30 | 0:00 | 13.57 | 31 | 20 | 10.54 | 101.3 | 7.57
2009-05-30 | 0:15 | 13.39 | 30 | 20 | 10.59 | 101.4 | 7.57
2009-05-30 | 0:30 | 13.22 | 30 | 19 | 10.6 | 101.1 | 7.55
2009-05-30 | 0:45 | 13.05 | 30 | 19 | 10.69 | 101.5 | 7.55
2009-05-30 | 1:01 | 12.87 | 30 | 19 | 10.75 | 101.7 | 7.55
2009-05-30 | 1:16 | 12.67 | 29 | 19 | 10.81 | 101.9 | 7.53
2009-05-30 | 1:31 | 12.47 | 29 | 19 | 10.84 | 101.6 | 7.53
2009-05-30 | 1:46 | 12.28 | 29 | 19 | 10.92 | 102 | 7.52
2009-05-30 | 2:01 | 12.1 | 29 | 19 | 10.98 | 102.2 | 7.52
2009-05-30 | 2:16 | 11.91 | 28 | 18 | 11 | 101.8 | 7.53
2009-05-30 | 2:31 | 11.71 | 28 | 18 | 11.04 | 101.8 | 7.51
2009-05-30 | 2:46 | 11.49 | 28 | 18 | 11.16 | 102.3 | 7.51
2009-05-30 | 3:01 | 11.26 | 28 | 18 | 11.19 | 102.1 | 7.49
2009-05-30 | 3:16 | 11.04 | 28 | 18 | 11.29 | 102.5 | 7.5
2009-05-30 | 3:31 | 10.82 | 27 | 18 | 11.32 | 102.3 | 7.49
2009-05-30 | 3:46 | 10.61 | 27 | 18 | 11.43 | 102.7 | 7.5
2009-05-30 | 4:01 | 10.4 | 27 | 18 | 11.51 | 103 | 7.48
2009-05-30 | 4:16 | 10.19 | 28 | 18 | 11.42 | 101.6 | 7.48
2009-05-30 | 4:31 | 9.99 | 27 | 18 | 11.66 | 103.2 | 7.48
2009-05-30 | 4:46 | 9.79 | 27 | 18 | 11.73 | 103.4 | 7.49
2009-05-30 | 5:01 | 9.59 | 27 | 18 | 11.74 | 103 | 7.49
2009-05-30 | 5:16 | 9.41 | 27 | 18 | 11.84 | 103.4 | 7.48
2009-05-30 | 5:31 | 9.24 | 27 | 18 | 11.91 | 103.6 | 7.47
2009-05-30 | 5:46 | 9.08 | 27 | 18 | 11.97 | 103.7 | 7.47
2009-05-30 | 6:01 | 8.93 | 27 | 18 | 11.98 | 103.4 | 7.46
2009-05-30 | 6:16 | 8.79 | 27 | 18 | 12.06 | 103.8 | 7.48
2009-05-30 | 6:31 | 8.66 | 27 | 18 | 12.12 | 104 | 7.48
2009-05-30 | 6:46 | 8.55 | 27 | 18 | 12.11 | 103.7 | 7.48
2009-05-30 | 7:01 | 8.44 | 27 | 18 | 12.13 | 103.5 | 7.47
2009-05-30 | 7:16 | 8.34 | 27 | 18 | 12.16 | 103.5 | 7.47
2009-05-30 | 7:31 | 8.25 | 27 | 18 | 12.19 | 103.6 | 7.47
2009-05-30 | 7:46 | 8.18 | 27 | 18 | 12.22 | 103.6 | 7.47
2009-05-30 | 8:01 | 8.11 | 28 | 18 | 12.31 | 104.2 | 7.48
2009-05-30 | 8:16 | 8.06 | 28 | 18 | 12.29 | 103.9 | 7.48
2009-05-30 | 8:31 | 8.03 | 28 | 18 | 12.36 | 104.4 | 7.49
2009-05-30 | 8:46 | 8.02 | 28 | 18 | 12.38 | 104.6 | 7.5
2009-05-30 | 9:01 | 8.05 | 28 | 18 | 12.39 | 104.7 | 7.5
2009-05-30 | 9:16 | 8.1 | 28 | 18 | 12.39 | 104.8 | 7.51
2009-05-30 | 9:31 | 8.18 | 28 | 18 | 12.37 | 104.9 | 7.51
2009-05-30 | 9:46 | 8.3 | 28 | 18 | 12.28 | 104.5 | 7.52
2009-05-30 | 10:01 | 8.45 | 28 | 18 | 12.27 | 104.8 | 7.53
2009-05-30 | 10:16 | 8.6 | 28 | 18 | 12.24 | 104.9 | 7.54
2009-05-30 | 10:31 | 8.8 | 28 | 18 | 12.14 | 104.5 | 7.55
2009-05-30 | 10:46 | 9.01 | 28 | 18 | 12.05 | 104.3 | 7.56
2009-05-30 | 11:01 | 9.24 | 28 | 18 | 12.03 | 104.7 | 7.57
2009-05-30 | 11:16 | 9.49 | 28 | 18 | 11.99 | 104.9 | 7.57
2009-05-30 | 11:31 | 9.75 | 28 | 18 | 11.87 | 104.6 | 7.58
2009-05-30 | 11:46 | 10.04 | 28 | 18 | 11.83 | 104.9 | 7.58
2009-05-30 | 12:01 | 10.33 | 28 | 18 | 11.76 | 105 | 7.59
2009-05-30 | 12:16 | 10.63 | 28 | 18 | 11.68 | 105 | 7.6
2009-05-30 | 12:31 | 10.95 | 28 | 18 | 11.58 | 104.9 | 7.59
2009-05-30 | 12:46 | 11.3 | 28 | 18 | 11.48 | 104.8 | 7.61
2009-05-30 | 13:00 | 11.64 | 28 | 18 | 11.39 | 104.8 | 7.62
2009-05-30 | 13:15 | 11.98 | 28 | 18 | 11.24 | 104.2 | 7.62
2009-05-30 | 13:30 | 12.32 | 29 | 19 | 11.12 | 103.9 | 7.63
2009-05-30 | 13:45 | 12.65 | 29 | 19 | 11.01 | 103.7 | 7.63
2009-05-30 | 14:00 | 12.97 | 29 | 19 | 10.91 | 103.5 | 7.63
2009-05-30 | 14:15 | 13.2 | 29 | 19 | 10.86 | 103.5 | 7.63
2009-05-30 | 14:30 | 13.32 | 29 | 19 | 10.83 | 103.5 | 7.64
2009-05-30 | 14:45 | 13.39 | 29 | 19 | 10.74 | 102.8 | 7.64
2009-05-30 | 15:00 | 13.68 | 29 | 19 | 10.69 | 103 | 7.64
2009-05-30 | 15:15 | 13.85 | 29 | 19 | 10.58 | 102.3 | 7.65
2009-05-30 | 15:30 | 13.99 | 29 | 19 | 10.51 | 101.9 | 7.65
2009-05-30 | 15:45 | 14.13 | 29 | 19 | 10.51 | 102.3 | 7.66
2009-05-30 | 16:00 | 14.23 | 30 | 19 | 10.5 | 102.4 | 7.66
2009-05-30 | 16:15 | 14.14 | 30 | 19 | 10.46 | 101.8 | 7.66
2009-05-30 | 16:30 | 13.98 | 30 | 19 | 10.48 | 101.6 | 7.65
2009-05-30 | 16:45 | 13.79 | 30 | 19 | 10.5 | 101.5 | 7.66
2009-05-30 | 17:00 | 13.69 | 30 | 20 | 10.56 | 101.8 | 7.65
2009-05-30 | 17:15 | 13.65 | 30 | 20 | 10.53 | 101.4 | 7.64
2009-05-30 | 17:30 | 13.55 | 30 | 20 | 10.53 | 101.1 | 7.63
2009-05-30 | 17:45 | 13.55 | 31 | 20 | 10.57 | 101.6 | 7.63
2009-05-30 | 18:00 | 13.61 | 31 | 20 | 10.56 | 101.6 | 7.64
2009-05-30 | 18:15 | 13.67 | 31 | 20 | 10.5 | 101.2 | 7.63
2009-05-30 | 18:30 | 13.71 | 31 | 20 | 10.52 | 101.4 | 7.62
2009-05-30 | 18:45 | 13.76 | 31 | 20 | 10.52 | 101.5 | 7.63
2009-05-30 | 19:00 | 13.84 | 31 | 20 | 10.49 | 101.4 | 7.63
2009-05-30 | 19:15 | 13.92 | 31 | 20 | 10.47 | 101.4 | 7.63
2009-05-30 | 19:30 | 13.94 | 31 | 20 | 10.42 | 101 | 7.63
2009-05-30 | 19:45 | 13.95 | 31 | 20 | 10.4 | 100.8 | 7.63
2009-05-30 | 20:00 | 13.96 | 31 | 20 | 10.43 | 101.1 | 7.63
2009-05-30 | 20:15 | 13.97 | 31 | 20 | 10.39 | 100.7 | 7.63
2009-05-30 | 20:30 | 13.97 | 31 | 20 | 10.37 | 100.6 | 7.63
2009-05-30 | 20:45 | 13.94 | 31 | 20 | 10.42 | 101 | 7.62
2009-05-30 | 21:00 | 13.92 | 31 | 20 | 10.44 | 101.1 | 7.61
2009-05-30 | 21:15 | 13.86 | 31 | 20 | 10.46 | 101.2 | 7.6
2009-05-30 | 21:30 | 13.76 | 31 | 20 | 10.5 | 101.3 | 7.6
2009-05-30 | 21:45 | 13.61 | 31 | 20 | 10.55 | 101.5 | 7.6
2009-05-30 | 22:00 | 13.39 | 31 | 20 | 10.57 | 101.2 | 7.59
2009-05-30 | 22:15 | 13.16 | 31 | 20 | 10.62 | 101.1 | 7.59
2009-05-30 | 22:30 | 12.97 | 31 | 20 | 10.71 | 101.6 | 7.58
2009-05-30 | 22:45 | 12.84 | 31 | 20 | 10.77 | 101.8 | 7.57
2009-05-30 | 23:00 | 12.72 | 31 | 20 | 10.81 | 102 | 7.57
2009-05-30 | 23:15 | 12.57 | 31 | 20 | 10.86 | 102.1 | 7.56
2009-05-30 | 23:30 | 12.4 | 30 | 20 | 10.92 | 102.2 | 7.55
2009-05-30 | 23:45 | 12.23 | 30 | 19 | 10.92 | 101.9 | 7.55
2009-05-31 | 0:00 | 12.11 | 30 | 19 | 10.99 | 102.2 | 7.54
2009-05-31 | 0:15 | 12.04 | 29 | 19 | 11.02 | 102.3 | 7.53
2009-05-31 | 0:30 | 12 | 29 | 19 | 11.04 | 102.5 | 7.53
2009-05-31 | 0:45 | 11.97 | 29 | 19 | 11 | 102 | 7.52
2009-05-31 | 1:01 | 11.91 | 29 | 19 | 11.05 | 102.3 | 7.51
2009-05-31 | 1:16 | 11.84 | 29 | 19 | 11.09 | 102.5 | 7.51
2009-05-31 | 1:31 | 11.78 | 29 | 19 | 11.05 | 102.1 | 7.51
2009-05-31 | 1:46 | 11.72 | 29 | 19 | 11.11 | 102.4 | 7.51
2009-05-31 | 2:01 | 11.65 | 29 | 19 | 11.14 | 102.5 | 7.51
2009-05-31 | 2:16 | 11.58 | 28 | 18 | 11.16 | 102.6 | 7.5
2009-05-31 | 2:31 | 11.48 | 27 | 18 | 11.2 | 102.7 | 7.49
2009-05-31 | 2:46 | 11.37 | 27 | 18 | 11.23 | 102.7 | 7.49
2009-05-31 | 3:01 | 11.24 | 27 | 17 | 11.23 | 102.4 | 7.49
2009-05-31 | 3:16 | 11.11 | 27 | 18 | 11.25 | 102.2 | 7.49
2009-05-31 | 3:31 | 10.97 | 27 | 18 | 11.33 | 102.7 | 7.47
2009-05-31 | 3:46 | 10.83 | 27 | 17 | 11.4 | 103 | 7.47
2009-05-31 | 4:01 | 10.69 | 27 | 18 | 11.45 | 103.1 | 7.48
2009-05-31 | 4:16 | 10.54 | 27 | 17 | 11.5 | 103.2 | 7.48
2009-05-31 | 4:31 | 10.38 | 27 | 17 | 11.55 | 103.3 | 7.48
2009-05-31 | 4:46 | 10.22 | 27 | 17 | 11.61 | 103.4 | 7.46
2009-05-31 | 5:01 | 10.07 | 27 | 17 | 11.61 | 103 | 7.47
2009-05-31 | 5:16 | 9.91 | 27 | 17 | 11.65 | 102.9 | 7.46
2009-05-31 | 5:31 | 9.75 | 27 | 17 | 11.74 | 103.4 | 7.46
2009-05-31 | 5:46 | 9.61 | 27 | 17 | 11.75 | 103.1 | 7.45
2009-05-31 | 6:01 | 9.47 | 27 | 17 | 11.78 | 103.1 | 7.45
2009-05-31 | 6:16 | 9.35 | 27 | 17 | 11.87 | 103.6 | 7.45
2009-05-31 | 6:31 | 9.24 | 26 | 17 | 11.88 | 103.4 | 7.45
2009-05-31 | 6:46 | 9.14 | 27 | 17 | 11.96 | 103.8 | 7.45
2009-05-31 | 7:01 | 9.05 | 27 | 17 | 12.01 | 104 | 7.46
2009-05-31 | 7:16 | 8.97 | 27 | 17 | 12 | 103.7 | 7.47
2009-05-31 | 7:31 | 8.91 | 27 | 17 | 12.06 | 104.1 | 7.47
2009-05-31 | 7:46 | 8.85 | 27 | 17 | 12.04 | 103.8 | 7.48
2009-05-31 | 8:01 | 8.8 | 27 | 17 | 12.1 | 104.2 | 7.48
2009-05-31 | 8:16 | 8.77 | 27 | 17 | 12.14 | 104.4 | 7.49
2009-05-31 | 8:31 | 8.73 | 27 | 17 | 12.1 | 104 | 7.49
2009-05-31 | 8:46 | 8.7 | 27 | 17 | 12.16 | 104.4 | 7.49
2009-05-31 | 9:01 | 8.7 | 27 | 17 | 12.18 | 104.6 | 7.5
2009-05-31 | 9:16 | 8.7 | 27 | 17 | 12.15 | 104.3 | 7.5
2009-05-31 | 9:31 | 8.68 | 27 | 17 | 12.19 | 104.7 | 7.51
2009-05-31 | 9:46 | 8.67 | 27 | 18 | 12.21 | 104.8 | 7.52
2009-05-31 | 10:01 | 8.67 | 27 | 17 | 12.22 | 104.9 | 7.52
2009-05-31 | 10:16 | 8.7 | 27 | 18 | 12.23 | 105 | 7.53
2009-05-31 | 10:31 | 8.74 | 27 | 18 | 12.18 | 104.7 | 7.53
2009-05-31 | 10:46 | 8.78 | 27 | 18 | 12.22 | 105.1 | 7.53
2009-05-31 | 11:01 | 8.85 | 27 | 18 | 12.21 | 105.3 | 7.54
2009-05-31 | 11:16 | 8.99 | 27 | 18 | 12.19 | 105.5 | 7.55
2009-05-31 | 11:31 | 9.16 | 27 | 18 | 12.12 | 105.2 | 7.56
2009-05-31 | 11:46 | 9.37 | 27 | 18 | 12.05 | 105.1 | 7.58
2009-05-31 | 12:01 | 9.68 | 27 | 18 | 11.96 | 105.2 | 7.56